vc/22-Sep-03

### Week 39 starting 22 Sept 03

### SM18 Operation SCHEDULE Meeting (Tuesdays)

Schedule for this week :

| BENCH | CURRENT MAGNET | STATUS Today(23/09/03) | Forward look, & Next Magnet |
| --- | --- | --- | --- |
| TBA1 | 2012 | Should be Warming Up | Bench to be de-activated – preparation for D bench |
| TBA2 | 3042 | Under Connection | Cold Tests Start Thurs night? |
| TBC1 | 2021 | Need to connect | Cooldown by Fri.Start Cold Tests start on Fri Night during w’end |
| TBC2 | SSS1 | cold | Cold tests |
| TBE1 | 3043 | Warm | Removal – bring new 2014 |
| TBE2 | 2010 | cold | Cold tests |
| TBF1 | 2025 after F2 done? |  | Prepare For Commissioning AFTER F2 |
| TBF2 | 2025 proposed |  | Prepare For Commissioning FIRST |

Test Coordinator this week:  N Smirnov;  MTM  Technician on call:  (16-3602) :

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Other General Issues

- New additional & easy  Tel nos.   70900  as well as 16 0060

- Paper Logbook  & Files

- Problems from Operation Last week (Stephane + Test Supervisor )

| Date | Problem | Type | Application/soft. | Action/advise |
| --- | --- | --- | --- | --- |
| Tuesday 16_M | TBE 1 : TEMA cannot be launched | AB/CO | TEMA | Script, which was managing the conflicts between applications, gave errors. Solved by Pedro. |
| Tuesday 16_A | TBE1 : IAP@cold Errors in compensation out of the limits defined for one channel. | MTM |  | Accepted by Giorgio. Origin not found. |
| Tuesday 16_A | TBE1: IDCCT current reading not working Problem in LHC cycle Load-line without read current. | AB/CO |  | Solved the 17 morning DDCT replaced by DDCT of cluster B . No calibration so read-out connected in the bottom of the DDCT. |
| Tuesday 16_N | TBE1 : IAP@warm cannot be performed | MTM |  | Connector inverted |
| Tuesday 16_N | MM : TBA2 MUX errors | AB/CO |  |  |
| Tuesday 16_N | TBA1 : IAP @ warm could not be performed. | MTM |  | Cards missing in HF/LF crates used by cluster F. |
| Wednesday 17_M | Full load line (TBE1): wrong current reading |  |  | Run recovery and then OK Origin not found |
| Thursday 18_M | TBE1: cable coupling current MM5 | MTM | TEMA? | Wrong cycle selected Magnetiz |
| Saturday 20_M Sunday 21_M | TBA1: unable to launch IAP cold_X2 from TEMA | AB/CO | IAP | IAP done manually. Problem under investigation |
| Saturday 20_A | TBE2: Tanticrysotat no visible in the control room Communication problem between PLCART and cryo supervision | MTM/ACR |  | Local check of the good functioning of the anticryostat. |
| Saturday 20_N | TBE1: Rotary switch blocked | MTM |  | The handling of the switch was not correctly done. See procedure. |
| Saturday 20_N | TBE1 : IAP at cold. One card faulty (8 channels out of range) in aperture 2 | MTM | IAP |  |
| Saturday 20_N Sunday 21_M | TBA1 : Vi corrupted in the MMP version because of an update of the library which was not working in TBA1 | AB/CO | MMP | Adriaan called : Launched the old version of MMP. Library updated Monday morning |
| Sunday 21_M | New PC : No macro available for the quench table. | MTM |  | Security setting corrected Solved (Daniel) |
| Sunday 21_M | Vclamp trigger could not be cleared in TBC1???? | MTM |  | ???? |
| Sunday 21_N | MM in TBA1 launched with MMP old version Full load line: Motor in aperture 1 is turning but not the shaft. | AB/CO |  | Then communication error with one motor during auto test. Solved by switching to another TRU and back on display??? |
| Monday 22_M | Full load line in TBA1: problem with end stops | MTM |  |  |
| Monday 22_M | Still problems with anti cryostat reading in the control room | MTM/ACR |  | Not solved. Local check |
| Monday 22_M | FTP errors: unable to launch TEMA |  |  | Automatic update of the FTP at CERN |
| Monday 22_M | TBE2: MM positioning All the sectors with b11 in orange and MM position OK? | MTM |  | AP1/AP2 connectors mixed when connected to TRU Table for MM errors in display . |
| Monday 22_A | TBE2: LHC cycle End field polarity error not seen | MTM |  | Switch on a instead of e coil in 0001 for shaft 25. |