vc/2-Oct-03

### Week 40 starting 29 Sept 03

### SM18 Operation SCHEDULE Meeting (Tuesdays)

Schedule for this week :

| BENCH | CURRENT MAGNET | STATUS Today(30/09/03) | Forward look, & Next Magnet |
| --- | --- | --- | --- |
| TBA1 | 3047 | Should be started cooling down today | Cold Tests Start Thurs ?After this magnet, Bench to be de-activated – preparation for D bench |
| TBA2 | 3042 | cold | Cold Tests StOP Thurs morn? |
| TBC1 | 2021 | Need to remove today | Next Magnet 3045; Connect & Cooldown by Fri. midday; Start Cold Tests on Fri during w’end |
| TBC2 | SSS1 | cold | Cold tests |
| TBE1 | 2002 | Finish Connection ; Should be started cooling down today | COOLDOWN START ; Cold Tests Start Thurs ? |
| TBE2 | 2010 | Need to remove today | Next Magnet 3039; |
| TBF1 | 2025 after F2 done? |  | Prepare For Commissioning AFTER F2 |
| TBF2 | 2025 proposed |  | Prepare For Commissioning FIRST |

Test Coordinator this week:  S.Sanfilippo;  MTM  Technician on call:  (16-3602) : Federico

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Other General Issues

- Problems from Operation Last week (Stephane + Test Supervisor )

| Date | Problem | Type | Application/soft. | Action/advise |
| --- | --- | --- | --- | --- |
| Tuesday 23_A | TBA2 :IAP@warm : Rheaters 25 ohms | MTM | IAP | Incorrect coefficients |
| Tuesday 23_A | TBA2:LF Acquisition : First led off but message acquisition armed | MTM | LF | PC reset and OK |
| Tuesday 23_A | TBA2 : HV test could not be launched because cable on passive bus bar was broken | MTM |  | Cable from TBA1 taken. Repaired by Richard. |
| Tuesday 23_A | TBA2 : HV test could not be launched because power supply current leads connected | MTM |  | Test coordinator has to plan the disconnection before |
| Thursday 25_M | Mobile rack for HV test used in C1 | MTM |  | Connection cables for passive burs bar broken. Richard was not informed |
| Thursday 25_A Friday 26_A | HV test at warm (TBE) capacitance problem. | MTM | HV | Equip.support called. New version of soft not ready. |
| Thursday 25_N | TBA2 : Problem with T anticryostat | MTM |  | Equip.support called. Resolved. |
| Friday 26_A | Shaft 21 (TBC1) : S02 coil e open. | MTM |  | Trace to be kept. |
| Friday 26_A | TBE2: Min energy at cold LF config not found | MTM | LF | ???? |
| Friday 26_N | TBC1: Provoked quench all LF Faulty quench in bus bar | AB/CO | AQA | No MIITS, No current , Origin of Quench +Vext Saturday : Nicolai decide to continue the training. Software problem on investigation. |
| Friday 26_N | TBA2: HV test Dipole–QH vs Ground >limit | MTM |  | OK the morning after. Humidity problem in leads? |
| Saturday 27_N | TBA2: IAP at cold D1-L-U (144 mV) AC compensation | MTM | IAP | Tuning of the channel? |
| Sunday 28_A | TBA2: Training quench 2 : Wrong extension file given | MTM |  | Extension .a02 |
| Sunday 28_N | TBA2: Motor problem during full load line MM on TBA2: Field polarity error on ap2 | MTM |  | Aborted and re-launched. Check carefully the cases explained in the table. |
| Monday 29_A | TBA12: TEMA not working for IAP at warm | AB/CO? | TEMA | TEMA launched in local OK |
| Monday 29_A | TBA2: Power supply did not start. Magnet shorted to the ground. | MTM |  | Cables coming from the patch panel used for RRR removed and OK. |