vc/21-Oct-03

### Week 43 starting 20 Oct 03

### SM18 Operation SCHEDULE Meeting (Tuesdays)

Schedule for this week :

| BENCH | CURRENT MAGNET | STATUS Today(21/10/03) | Forward look, & Next Magnet |
| --- | --- | --- | --- |
| TBA1 | 3042 | Leak test finished –remove 3042 | Bench de-activated – preparation for D bench “Park” here 2001 ( moved from D bench ) |
| TBA2 | 1020 | Disconnect & Remove | Next Magnet 1049 |
| TBC1 | 3044 | Disconnect & Remove | Next Magnet 2024 |
| TBC2 | SSS1 | cold | Diode Cold tests then warm up & remove; bring new SSS |
| TBE1 | 2028 | Connect & cooldown | START COLD TESTS |
| TBE2 | 1050 | Cold | Need to undergo Thermal Cycle |
| TBF1 | 2002 | cold | Commissioning/powering tests |
| TBF2 | 2025 | Warm | Next Magnet 3058 : as first time Usage of F bench for Operation |

Test Coordinator this week:  Olav Berrig;  MTM  Tech. on call:  (16-3602) : Guy Deferne

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Other General Issues

  - Loading up time &  AQA Analyses files(Daniel)

  - Traveller Followup – “I” shift 0830-1730

  - F Bench Commissioning

  - Interactions in Traveller & Info  ( Q-Table, HV Tests)

  - AOB

Current Bench Status:

# SM18 OPERATION WEB PAGE

| Bench | Magnet | Test | Temp. | Time | Current |
| --- | --- | --- | --- | --- | --- |
| TBA1 | _ | _ |  | 10/10/2003 9:20:27 AM | 0 |

| TBA2 | MBBR1020 | Finished Ready to go |  | 10/20/2003 6:31:23 PM | 0 |
| --- | --- | --- | --- | --- | --- |

| TBB1 | _ | _ |  |  | 0 |
| --- | --- | --- | --- | --- | --- |

| TBB2 | _ | _ |  |  | 0 |
| --- | --- | --- | --- | --- | --- |

| TBC1 | MBAL3044 | Finished Ready to go |  | 10/20/2003 6:31:40 PM | 0 |
| --- | --- | --- | --- | --- | --- |

| TBC2 | SSS001 | _ |  | 10/19/2003 12:35:33 PM | 0 |
| --- | --- | --- | --- | --- | --- |

| TBD1 | _ | _ |  |  | 0 |
| --- | --- | --- | --- | --- | --- |

| TBD2 | _ | _ |  |  | 0 |
| --- | --- | --- | --- | --- | --- |

| TBE1 | MBAR2028 | Prep6.3 QH Measure | 300 K | 10/21/2003 10:15:00 AM | 0 |
| --- | --- | --- | --- | --- | --- |

| TBE2 | MBBL1050 | PT30 HV Thermal Cycle | Warming | 10/21/2003 11:27:52 AM | 0 |
| --- | --- | --- | --- | --- | --- |

| TBF1 | MBBR2002 | PT3 HV Cold Test | 1.9 K | 10/20/2003 1:57:42 AM | 0 |
| --- | --- | --- | --- | --- | --- |

| TBF2 | MBBR2025 | Finished Ready to go | 300 K | 10/21/2003 10:00:00 AM | 0 |
| --- | --- | --- | --- | --- | --- |

- Problems from Operation Last week (Stephane + Test Supervisor )

| Date | Problem | Type | Application/soft | Action/advise |
| --- | --- | --- | --- | --- |
| Tuesday 13_A | TBF2 Training quench ; Quench at 10437A. Je ne sais pas ou est passe le fichier de ce quench : L0310142027.a02 |  |  |  |
| Tuesday 13_N | TBA2: Final HV test at cold. There is no test sequence for HV at cold with Power leads connected to power converter. Test launched and sequence 4 onwards selected after run recovery. |  |  |  |
| Wednesday 14_M | TBE1: SSL. Server lost on sunmta26 |  |  | restarted OK |
| Thursday 15_M | TBE2: IAP @ Warm ; HF Discharge; The test could not launched as magair300_x2 and mag300_x2 interlock signals are not clearing. Equipment support was contacted and Richard is on the job. |  |  | - Cryo work on Cluster D took the compressed air supply from Cluster E, causing the interlock. OK after cryo had finished. |
| Thursday 16_N | TBF2: No HV test at cold and no Thermal cycle because the rotary switch is blocked |  |  |  |
| Friday 17_A | TBA2 - DISCONNECTION start for ICS TBE1 - CONNECTION of 2028 CONTINUE for ICS FOR BOTH POINTS ICS WILL NOT COME during weekend & so now WE ARE WAITING FOR Them for 2 days !! |  |  | ICS will soon (November) have 3 shifts i.e. 7/7 & 24/24 |
| Friday 17_A | *** Cold Tests on TBE2 ( 1.9 K tomorrow (saturday) : Here Shafts need to be installed but there is a problem due to anticryostat being too warm for reasons unknown. Patrick Legrand informed |  |  |  |
| Friday 17_A | TBC1 : Two Q Loc cards not getting compensated. D1_L-U potaim card was not getting compensated. Exchanged the card from A1 crate as suggested by Richard. Problem solved. The card is kept in the almirah in the control room on which rota is pasted. |  |  |  |
| Saturday 18_A | TBE2: IAP at cold. Vgas_Int_lead exceeds from 10 mv( offset) to 10.7 mv but could not be adjusted. Some errors in Qloc cards could not be adjusted. |  |  |  |
| Saturday 18_N | TBE1:MM position problem CPU encoder error. As per the instruction from Richard, we reset the CPU's by switching off the power supply. The MM position measurement could be carried out without any further problem. |  |  | In case of MM problems, do not switch off the MMP program, but go to the hardware panel. In the hardware panel, it is possible to switch off/on the VME crates. It is not necessary to switch off the MMP program. |
| Saturday 18_N | TBC1: We wanted to carry out the HV insulation test at cold ; However, "Run Sequence " did not succeed for the first test. It was not even giving the voltage of 650 V. The megger was giving red beeps and displaying "burn". We checked all cables and found to be alright. We found that the ground cable from the magnet to the ground bus cable had not been properly connected. We have connected it by twisting the cables but this requires improvement. Though this may not be the reason for not able to proceed further, this requires to be improved. After this, the megger was able to give voltage but once the measurement for that step was over, the message "Capacitance values exceed exspectation " came which could not be cleared despite repeated attempts . Hence we could not proceed further. |  |  | We tried to do tests step-by-step; Steps involving dipoles did not succeed while others could be done. |
| Saturday 18_N | TBE2: Minimum Energy Quench at 3000 A; 1.9 K. Vquench is between 700 and 710 volts. The data was analyzed for the diode test as required by Eric Flock. A print-out of the same is kept in the log. However, we did not understand how to calculate the diode joint resitance. Next shift may talk to Eric Flock Or Nick for further instructions on this. THE DESCRIPTION OF THIS TEST IS KEPT IN "Quench Procedures" FOLDER(Name : "Brief Description of the method used to measure the diode contact resistances") |  |  |  |
| Sunday 19_M | TBC1: Mimimun Energy @6.5kA +RRR, Vquench=540V NOTE: LF configuration file is not found. |  |  | Remember to wait for the LF data to be transferred, before doing the RRR test. |
| Sunday 19_M | TBE1: We suspect some problem with anticryostat A1L showing a 0 value since 7:00 pm Problem is confirmed. Anticryostat temperature is 58 C. Richard, Nick informed P. Legrand contacted by Nick. Didier Cotte contacted by myself, he will come. Everybody is here. Didier removes the shafts. P. Legrand found a bad contact on A2L connector. |  |  |  |
| Monday 20_M | TBF2: IAP @ warm did not work from TEMA. Giorgio believes it is a problem with the PLCSE. He has sent an e-mail to Claude Dehavay to solve it. |  |  |  |
| Monday 20_A | TBE2: Terminal screen kaput - switches off by itself. Magnus Bjork called by Adriaan. |  |  |  |

| Date | Problem | Type | Application/soft. | Action/advise |
| --- | --- | --- | --- | --- |
| Tuesday 14_M | TBA2: MM LHC export blocked | MTM | MMP | Problem under investigation. |
| Tuesday 14_M | TBF2: Quench at 6.5 kA MM positioning instead of Q.loc | MTM |  |  |
| Tuesday 14_M | TBF2: Training quench 2 File L0310142027.a02 lost | MTM |  |  |
| Tuesday 14_M | TBA2: No setting HV test final at cold with current leads connected. | MTM |  | Test launched and start and sequence 4 selected after run recovery (instruction from Richard). |
| Tuesday 14_M | TBA2: No setting HV test final at cold with current leads connected. | MTM |  | Test launched and sequence 4 selected after run recovery (instruction from Richard). |
| Wednesday 15_N | TBF2: No data for anti-cryostat in the control room. | ACR? |  |  |
| Wednesday 15_N | TBF2: Cryo problem | ACR |  |  |
| Thursday 16_M | TBE2: (IAP @ warm) HF discharge could not be launched. Mag_air_300_X2, Mag_300_X2 interlocks. | ACR |  | Compress air supply from cluster E used for work in cluster D? |
| Friday 17_M | TBF2: Anti-cryostat at –180 C due to a lost of vacuum inside the cold mass! (information from P.Legrand) | ACR |  |  |