Tss16-Feb-04

### Week 08  starting 16 Feb 04

### SM18 Operation SCHEDULE Meeting (Tuesdays)

Bench status  and tests for this week :

Test Coordinator this week:  Nicolay 164674 ;  MTM  Tech. on call Giorgio:  (16-3602) :

Magnet coming soon:

3092, 3077,(without anticryostat) , 3078,  2050 (with anticryostat).

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Other General Issues

  - TBD1 (1037) : Insulation  dipole vs. ground, good at warm, failed below 250 K. Special investigation.

  - Quench for warm up : Faulty quench performed for 1044. It has to be performed correctly. Useful to gain information (rf: 1033 with change of Iearth).

  - Case of 3030: After the first training  (Iq>12000 A) no ramp to 9 T was recorded. A second ramp to 9 T is needed in that case.

  - Scan of equipment ( do not forget to scan the gamelles,  special attention when shafts are entered during the training).

  - HV tests : megger stops during the tests. It happens with several mobile racks.

  - Truck left without key: difficult to remove shaft from D2.

| Date | Problem | Type | Application/soft. | Action/advise |
| --- | --- | --- | --- | --- |
| Tuesday_10_M | TBE2: Scanning not possible for the mobile rack TBD2: Full load-line: HWRC says that config file missing. | MTM | HWREC | Entered manually. Maryline fixed the problem. MMP Config. files are done manually for the moment. Guy performed manually the file. |
| Tuesday_10_N | TBD2: Launch of thermal cycle. TEMA blocked. HRWC in TEMA indicates “gamelles missing”. | MTM | HWREC | The scan of the gamelles was forgotten by OP team. |
| Tuesday_10_N | MTF does not up-load files Oracle error =6508 You are not identified as a valid user. | MTM | HWREC | SM18 op not valid as login for HWREC in EDMS. Operators have to use their own login to keep a trace. |
| Tuesday_10_N Wednesday_11_M Wednesday_11_A Friday _13_M | TBD1: HV test at cold failed: dipole vs ground failed at 320 V | MTM |  | Investigation at cold performed by Guy, Richard , Kevin and Stephane. Problem not in the IFS box Magnet warm up. Disconnected and reconnected again. |
| Thursday_12_N | TBF1: No Qloc for quench number 3 | MTM |  | Guy found that two cables from the switch to the chassis (Qloc, MM, FA) were not connected (new mobile rack). |
| Wednesday_11_A Thursday_12_M Thursday_12_N | TBF1: IAP @ cold : Two identical files created while triggering. The second file did not contain useful data. Problems in Qloc compensation (TBF1), offsets in potaim cards ( TBE2). | MTM | IAP |  |
| Thursday_12_N | TBC1:MM positioning not possible. CPU address error | AB/CO |  | Hubert/Adrian changed the CPU card. |
| Thursday_12_N | TBF1: MM positioning Over-range in one head sector | MTM |  | Cable connecting switch to amplifiers not properly pushed. Guy advised to verify that the LEMO are pushed properly when this type of problem occurred. Change from a to e coil for the two head sectors during the night? |
| Friday_13_N | TBC1:Using TEMA : message error receiving “ INFO from UDP socket” | ? |  | Maryline advised to stop all the applications and restart sun machine connected to C bench.. |
| Friday_13_N | TBC1: Quench for warm up. Start training then stop at 12100 A and trigger cards. Faulty quench .No information | MTM |  | Request of Stephane: Quench for warm up is also important: we can gain information and help test coordinator. It should be performed correctly. |
| Saturday_14_M | TBF2 : Server crashed. | IT? AB/CO? |  | Maryline came and fixed the problem. |
| Saturday_14_A | Cluster E : Access extremely slow to analyse the quenches from the control room. | IT? AB/CO? |  | Analyse in local. (cluster E) |
| Sunday _15_M | TBE1: Training quenches. MIITS too high (33.9 then 35). Pick voltage at 128 V. | MTM |  | Check :Discharge of heaters. Signals from V taps . Iearth HV test OK Stephane asks to proceed with minimum energy. |
| Saturday_14_N | TBF2, TBE1: unable to upload CDPT to MTF | TS/CSE | MTF | Quit and re-launch the browser. If problem persists send e-mail + screen shots to EDMS support. |