VC/ 9-Mar-04

### Week 10  starting 01 March 04

### SM18 Operation SCHEDULE Meeting (Tuesdays)

Bench status  and tests for this week :

| Bench | Name |  |  |
| --- | --- | --- | --- |
| TBC1 | MBAR3062 | Needs new before W'end |  |
| TBC2 | SSS002 | Cold Tests This Week | cooling? |
| TBD1 | 3032 | Arrived Monday 1/03/04 |  |
| TBD2 | MBAR3051 |  | cold |
| TBE1 | 1054 | arrived Monday 1/03/04 |  |
| TBE2 | MBBR3077 |  | cold |
| TBF1 | MBAR2034 |  | cooling? |
| TBF2 | MBBR3037 |  |  |
|  | Forward look : 3068 , 2027 with AntiC |  |

Test Coordinator this week:  John Miles 16-4657 ;  MTM  Tech. on call :  (16-3602) .

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Cryogenics  2/29/2004 9:30:00 AM     Quench Recovery very slow  took 8 hours for recovery from previous quench

E1

- Similar problem C1 for recovery Sun night/Mon morning on 3062  ( > 7 hr Quench recovery)

- Cryogenics  2/29/2004 12:42:12 PM     Cooling down very slow.  
Expected 4.5K on 28th evening 6:00 PM still at 4.5K

C1

- Other General Issues

  - Water Leak on Tues night :  OP Actions + H.Thiesen , Giorgio

  - Priority List in SMCR includes shafts/no shafts

  - ICS callout in e-logbook for info to next shift

  - ICS 06 numbers & pincode on 70900

  - AOB

| Date | Problem | Type | Application/soft. | Action/advise |
| --- | --- | --- | --- | --- |
| Tuesday 24_A | TBE1: heavy water leak above power converter. Break of a the water cable connection. | TS/EL |  | Intervention of Ewald, Giorgio. Water taps closed by J.P Lamboy. Fire brigade called. RME and RMB disconnected by H.Thiesen. Restart TBE Friday evening. |
| Thursday 26_A | TBF2: Sun MTA28 blocked. Server down. | AB/CO-MTM |  | Advise from Alessandro : Use the exit button but then wait for a complete exit without clicking in the abort button. |
| Thursday 26_N | TBF1: LHC cycle (at the end) Full load line. Aperture 2: the shaft is not rotating well. Consequence: CPU error address. Encoder pulse is not received so current reading was missing Cannot re-start with run-recovery | MTM |  | Ewald reduced the speed of rotation to decrease the scratching. Action : Frederic Flamand will repair the TRU 3 after the test. |
| Friday 27_A | TBE1: PC fault launching IAP at cold. Problem with the position detector of the rotary switch. | MTM |  | Fault disappeared after several trials. Giorgio will fix this problem. |
| Saturday 28_A | TBE1: Local sun blocked. Server blocked. | AB/CO-MTM? |  | Ewald asks to kill the applications with “xkill”. |
| Saturday 28_N | TBD1: No signal quench received by Cryo. | ACR |  | J.P.Lamboy is investigating. |
| Saturday 28_N | Cooling down too long : Message from J.P. Lamboy: Nominal 3 magnets cold . 2 at 1.9 K and 1 recovery after a quench. | ACR/MTM |  | Advise : Shift leader will set priorities . |
| Sunday 29_N | TBE1: No power in the 220 V outlet test bench. The multipole socket is no powered for safety following the accident of Tuesday. | AB/PO |  | R.Mompo aware. Action with P.Legrand. |
| Sunday 29_N | TBD1: No HV test sequence in MATEST on cluster D ( to put off hf , lf acquisitions) | MTM |  | R.Mompo aware : Implementation with Giorgio. |
| Sunday 29_N | TBD1: TEMA has to be re-started many times. | AB/CO |  | Restart applications. |
| Monday 1_M | TBD2: lost of cryo OK during a current ramp | ACR |  |  |