VC/ 9-Mar-04

### Week 11  starting 08 March 04

### SM18 Operation Meeting (Tuesdays)

Bench status  and tests for this week :

| Bench | Name |  | Temp | - | - |
| --- | --- | --- | --- | --- | --- |
| TBC1 | MBBR2052 | NEW |  |  |  |
| TBC2 | SSS002 | EX1 Extra Test |  |  |  |
| TBD1 | MBBL3071 | PT 1 IAP @ cold |  |  |  |
| TBD2 | MBBR3063 | NEW |  |  |  |
| TBE1 | MBAL3105 | ICS 3 |  |  |  |
| TBE2 | MBAL1056 | NEW |  |  |  |
| TBF1 | MBBR1041 | PT 12 Warm Up |  |  |  |
| TBF2 | MBBR2014 | CR 1 Leak Test |  |  |  |

Test Coordinator this week:  Pierre.Pugnat  ;  MTM  Tech. on call :  (16-3602) .

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Other General Issues

  - Priority List in SMCR includes shafts/no shafts

  - AOB

| Date | Problem | Type | Application/soft. | Action/advise |
| --- | --- | --- | --- | --- |
| Tuesday_02_A Saturday_06_M | TBE1:C11 pins damaged on magnet 1054. Bad manipulation by ICS during HV test Same on 2027 (TBD2) | ICS |  | Richard replaced the connectors. OK But no spare now. ICS responsible aware. |
| Tuesday_02_A | TBE1:HF not getting armed. | MTM | HF, LF | HF, LF switched off. Crates powered off. OK |
| Tuesday_02_A | Mobic was not functioning correctly. Windows was hung. | MTM |  | Richard advised to follow the procedure written by Maryline :Reset the program. Then OK |
| Wednesday_03_M | TBF2 : IAP @cold . HRWREC mentioned that a fixed equipment was not recognized (the generator for AC compensation) | MTM |  | Solved by Federico. Keep the equipment in place. |
| Wednesday_03_A | TBD1: Trigger not received by CTH during a quench (mentioned last week) | ACR |  | Process corrected by A.Tovar. |
| Thursday_04_N | TBF1: IAP @warm : Permanent trigger originating from the ext-int V clamp card. | ICS |  | Richard contacted. He called ICS team leader. It was found that the soldering of the Vtaps VT1_5, VT2_5 in the bus bar were not correct. Then OK |
| Thursday_04_N | TBC2; Cryo problem : quench not possible | ACR |  | Problem with cold compressor. |
| Friday_05_A | TBF2: IAP final at warm. LF, HF DAQ not get armed. Reset all the crates. Exit from TEMA. When doing exit: meassage : wait a moment. TEMA blocked. | MTM/AB-CO | TEMA | Richard contacted. He called Magnus who killed the process. A.Raimondo is preparing a tool to be used by equipment support to do it. |
| Friday_05_A | TBE1: End of HV test, quench signal received because of A.Tovar accidentally stopped the Communication PLC/CTH. CFB pressure increased ????. | ACR |  | No explanation for this increase of pressure. Valve was closed ΔT ->11k |
| Friday_05_N | TBC1: Unable to adjust amplitude of D2_U-L below 77 mV (AC comp cold) TBD1: Same for D1 l-U below 100 mV Initially D20 was not connected and 9v at potaim card | MTM | IAP | Advise of Richard. Check if all the connectors are correctly plugged in particular the D20. Revise IAP criteria. |
| Friday_05_N | TBE1:Min energy launched from TEMA . Ramped to 11850 A the ramp down after 20 s then ramp up without intervention. | MTM/ABCO | TEMA | Cycle is OK (checked by Maryline). Slow PA |
| Friday_05_N | TBC1: no current reading in Aperture 1 One cable remains plugged in the SSS crates. | MTM |  | Richard solved the problem. |
| Friday_05_N | TBC1: Quench at 1.5 kA Error in AQA. Triggered by Vbus. | MTM |  | Symmetric quench? |
| Saturday_06_N | TBC1: During LHC cycle, VME error. Cycle interrupted The auto test was good. | AB/CO |  | Richard advises to reset VME crates then repeat a small quench and re-start LHC cycle. He asked to confirm with the test coordinator. Test coordinator decided to stop the magnetic measurements because VME errors happened frequently on C1 and magnet was good. Should we change the crates? |
| Saturday_06_N | TBC1: HV test at cold stopped four times Message read : The megger stopped the measurement. Internal problem of Megger | MTM |  | Perform run-recovery. Next version of HV : Automatic run recovery implemented in that case. After filtering errors. |
| Sunday_07_N | TBF2 (2014 warm): In the traveller is is mentioned that the times when the electrical test done by MTM are not reported then delaying the test sequences. This is not the case. The traveller was filled in time (Kevin) | MTM/ACR/ICS. |  |  |