Vc/6-Nov-04

### Week45 starting 1Nov 04

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday 1 Nov 04  ~15.00 hrs

LHC SM18 Magnet Test Facility

| 7 | A1 | MBBL2098 | Prep 5 Pumping + Cool Down | 299.1 |
| --- | --- | --- | --- | --- |
| 4 | A2 | MBBL3219 | PT1 IAP@1.9K, Part III | 1.9 |
| 3 | B1 | SSS067 | Stretched Wire Cold SC25 | 1.9 |
| 8 | B2 | _ | _ | 298 |
| 12 | C1 | MBBL2115 | ICS 5 Traveller Last | 298 |
| 6 | C2 | MBBL1163 | Prep 5 Pumping + Cool Down | 281.8 |
| 9 | D1 | SSS027 | ICS 3 Traveller after warm test | 296.8 |
| 2 | D2 | SSS065 | Stretched Wire Cold SC25 | 1.91 |
| 10 | E1 | MBAL1094 | ICS 4 Final connection | 297.4 |
| 5 | E2 | MBBL3099 | Prep 5 Pumping + Cool Down | 82.4 |
| 11 | F1 | _ | _ | 297.8 |
| 1 | F2 | MBBR3188 | PT 12 Warm Up | 204.3 |

11/1/2004 2:59:30 PM    Temperature data are taken from TBSM18 site

Coordinator this week: V.Chohan; MTM support on call :  (16-3602 … Richard).

13 Magnets [ 10 dipoles( 3 with MM)+ 3 SSS +  0 SSSS]   Disconnected after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

HV Insulation re-test on SSS after corrector powering ? (observation Richard )

AOB

| From 24-10-2004 to 30-10-2004 |
| --- |
| - | Bench | Magnet Name | shafts | Start Date . | Connect Time | Wait +Cooling | Cold Tests | Wait +Warming | Disconnect Time | End Date . | Total | Q |
| 1 | TBF2 | MBAR3049 | False | 24-10-2004 | 25 h. | 28 h. | 26 h. | 20 h. | 1 h. | 28-10-2004 | 4 days | 1 |
| 2 | TBF1 | MBAL1158 | False | 22-10-2004 | 28 h. | 33 h. | 26 h. | 36 h. | 7 h. | 27-10-2004 | 5 days | 2 |
| 3 | TBE2 | MBAL3232 | False | 18-10-2004 | 30 h. | 63 h. | 25 h. | 19 h. | 8 h. | 24-10-2004 | 6 days | 2 |
| 4 | TBE2 | MBBL1128 | False | 24-10-2004 | 23 h. | 41 h. | 14 h. | 14 h. | 8 h. | 28-10-2004 | 4 days | 2 |
| 5 | TBE1 | MBBL2098 | False | 21-10-2004 | 29 h. | 35 h. | 47 h. | 17 h. | 10 h. | 27-10-2004 | 6 days | 9 |
| 6 | TBD2 | SSS066 | False | 14-10-2004 | 78 h. | 28 h. | 76 h. | 11 h. | 36 h. | 24-10-2004 | 10 days | 2 |
| 7 | TBC2 | MBBL3164 | True | 17-10-2004 | 33 h. | 87 h. | 66 h. | 13 h. | 5 h. | 26-10-2004 | 9 days | 2 |
| 8 | TBC2 | MBBL3236 | True | 26-10-2004 | 28 h. | 23 h. | 30 h. | 14 h. | 5 h. | 30-10-2004 | 4 days | 2 |
| 9 | TBC1 | MBBL1150 | False | 20-10-2004 | 48 h. | 62 h. | 17 h. | 39 h. | 12 h. | 28-10-2004 | 8 days | 2 |
| 10 | TBB2 | SSS064 | False | 13-10-2004 | 52 h. | 79 h. | 90 h. | 18 h. | 10 h. | 25-10-2004 | 12 days | 2 |
| 11 | TBB1 | SSS068 | False | 21-10-2004 | 40 h. | 18 h. | 71 h. | 5 h. | 7 h. | 27-10-2004 | 6 days | 2 |
| 12 | TBA2 | MBAL3230 | True | 23-10-2004 | 29 h. | 31 h. | 41 h. | 43 h. | 10 h. | 29-10-2004 | 6 days | 2 |
| 13 | TBA1 | MBAL3226 | False | 21-10-2004 | 33 h. | 32 h. | 36 h. | 19 h. | 14 h. | 27-10-2004 | 6 days | 5 |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAL3226 | 9548 A. (D1_L) | 10280 A. (D1_U) | 11729 A. (D2 - L) | 12136 A. (D2-L) | 12667 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - |
| 2 | MBAL3230 | 12599 A. (D2_U_S01) | 12486 A. (D1-U) |  | - | - | - | - | - | - | - | - | - | - |
| 3 | SSS068 | 11914 A. (QI-12) | 12160 A. (QI_12) |  | - | - | - | - | - | - | - | - | - | - |
| 4 | SSS064 | 11703 A. (QI-12) | 12349 A. (QE-34) |  | - | - | - | - | - | - | - | - | - | - |
| 5 | MBBL1150 | 12575 A. (D1_U) | 12850 A. (No Data) |  | - | - | - | - | - | - | - | - | - | - |
| 6 | MBBL3164 | 10418 A. (D2_U_S01) | 11824 A. (D1_S01) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 7 | MBBL3236 | 9849 A. (D2-L-S13) | 12208 A. (D1-L-S01) |  | - | - | - | - | - | - | - | - | - | - |
| 8 | SSS066 | 12148 A. (Q_E-12) | 12361 A. (QE_34) |  | - | - | - | - | - | - | - | - | - | - |
| 9 | MBBL2098 | 9818 A. (D2-U) | 10967 A. (D2_L) | 12216 A. (D2_U) | 12285 A. (D2-U) | 12260 A. (D2-U) | 11118 A. (D2_U) | 11072 A. (D2_U) | 11197 A. (D2-U) | 11224 A. (D2-U) |  | - | - | - |
| 10 | MBAL3232 | 10125 A. (D2_U) | 12022 A. (D2_U) |  | - | - | - | - | - | - | - | - | - | - |
| 11 | MBBL1128 | 12108 A. (D2_U) | 12682 A. (D2-U) |  | - | - | - | - | - | - | - | - | - | - |
| 12 | MBAL1158 | 12342 A. (D2-L) | 12606 A. (D2-L) |  | - | - | - | - | - | - | - | - | - | - |
| 13 | MBAR3049 | 12084 A. (D2-L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |

PROBLEMS OF LAST WEEK

| Date | Bench | Problem | Type | Application/soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _26_M | TBA2 | Iearth to high (6-7 mA )recorded during diode test of the MBAL3230. | MTM |  | Pierre advised to check the electrical integrity of the magnet (HV test) |
| Tuesday _26_A | TBF2 | CDAP report shows results not OK : but no step is indicated as “failed”. | AB/CO? | HV | Step(s) not executed in the sequence? |
| Wednesday _27_A | TBD1 | T-coil recordings not correct. | MTM |  | Giorgio informed. Sector 04 BCD broken (not use for MQ magnetic measurements). Giorgio found also a problem in the “nappe” which disappeared. |
| Wednesday _27_A | TBD1 | SPA abort during LHC cycle. No cryogenic problem and not detected by the PLC-SE |  |  | Pierre’s advise: Training quench and repeat LHC cycle. |
| Thursday _28_M | TBD1 | MM3 (full load-line): no Temperature recordings, shaft position alert at high current. MM4: Correctors measurements stopped : Maryline’s investigation on the MM measurements. | MTM |  | Pierre advised to proceed. |
| Thursday _28_A | TBC2 | MM1: not good value for TT821 | MTM |  | Giorgio investigated the problem: pin broken in the chassis: Chassis changed. Then the “0” of the coder not recognized. 1 cable connecting one integrator to the chassis not working +sliding of one aperture (mechanical problem in the TRU). Giorgio solved all the problems but time lost ~4 h |
| Thursday _28_A | TBC2 | PC blocked in “stand-by mode”’ | AB/PO |  | Reset the server and re-start? |
| Thursday _28_N | TBB1 | Vacuum leak? | ICS |  | ICS asked to connect the blue cable for the anti-cryostat. |
| Thursday _28_N | TBD1 | CDAP, HV test: Stop of the megger during the measurements (rack MXT.MY95) |  |  | Cable connection between Megger and the bench not working properly.? Richard informed. |
| Thursday _28_N | TBB1 | Not enough wire for SSW measurements. Nitrogen cylinder empty. Stage B not found. Wire broken. | MTM |  |  |
| Friday_29_M | TBB2 | The sequence Ramp_MQ-760 A and Ramp_MQ_11850 A were missing for B2 bench. | MTM |  | Giorgio add these sequences. |
| Friday_29_M | TBD1 | No corrector signal seen with shaft measurements. | MTM | MMP | Origin current in one aperture and measurement in the second aperture. Giorgio modified the run type: measurement on both apertures and then OK. |
| Friday_29_A | TBD1 | Motor controller of SSW damaged. | MTM |  | Guy informed: The motor controller needs to be switched "ON" after switching "ON" Kepco. After use, we need to switch "OFF" the motor controller before switching "OFF" Kepco. Motor controller has been replaced |
| Friday_29_N | TBB2 | Sextupole measurements at cold not correct, repeated a second time: still outside the window. Cryo OK lost 2 times | MTM ACR |  | Giorgio informed: the wrong circuit was powered. Octupole powered instead of sextupole one. |
| Friday_29_N | TBD1 | During strength measurements wire broke. X-axis result data not found. | MTM |  | New wire and test re-launched. |
| Saturday 30_N | TBB1 | Powering of the EF circuit (octupole) not possible (SSS67) | ICS? AB/PO? |  | Proceed with the rest of the cod test program .AB/PO called monday. Investigation on the run. |
| Sunday_31_N | TBB1 | Wire broke during rot-wire measurements (sextupole) . | TS/SU |  | New wire, XY measurements and rot-wire measurements repeated. |
| Monday_1_M | TBD2 | Survey called. No answer. | TS/SU |  | Survey came 2h30 after. |

# Text boxes

- Light green : Connect (up to Cryo pump down)	24 h

- Blue :  Wait + Cool down				26 h

- Ivory: Tests at Cold					36 h

- Red :  Wait + Warm up					12 h

- Dark Green: Disconnect					12 h