Vc/9-Nov-04

### Week46 starting 8Nov 04

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday 8 Nov 04  ~10.00 hrs

LHC SM18 Magnet Test Facility

| 4 | A1 | MBBL2098 | PT 4 Slow power abort check | 1.9 |
| --- | --- | --- | --- | --- |
| 1 | A2 | MBBL3219 | PT 12 Warm Up | 116 |
| 6 | B1 | SSS042 | Prep 5 Pumping + Cool Down | 295.6 |
| 12 | B2 | SSS069 | ICS 5 Traveller Last | 297.4 |
| 5 | C1 | MBBR3243 | Prep 5 Pumping + Cool Down | 148.6 |
| 9 | C2 | _ | _ | 296.3 |
| 3 | D1 | SSS075 | PT 8 Fast De-excitation | 1.9 |
| 10 | D2 | SSS071 | CR 1 Leak Test | 295.9 |
| 2 | E1 | MBBL2116 | PT 6.2 Training 2 | 5.2 |
| 8 | E2 | MBAL2097 | ICS 3 Traveller after warm test | 295.8 |
| 7 | F1 | MBAL2092 | ICS 4 Final connection | 294.5 |
| 11 | F2 | MBBL3221 | ICS 5 Traveller Last | 296.9 |

11/8/2004 9:50:45 AM    Temperature data are taken from TBSM18 site

Coordinator this week: Gijs De Rijk; MTM support on call :  (16-3602 Maryline).

14 Magnets:

9 dipoles( 0 with MM)+ 5 SSS(1 with MM) +  0 SSSS ]   Disconnected after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

Traveller & SSS & Warm SSW

AOB

| Magnets under Test from 31-10-2004 to 06-11-2004 |
| --- |
| - | Bench | Magnet | Shafts | Start Date . | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total |  |  |
| 1 | TBA1 | MBBL1149 |  | 27-10-2004 | 21 h. | 25 h. | 13 h. | 18 h. | 7 h. | 31-10-2004 | 96 hours |  |  |
| 2 | TBB1 | SSS067 |  | 27-10-2004 | 38 h. | 38 h. | 65 h. | 14 h. | 7 h. | 03-11-2004 | 168 hours |  |  |
| 3 | TBB2 | SSS050 |  | 25-10-2004 | 43 h. | 24 h. | 56 h. | 12 h. | 14 h. | 01-11-2004 | 168 hours |  |  |
| 4 | TBC1 | MBBL2115 |  | 28-10-2004 | 32 h. | 29 h. | 17 h. | 18 h. | 8 h. | 01-11-2004 | 96 hours |  |  |
| 5 | TBC1 | MBBR3222 |  | 01-11-2004 | 24 h. | 43 h. | 17 h. | 15 h. | 5 h. | 06-11-2004 | 120 hours |  |  |
| 6 | TBC2 | MBBL1163 |  | 30-10-2004 | 24 h. | 37 h. | 17 h. | 19 h. | 5 h. | 04-11-2004 | 120 hours |  |  |
| 7 | TBD1 | SSS014 |  | 21-10-2004 | 115 h. | 32 h. | 91 h. | 10 h. | 9 h. | 31-10-2004 | 240 hours |  |  |
| 8 | TBD1 | SSS027 |  | 31-10-2004 | 35 h. | 27 h. | 44 h. | 8 h. | 9 h. | 05-11-2004 | 120 hours |  |  |
| 9 | TBD2 | SSS065 |  | 24-10-2004 | 61 h. | 96 h. | 62 h. | 14 h. | 14 h. | 03-11-2004 | 240 hours |  |  |
| 10 | TBE1 | MBAL1094 |  | 31-10-2004 | 30 h. | 30 h. | 17 h. | 14 h. | 6 h. | 05-11-2004 | 120 hours |  |  |
| 11 | TBE1 | MBAL1157 |  | 27-10-2004 | 26 h. | 40 h. | 20 h. | 13 h. | 10 h. | 31-10-2004 | 96 hours |  |  |
| 12 | TBE2 | MBBL3099 |  | 29-10-2004 | 42 h. | 49 h. | 12 h. | 13 h. | 6 h. | 03-11-2004 | 120 hours |  |  |
| 13 | TBF1 | MBAL1160 |  | 27-10-2004 | 27 h. | 29 h. | 17 h. | 25 h. | 8 h. | 01-11-2004 | 120 hours |  |  |
| 14 | TBF2 | MBBR3188 |  | 28-10-2004 | 25 h. | 44 h. | 16 h. | 17 h. | 10 h. | 01-11-2004 | 96 hours |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBL1149 | 12300 A. (D1-L) | 12706 A. (D1_U) |  | - | - | - | - | - | - | - | - | - | - |
| 2 | SSS067 | 11783 A. (QE-34) | 12848 A. (QI-12) |  | - | - | - | - | - | - | - | - | - | - |
| 3 | SSS050 | 12200 A. (QI-34) | 12732 A. (QE-34) |  | - | - | - | - | - | - | - | - | - | - |
| 4 | MBBL2115 | 10700 A. (D2-U) | 12100 A. (D1-L) |  | - | - | - | - | - | - | - | - | - | - |
| 5 | MBBR3222 | 12057 A. (D2-U) | 12469 A. (D2_U) |  | - | - | - | - | - | - | - | - | - | - |
| 6 | MBBL1163 | 12042 A. (D1_U) | 12541 A. (D2_L) |  | - | - | - | - | - | - | - | - | - | - |
| 7 | SSS014 | 12137 A. (QI-34) | 12314 A. (QE-12) |  | - | - | - | - | - | - | - | - | - | - |
| 8 | SSS027 | 11700 A. (QE_12) | 12498 A. (QE_34) |  | - | - | - | - | - | - | - | - | - | - |
| 9 | SSS065 | 11653 A. (QE-12) | 12553 A. (QI-12) |  | - | - | - | - | - | - | - | - | - | - |
| 10 | MBAL1094 | 12195 A. (D1-U) | 12835 A. (D2-L) |  | - | - | - | - | - | - | - | - | - | - |
| 11 | MBAL1157 | 11724 A. (D1_U) | 12420 A. (D1-U) |  | - | - | - | - | - | - | - | - | - | - |
| 12 | MBBL3099 | 12055 A. (D2_L) | 12112 A. (D2-U) |  | - | - | - | - | - | - | - | - | - | - |
| 13 | MBAL1160 | 11511 A. (No Data) | 12483 A. (D1_L) |  | - | - | - | - | - | - | - | - | - | - |
| 14 | MBBR3188 | 12307 A. (D1_U) | 12383 A. (D1-U) |  | - | - | - | - | - | - | - | - | - | - |

| From 31-10-2004 to 06-11-2004 |
| --- |
| - | Bench | Magnet Name | shafts | Start Date . | Connect Time | Wait +Cooling | Cold Tests | Wait +Warming | Disconnect Time | End Date . | Total | Q |
| 1 | TBF2 | MBBR3188 | False | 28-10-2004 | 25 h. | 44 h. | 16 h. | 17 h. | 10 h. | 01-11-2004 | 4 days | 2 |
| 2 | TBF1 | MBAL1160 | False | 27-10-2004 | 27 h. | 29 h. | 17 h. | 25 h. | 8 h. | 01-11-2004 | 5 days | 2 |
| 3 | TBE2 | MBBL3099 | False | 29-10-2004 | 42 h. | 49 h. | 12 h. | 13 h. | 6 h. | 03-11-2004 | 5 days | 2 |
| 4 | TBE1 | MBAL1094 | False | 31-10-2004 | 30 h. | 30 h. | 17 h. | 14 h. | 6 h. | 05-11-2004 | 5 days | 2 |
| 5 | TBE1 | MBAL1157 | False | 27-10-2004 | 26 h. | 40 h. | 20 h. | 13 h. | 10 h. | 31-10-2004 | 4 days | 2 |
| 6 | TBD2 | SSS065 | False | 24-10-2004 | 61 h. | 96 h. | 62 h. | 14 h. | 14 h. | 03-11-2004 | 10 days | 2 |
| 7 | TBD1 | SSS014 | True | 21-10-2004 | 115 h. | 32 h. | 91 h. | 10 h. | 9 h. | 31-10-2004 | 10 days | 2 |
| 8 | TBD1 | SSS027 | False | 31-10-2004 | 35 h. | 27 h. | 44 h. | 8 h. | 9 h. | 05-11-2004 | 5 days | 2 |
| 9 | TBC2 | MBBL1163 | False | 30-10-2004 | 24 h. | 37 h. | 17 h. | 19 h. | 5 h. | 04-11-2004 | 5 days | 2 |
| 10 | TBC1 | MBBL2115 | False | 28-10-2004 | 32 h. | 29 h. | 17 h. | 18 h. | 8 h. | 01-11-2004 | 4 days | 2 |
| 11 | TBC1 | MBBR3222 | False | 01-11-2004 | 24 h. | 43 h. | 17 h. | 15 h. | 5 h. | 06-11-2004 | 5 days | 2 |
| 12 | TBB2 | SSS050 | False | 25-10-2004 | 43 h. | 24 h. | 56 h. | 12 h. | 14 h. | 01-11-2004 | 7 days | 2 |
| 13 | TBB1 | SSS067 | False | 27-10-2004 | 38 h. | 38 h. | 65 h. | 14 h. | 7 h. | 03-11-2004 | 7 days | 2 |
| 14 | TBA1 | MBBL1149 | False | 27-10-2004 | 21 h. | 25 h. | 13 h. | 18 h. | 7 h. | 31-10-2004 | 4 days | 2 |

from 31-10-2004 to 06-11-2004

PROBLEMS OF LAST WEEK

Problem chronology  of the SSW Measurements done at Cold for SSS 027 on TBD1.

Test Done                                                    No of times Repeated                          Date and time                                               Status

SC21:Installation of SSW in AP 2                                   1                                                    3/11/04         9:30 PM                                   OK

FS21: Field Strength Measurement in AP 2 @ 760 A       4                                                   4/11/04         00.20 AM                            Failed

FS22: Field Strength Measurement in AP 2 @ 5000 A     3                                                   4/11/04         00:53 AM                            Failed

FS23:Field Strength Measurement in AP 2 @ 5000 A       2                                                   4/11/04         3:02 AM                              Failed

SC22:Roll angle Measurement in AP2 - Quadrapole          1                                                   4/11/04         5:29 AM                              Failed
                                                                                                                                                                                         (Wire broke and it is reinstalled)

SC21:Installation of SSW in AP 2                                   1                                                   4/11/04         3:30 PM                               OK

SC22:Roll angle Measurement in AP2 - Quadrapole         7                                                    4/11/04         3:40 PM                               Failed
                                                                                                                                                                                                     (Uncertainty very high)

SC11:Installation of SSW in AP 1                                1                                                    4/11/04         7:00 PM                              OK

SC12:Roll angle Measurement in AP1 - Quadrapole        1                                                     4/11/04         7:10 PM                              Failed
                                                                                                                                                                                            (Readings out of range)

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _02_N | TBD2 | CDW2 on SSS065: resistance values of sextupoles are out of ranges. Already observed before cold tests | MTM |  | Stripping not given before investigation. |
| Tuesday _02_N | TBB1 | IAP at warm final: results in bothMQAX-R67 and MQXX-R67. | AB/CO | IAP | Results merged in the proper folder. |
| Wednesday _03_A | TBA2 | LHC cycles on 3219 magnet. Angles between forward and downward rotation too high in the current ramp. | MTM | MMP | Stephane contacted. The values measured during the current ramp are not correct We should consider only those on current plateau. After check he advised to proceed. |
| Wednesday_03_N | TBD1 | SSW measurements on SSS065. Motor blocked. Bad results obtained for strength measurements at cold. Wire broken. | MTM |  | Test stopped and re-lauunched. Check by N.Smirnov. |
| Wednesday_03_N |  | Rocla blocked in front of the C bench |  |  | 9 hours of stop. |
| Wednesday_03_N Sunday_07_A | TBF1 TBF2 | Wrong HV test sequence performed at cold : The sequence of HV at warm was applied. | MTM |  | Error found Test with sequence at cold re-launched. |
| Wednesday_03_N | TBA2 | Problem in the scan of the long anti-cryostat occurring in the middle of the cold test (after 6 LHC cycles)? | MTM |  | Anti-cryostat coefficients recognized now for the rest of the test. Maryline will trace back the problem. |
| Thursday_04_M |  | PT4: Communication problem between PLC and PC | AB/CO? |  |  |
| Thursday_04_M |  | PT5.1: Data not visible in TEMA .Analysis done in local. | AB/CO? |  |  |
| Thursday_04_M | TBF1 | Connector of the cable connecting the TRU and the mobile rack is found damaged | MTM |  |  |
| Thursday_04_M | TBD1 | SPA occurred during roll angle measurements. Roll angle values and uncertainty out of the limits | MTM |  |  |
| Thursday_04_A | TBD2 | HV test failed on SSS071 Resistance value out of range for Ext.dipole and MQT. | CRI | HV | Richard called. Investigation 2 non conformities opened. -VTaps inverted in the MQT: proceed as is. -Inversion of the Vtaps of the dipole performed by Richard. |
| Thursday_04_A | TBD2 | SSW at cold giving bad results. Good at warm |  |  | Nicolay, Guy contacted. Software problem. J.Di Marco contacted. Still unsolved |
| Thursday_04_N | TBD1 | HV test launched with XMT.MY95 rack. Communication problem. | AB/CO? | HV | Rack changed. OK. Richard informed. |
| Friday_05_M | TBF1 | LHC cycle not inserted during the training. | MTM |  | TC informed. Provoked quench at 10 kA and test done. |
| Friday_05_M | TBA2 | SPA during LHC cycle at 11850 A | ? |  |  |
| Friday_05_M | TBF1 | SPA during LHC cycle: DCCT falls down. | AB/PO |  | Richard put the DCCT on. Quench performed and LHC done |
| Friday_05_A | TBF1 | Fault on the motor during LHC cycle. | MTM |  | Cable rolled around the motor axis. Richard and Guy solved the problem. |
| Friday_05_A | TBB2 | TEMA not responding. Sun crashed. | AB/CO |  | Richard investigated. Power supply of the sun out of work. Richard phoned to Magnus and Alessandro. Richard took the power supply of the SUN in the cluster G and replaced the dead one in that of the cluster B. |
| Friday_05_N | TBB2 | Permanent trigger coming from the chassis 4 of the Potaim-cards. | MTM |  | Richard called. He found a problem with one connection cable. He decide to un-validate the concerned card. He will see P.Legrand Monday to check the cable. |
| Saturday_06_M | TBD1 | Lyra test: not a good value of Vmagnet at warm. (2.68 V instead of 3.54 V) | ICS |  | Richard suspected a problem in the VTaps soldered by ICS. He put a strapping in the chassis to by pass these to taps. Then OK for Lyra test. |
| Saturday_06_A | TBC1 | One Vtaps cut in the CFB | MTM |  | Richard informed and by pass it in the chassis. Pot-Aim. |
| Sunday 07_A | TBD2 | SSS071: leak test failed. |  |  | Problem coming from the line X? leak on the MRB? TC informed. Disconnection of the magnet decide Monday.. |
| Sunday 07_A | TBF2 | CDAP HV : Warm sequence followed (voltage up to 600 V). | MTM | HV | Test repeated with the cold sequence. |
| Sunday 07_A | TBB1 | IAP at warm: resistance measurements not correct. | MTM | IAP | Richard called. SSS type B-F. Adaptor needed plugged on the connector P11 to redistribute the V taps. |