Vc/16-Nov-04

### Week47 starting 15Nov 04

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday 15 Nov 04  ~12.20 hrs

LHC SM18 Magnet Test Facility

| 10 | A1 | MBBL3220 | ICS 4 Final connection | 295.2 |
| --- | --- | --- | --- | --- |
| 3 | A2 | MBBL2112 | PT 2 HV Cold Test | 1.9 |
| 5 | B1 | SSS042 | PT 10.1: Prov. Quench 1/2 HF 11850 | 11.6 |
| 6 | B2 | MBBR3241 | PT 12 Warm Up | 267.4 |
| 2 | C1 | MBAR1138 | PT 6.1 Training 1 | 1.9 |
| 8 | C2 | MBBL2119 | Prep 5 Pumping + Cool Down | 205.4 |
| 9 | D1 | SSS039 | ICS 4 Final connection | 294.5 |
| 1 | D2 | SSS071 | Survey @ Cold SC14 | 1.91 |
| 12 | E1 | MBBL1164 | Finished Ready to go Final scan | 298 |
| 4 | E2 | MBAL2097 | PTC 9 HV @ Cold LAST | 1.9 |
| 7 | F1 | MBBL1147 | Prep 5 Pumping + Cool Down | 2.35 |
| 11 | F2 | MBBR3244 | ICS 4 Final connection | 293.8 |

11/15/2004 12:21:40 PM    Temperature data are taken from TBSM18 site

Coordinator this week: Vittorio; MTM support on call :   16-3602 .

12 named  Magnets Tested or 13 Magnets tests :

10 dipoles( 2 with MM & long MM)+ 2 SSS(1 with MM) +  0  SSSS     Disconnected after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

SSW system & Guy

AOB

## Tests

| Magnets under Test from 07-11-2004 to 13-11-2004 |
| --- |
| - | Bench | Magnet | Shafts | Start Date . | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total |  |  |
| 1 | TBA1 | MBBL2098 |  | 31-10-2004 | 28 h. | 167 h. | 32 h. | 15 h. | 7 h. | 10-11-2004 | 240 hours |  |  |
| 2 | TBA2 | MBBL3219 |  | 29-10-2004 | 24 h. | 44 h. | 160 h. | 18 h. | 0 h. | 09-11-2004 | 264 hours |  |  |
| 3 | TBA2 | MBBL3219 |  | 09-11-2004 | 7 h. | 51 h. | 12 h. | 19 h. | 4 h. | 11-11-2004 | 48 hours |  |  |
| 4 | TBB2 | SSS069 |  | 01-11-2004 | 57 h. | 31 h. | 51 h. | 13 h. | 13 h. | 08-11-2004 | 168 hours |  |  |
| 5 | TBC1 | MBBR3243 |  | 06-11-2004 | 29 h. | 28 h. | 15 h. | 23 h. | 9 h. | 11-11-2004 | 120 hours |  |  |
| 6 | TBC2 | MBBL2113 |  | 04-11-2004 | 41 h. | 20 h. | 15 h. | 13 h. | 6 h. | 08-11-2004 | 96 hours |  |  |
| 7 | TBD1 | SSS075 |  | 05-11-2004 | 18 h. | 32 h. | 76 h. | 19 h. | 13 h. | 12-11-2004 | 168 hours |  |  |
| 8 | TBE1 | MBBL2116 |  | 05-11-2004 | 15 h. | 36 h. | 18 h. | 20 h. | 7 h. | 09-11-2004 | 96 hours |  |  |
| 9 | TBE2 | MBAL1161 |  | 03-11-2004 | 28 h. | 23 h. | 15 h. | 16 h. | 10 h. | 07-11-2004 | 96 hours |  |  |
| 10 | TBE2 | MBAL2097 |  | 07-11-2004 | 27 h. | 27 h. | 51 h. | 20 h. | 0 h. | 12-11-2004 | 120 hours |  |  |
| 11 | TBF1 | MBAL2092 |  | 07-11-2004 | 40 h. | 7 h. | 16 h. | 19 h. | 14 h. | 11-11-2004 | 96 hours |  |  |
| 12 | TBF1 | MBBL3218 |  | 01-11-2004 | 27 h. | 27 h. | 57 h. | 15 h. | 8 h. | 07-11-2004 | 144 hours |  |  |
| 13 | TBF2 | MBBL3221 |  | 02-11-2004 | 24 h. | 69 h. | 29 h. | 15 h. | 5 h. | 08-11-2004 | 144 hours |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBL2098 | 11778 A. (D2_U) | 12606 A. (A2-CF-S3) | 10753 A. (D2-U) | 11081 A. (D2-U) | 10 A. (No Data) | 10966 A. (D2_U) | 11054 A. (A2 CF S3) | 10980 A. (D2_U) |  | - | - | - | - |
| 2 | MBBL3219 | 10660 A. (D2-L-S01) | 10417 A. (D1-U-S01) | 11741 A. (D2_L_S01) | 12460 A. (D2_U_S01) | 12779 A. (D@_U_S01) | 12365 A. (D2-U-?) | 12345 A. (D2-U) | 12654 A. (D2 U) | 12740 A. (D2-U) |  | - | - | - |
| 3 | MBBL3219 | 10791 A. (D1-L) | 11631 A. (D2_L_MR) | 12696 A. (D2_L) |  | - | - | - | - | - | - | - | - | - |
| 4 | SSS069 | 12250 A. (QE_12) | 12248 A. (QI_34) |  | - | - | - | - | - | - | - | - | - | - |
| 5 | MBBR3243 | 12369 A. (D2_U) | 12849 A. | - | - | - | - | - | - | - | - | - | - | - |
| 6 | MBBL2113 | 11916 A. (D2-L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 7 | SSS075 | 12248 A. (QI_12) | 12649.6 A. (QE-34) |  | - | - | - | - | - | - | - | - | - | - |
| 8 | MBBL2116 | 10789 A. (D1-L) | 12435 A. (D1-U) |  | - | - | - | - | - | - | - | - | - | - |
| 9 | MBAL1161 | 12030 A. (D1-U) | 12327 A. (D2_L) |  | - | - | - | - | - | - | - | - | - | - |
| 10 | MBAL2097 | 11903 A. (D2-U) | 12289 A. (D1-U) | 12532 A. (D1_L) | 12857 A. | - | - | - | - | - | - | - | - | - |
| 11 | MBAL2092 | 11410 A. (D1-U-MRB) | 12708 A. (D2-L MRB) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 12 | MBBL3218 | 10289 A. (No Data) | 10402 A. (D2-L-S01) | 12266 A. (D2 L) | 12491 A. (D2-L) | 12850 A. | - | - | - | - | - | - | - | - |
| 13 | MBBL3221 | 9207 A. (D1-L) | 10855 A. (D2-L) | 11525 A. (D1-U) | 12773 A. (D1-L) | 12523 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - |

| From 07-11-2004 to 13-11-2004 |
| --- |
| - | Bench | Magnet Name | shafts | Start Date . | Connect Time | Wait +Cooling | Cold Tests | Wait +Warming | Disconnect Time | End Date . | Total | Q |
| 1 | TBF2 | MBBL3221 | False | 02-11-2004 | 24 h. | 69 h. | 29 h. | 15 h. | 5 h. | 08-11-2004 | 6 days | 5 |
| 2 | TBF1 | MBAL2092 | False | 07-11-2004 | 40 h. | 7 h. | 16 h. | 19 h. | 14 h. | 11-11-2004 | 4 days | 2 |
| 3 | TBF1 | MBBL3218 | True | 01-11-2004 | 27 h. | 27 h. | 57 h. | 15 h. | 8 h. | 07-11-2004 | 6 days | 4 |
| 4 | TBE2 | MBAL1161 | False | 03-11-2004 | 28 h. | 23 h. | 15 h. | 16 h. | 10 h. | 07-11-2004 | 4 days | 2 |
| 5 | TBE2 | MBAL2097 | False | 07-11-2004 | 27 h. | 27 h. | 51 h. | 20 h. | 0 h. | 12-11-2004 | 5 days | 3 |
| 6 | TBE1 | MBBL2116 | False | 05-11-2004 | 15 h. | 36 h. | 18 h. | 20 h. | 7 h. | 09-11-2004 | 4 days | 2 |
| 7 | TBD1 | SSS075 | True | 05-11-2004 | 18 h. | 32 h. | 76 h. | 19 h. | 13 h. | 12-11-2004 | 7 days | 2 |
| 8 | TBC2 | MBBL2113 | False | 04-11-2004 | 41 h. | 20 h. | 15 h. | 13 h. | 6 h. | 08-11-2004 | 4 days | 1 |
| 9 | TBC1 | MBBR3243 | False | 06-11-2004 | 29 h. | 28 h. | 15 h. | 23 h. | 9 h. | 11-11-2004 | 5 days | 1 |
| 10 | TBB2 | SSS069 | False | 01-11-2004 | 57 h. | 31 h. | 51 h. | 13 h. | 13 h. | 08-11-2004 | 7 days | 2 |
| 11 | TBA2 | MBBL3219 | True | 29-10-2004 | 24 h. | 44 h. | 160 h. | 18 h. | 0 h. | 09-11-2004 | 11 days | 9 |
| 12 | TBA2 | MBBL3219 | True | 09-11-2004 | 7 h. | 51 h. | 12 h. | 19 h. | 4 h. | 11-11-2004 | 2 days | 3 |
| 13 | TBA1 | MBBL2098 | False | 31-10-2004 | 28 h. | 167 h. | 32 h. | 15 h. | 7 h. | 10-11-2004 | 10 days | 8 |

from 07-11-2004 to 13-11-2004

| Date | Bench | Problem | Type | Appl. | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _09_N | TBE1 | Change of type for the magnet 1164 from type L to type R. Name not change in the bare code. Scan not correct | MTM |  | Magnet name entered manually. |
| Tuesday _09_N | TBD2 | Wire broken during rot-wire measurements | MTM |  | XY axis + test repeated. |
| Wednesday _10_M | TBA1 | Server crashed while using TEMA Disk full. | AB/CO |  | TEMA re-launched. |
| Wednesday _10_M | TBD1 | LF DAQ not working. SPA occurred. | AB/CO | LF |  |
| Wednesday _10_N | TBE2 | 1.5 kA provoked quench (all HF). Alert jumpv>1.0. | MTM | AQA | Maryline, TC contacted. They decided to repeat Resistance measurements and HV test. OK. Proceed with the first training |
| Wednesday _10_N | TBE2 | TEMA blocked because of HWREC | AB/CO | HWREC | Maryline called. She gave the password to put HWREC off. |
| Wednesday _10_N | TBB1 | Powering EF circuit not possible. Inductance not corresponding to 16 mH | AB/PO |  | Test went one with training. Maryline changed the inductance values. |
| Thursday _11_M |  | Net work problem. AFS server is not working. | IT |  | Wait for re-start of AFS server. |
| Thursday _11_M | TBE2 | First quench at 12000 A. Oscillations in the direct signals. | MTM |  | Stephane and Gijs contacted Pierre. He asked for another quench before investigation. Training pursued up to 9 T + thermal cycle. |
| Thursday _11_N Friday_12_N | TBC2 | During charge of QH, QH1&2 trigger and discharge | MTM |  | Maryline informed. She will check the inter-out card. |
| Thursday _11_N | TBD | The power supply 18 V is missing | MTM |  | Borrow that from TBB. |
| Friday_12_M | TBF | HV test execution slow. | AB/CO | HV | Alessandro found the raison: Bug due to a too big log file. Corrected also in MMP |
| Friday_12_N | TBB1 | MQT powering (EF). Ramp to 550 A but display of power showed wrong inductance value (MO). | AB/CO |  | Wrong name in the power supply controller. Action Maryline to trigger the changes. |
| Friday_12_N | TBB1 | SSW measurement: Difficulty to fix the new wire: the washers on the tension motor spindle did not grip the wire tight enough. Slip of the wire when the tension increases | MTM |  | Action Guy Monday morning. |
| Saturday_13_M | TBA1 | Cryo loss on MB1134 during 1,5 kA (HF) provoked quench | ACR |  |  |
| Saturday_13_M | TBF2 | Current reading in Power Application in TEMA not working during ramping. | MTM | Power |  |
| Saturday_13_M | TBB1 | Sun workstation (menu selection) not working correctly. | AB/CO |  | Adrian called. Disk full. He found some place. |
| Saturday_13_M | TBF2 | Bench selections rotary switch not working. | AB/PO |  | .Maryline contacted. IDCCT crashed. Problem solved by putting on the DCCT. |
| Saturday_13_A | TBD2 | PLC-SE blocked. SSW measurement: Execution blocked when SSW window appeared. The same in the system installed on TBB1 |  |  | Maryline informed. Config file corrupted. Martin copied correct file from another measurements. |
| Sunday_14_M | TBD2 | Powering EF (MQT) circuit not possible (SSS071). 2 triggers are coming. | MTM |  | Powered with 2 A/s instead of 5 A/s but with 16 m H of inductance. Maryline contacted. |
| Sunday_14_N | TBE2 | Lf application did not create correct config.file Problem with extension of training quench. Extension has to be changed from .a to .u | AB/CO | lf | Maryline called. She advised to copy correct one from another test file. |
| Sunday_14_N | TBB1 | SSS042 did not reach 12850 A after 9 ramps. | MTM |  | TC decided to do 4.4 K quench before thermal cycle. Quench in the clamp at 6 kA. SS limit study at 4.4 K on SSS not possible because main leads not in He liquid. |
| Monday_15_A | TBC1 | MB1138: Anomalous decrease of I magnet. | MTM |  | Origin not known. Pierre informed. |
| Monday_15_A | TBC1 | Difficulties to recover the cryo OK at 1.9 K | ACR |  | Oil leak in the pumping group. |