### Week 34

### SM18 Operation SCHEDULE Meeting (Tuesdays)

Bench status  and tests for this week :

| TBA1 | MBBL3192 | ICS 5 Traveller Last | 300.8 |
| --- | --- | --- | --- |
| TBA2 | MBBL1107 | Prep 5 Pumping + Cool Down | 1.9 |
| TBB1 | SSS041 | ICS 5 Traveller Last | 300.3 |
| TBB2 | MBBL1121 | PT 12 Warm Up | 152.9 |
| TBC1 | MBAL3079 | PT 1 IAP @ cold | 1.91 |
| TBC2 | MBBL3198 | Prep 5 Pumping + Cool Down | 301.3 |
| TBD1 | SSS007 | MM3 Full Load Line + Joints | 1.9 |
| TBD2 | SSS012 | Stretched Wire Warm SW24 survey | 288.7 |
| TBE1 | MBBL2075 | PT 12 Warm Up | 276.3 |
| TBE2 | MBBR1019 | Prep 5 Pumping + Cool Down | 302.1 |
| TBF1 | MBBL3136 | CR 1 Leak Test | 223.4 |
| TBF2 | MBBL1131 | Prep 5 Pumping + Cool Down | 104.8 |

Test Coordinator this week:  M.Buzio;  MTM  Support. Maryline:  (16-3602) .

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Problem with CFU: cooling down to 80 K was slow Sunday. One parameter changed in the process.

- SSS

  - SSS: Survey called Sunday morning : not responding all the day. Finally could be reached Sunday night. One CERN number now 165010.

  - Survey on SSS41: delay of one day to perform the survey in aperture 2.

- AOB

  - DAQ for Training quench : Pre-trigger time increased to 154 s: saving DAQ file will take longer.

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Wednesday_11_M | TBC2 | IAP @ warm : interlock present in pot-aim cards | ICS |  | Guy informed. He asked ICS to check the connection. A wire was broken on CFB side. Repaired by ICS |
| Thursday_12_A | TBB1 | PLC-Art communication problem. | AB/CO |  | Put off and on. Then OK |
| Thursday _12_N | TBE1 | MM1: VME error. Remains after resetting the crates. | AB/CO |  | Guy informed. |
| Friday_13_M | TBB2 | Tcoil measurement: wrong value of T on slots 4 and 8 aperture 2 | MTM |  | Guy changed the faulty cards. |
| Friday_13_M | TBB1 | SSS41: SPA: bench configured in dipole for the inductance | MTM |  | Giorgio changed for SSS. |
| Friday_13_M | TBB2 | Permanent trigger on LF: | MTM |  | Giorgio found that card Crate1 slot 16 was faulty. |
| Friday_13_N Saturday_14_M | TBE1 | Saving of training quench files is longer than usual. 2074: AQA warning ; Tau jumps ~9 V VT exceeds 100 V (D2_L~109 V) | MTM |  | Reason: pre-trigger time increased from 9s to 157 s with respect to previous config file. It will take now longer to receive the HF files. Giorgio informed.He asked to check QH discharge and MIITS. |
| Friday_13_N | TBD1 | All 4 heaters should fire after trigger. | MTM |  | Giorgio consulted. |
| Friday_13_N | TBD1 | LHC cycle not OK. DAQ trigger not received Motor touching the end stops Motor 2 touching the end stops Over-range in many sectors | MTM |  | Wrong setting was selected in MMP (dipole instead of quadrupole). GHH restart the measurements. |
| Saturday_14_A | TBE2 | IAP @ warm on 1019 magnet : Temperature sensor reading not OK (173 K instead of 300 K). |  |  | Giorgio consulted. He checked the resistance: not OK 113 ohms instead of 60 ohms. NC opened. |
| Saturday_15_A | TBD1 | 600 A power supply cannot be controlled automatically. Piloted only through a terminals Then coil A in sector 0001 in ap2 not working : | AB/PO |  | Problem in gateway. Changed in E-coil. |