Vc/5 april 05

### Week 14  starting 4 april 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  4 April 05 @ 4/4/2005 2:44:02 PM

| 7 | A1 | MBBL1193 | PT 10 Min. NRJ 11850 A | 2.35 HV |
| --- | --- | --- | --- | --- |
| 6 | A2 | MBBR3307 | Prep 0 Scan Bench | 295.89 |
| 12 | B1 | SSS104 | Prep 3 CDW1 HV | 294.28 |
| 11 | B2 | MBBL1211 | PT 12 Warm Up | 61.61 |
| 4 | C1 | MBBL1212 | Prep 5 Pumping + Cool Down | 106.03 |
| 3 | C2 | MBAR3254 | PT 12 Warm Up | 26.27 |
| 2 | D1 | SSS117 | heater delay - 1/2 HF | 1.92 |
| 8 | D2 | SSS022 | PT 3 QH Measure | 2.35 |
| 5 | E1 | MBBR3315 | Prep 2.1 IAP @ warm LF | 296.52 |
| 9 | E2 | MBAL1210 | Prep 5 Pumping + Cool Down | 1.90 |
| 10 | F1 | MBAL2136 | Prep 5 Pumping + Cool Down | 201.39 |
| 1 | F2 | MBBR3267 | PT 12 Warm Up | 286.25 |

Temperature data are taken from TBSM18 site

Coordinator this week: Pierre; MTM support on call :   16-3602 .

14 Magnets Tested  ( not counting 1214)

11  dipoles( 1 with MM)+ 3  SSS(0  with MM) +  0  SSSS (with MM)

Disconnected after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Situation on Repair :

General Points

…

New AQA

> > 
> > 
> > -------- Original Message -------- Subject:  Improvements in 
> > software and times for operation friendly systems 
> > Date:  Tue, 05 Apr 2005 13:40:35 +0200 
> > From:  chohan <vinod.chohan@cern.ch> 
> > To:  Louis Walckiers <Louis.Walckiers@cern.ch> 
> > 
> > 
> > Quelques modification seront envisageable si nous avons plus de 
> > flexibilite sur les application SM18 : 
> > 
> > *Pour TEMA : 
> > -Le message ("Are you sure that the magnet is not a repeated 
> > magnet") arrivant sur le TEMA lors du lancement de chaque test. 
> > Dois etre place que pour les "IAP all HF" et "IAP at cold" dans le 
> > cas d'un cycle thermique de l'aimant. 
> > -Un message lors du changement de banc actif entre SSS a Dipole 
> > devrait avertir les operateurs a verifier "quench heater connetion" 
> > et "Henry's setting power converter", sur cluster B seulement. 
> > -La possibilite de switcher OFF Hardware Recognition. 
> > 
> > 
> > *Pour AQA v.6 : 
> > -La selection du pens en analyse manuel devrait etre automatique sur 
> > l'application HFoff . 
> > -La selection de "scale global" et les curseurs "free" par defaut a 
> > l'ouverture de l'application. 
> > -Une valeur affiche dans le MIITS manuel pour le delta T entre 
> > l'origine du quench et la detection potaim card (200 mv). Demander 
> > depuis debut decembre. 
> > -Certaine valeurs devraient etre plus lisible (MIITS manuel). 
> > 
> > *Pour IAP sur SSS : 
> > -La lecture des "voltage taps" devrait se faire a partir d'une 
> > application IAPSSS, et non part lfoff, ce qui a chaque fois 
> > complique le travail pour les operateurs lorsque l'on passe de 
> > dipole a SSS 
> > 
> >

AOB

Back to Homepage

| Magnets under Test from 27-03-2005 to 02-04-2005 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total |  |  |
| 1 | TBA1 | MBBL1196 |  |  | 21-03-2005 | 34 h. | 31 h. | 85 h. | 16 h. | 17 h. | 29-03-2005 | 183 h. |  |  |
| 2 | TBB1 | SSS125 |  |  | 25-03-2005 | 33 h. | 23 h. | 19 h. | 10 h. | 22 h. | 30-03-2005 | 107 h. |  |  |
| 3 | TBB2 | MBBL2157 |  |  | 26-03-2005 | 34 h. | 30 h. | 28 h. | 17 h. | 12 h. | 31-03-2005 | 121 h. |  |  |
| 4 | TBC1 | MBAL1200 |  |  | 28-03-2005 | 18 h. | 24 h. | 22 h. | 14 h. | 8 h. | 01-04-2005 | 86 h. |  |  |
| 5 | TBC1 | MBAL2155 |  |  | 21-03-2005 | 31 h. | 30 h. | 26 h. | 14 h. | 34 h. | 28-03-2005 | 135 h. |  |  |
| 6 | TBC1 | MBBL1214 |  |  | 01-04-2005 | 25 h. | -25 h. | 27 h. | 1 h. | 4 h. | 02-04-2005 | 32 h. |  |  |
| 7 | TBC2 | MBAL2125 |  |  | 25-03-2005 | 32 h. | 29 h. | 14 h. | 14 h. | 17 h. | 30-03-2005 | 106 h. |  |  |
| 8 | TBD1 | SSS028 |  |  | 23-03-2005 | 22 h. | 20 h. | 17 h. | 9 h. | 77 h. | 29-03-2005 | 145 h. |  |  |
| 9 | TBD2 | SSS025 |  |  | 25-03-2005 | 40 h. | 14 h. | - | - | - | 31-03-2005 | 118 h. |  |  |
| 10 | TBE1 | MBBR3294 |  |  | 19-03-2005 | 26 h. | 97 h. | 20 h. | 11 h. | 35 h. | 29-03-2005 | 189 h. |  |  |
| 11 | TBE2 | MBAL3080 |  |  | 29-03-2005 | 26 h. | -19 h. | 58 h. | 13 h. | 8 h. | 02-04-2005 | 86 h. |  |  |
| 12 | TBE2 | MBBR3312 |  |  | 22-03-2005 | 27 h. | 32 h. | 31 h. | 15 h. | 15 h. | 29-03-2005 | 120 h. |  |  |
| 13 | TBF1 | MBAR3299 |  |  | 24-03-2005 | 37 h. | 22 h. | 30 h. | 15 h. | 12 h. | 29-03-2005 | 116 h. |  |  |
| 14 | TBF1 | MBAR3304 |  |  | 29-03-2005 | 27 h. | 30 h. | 14 h. | 13 h. | 10 h. | 02-04-2005 | 94 h. |  |  |
| 15 | TBF2 | MBBL2162 |  |  | 25-03-2005 | 50 h. | 24 h. | 17 h. | 15 h. | 17 h. | 30-03-2005 | 123 h. |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBL1196 | 12389 A. (D1_L) | 12382 A. (D1_L_S04) | 12263 A. (D1_L) | 12499 A. (D1_U_MRB) | 12708 A. (D1_L_MRB) | 0 A. (No Data) | 12772 A. (D1-L) | 12846 A. (D1_L) | 12850 A. | - | - | - | - |
| 2 | SSS125 | 11814 A. (QI-34) | 12285.9 A. (QE-12) |  | - | - | - | - | - | - | - | - | - | - |
| 3 | MBBL2157 | 11573 A. (D1_U) | 11959 A. (D2_U) | 12462.3 A. (D2 - U) | 12758.7 A. (D1-L) |  | - | - | - | - | - | - | - | - |
| 4 | MBAL1200 | 11989.3 A. (D1-U) | 12558 A. (D2_U) |  | - | - | - | - | - | - | - | - | - | - |
| 5 | MBAL2155 | 11255 A. (D1-L) | 12675 A. (D2-L) |  | - | - | - | - | - | - | - | - | - | - |
| 6 | MBBL1214 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 7 | MBAL2125 | 11913 A. (D2_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 8 | SSS028 | 11634 A. (QI - 12) | 12227 A. (QE - 34) |  | - | - | - | - | - | - | - | - | - | - |
| 9 | SSS025 | 11702.8 A. (QI_34) | 11759 A. (QE_12) | 12202 A. (QE-12) | 12397 A. (QE-12) | 12550 A. (QI_34) | 12643 A. (QE_34) | 12773.1 A. (QI_12) | 12850 A. | - | - | - | - | - |
| 10 | MBBR3294 | 11917.4 A. (D2_L) | 12302.4 A. (D1_L) |  | - | - | - | - | - | - | - | - | - | - |
| 11 | MBAL3080 | 10540.7 A. (D2_L) | 12197 A. (D1_L) |  | - | - | - | - | - | - | - | - | - | - |
| 12 | MBBR3312 | 8477 A. (D2 - L) | 12164 A. (D2 - U) | 12441.5 A. (D2-U) |  | - | - | - | - | - | - | - | - | - |
| 13 | MBAR3299 | 10237 A. (D2_U) | 11116 A. (D1_U) | 12159 A. (D2_L) | 12690 A. (D2_U) | 12858 A. | - | - | - | - | - | - | - | - |
| 14 | MBAR3304 | 12281 A. (D1-U) | 12581 A. (D2-U) |  | - | - | - | - | - | - | - | - | - | - |
| 15 | MBBL2162 | 11222 A. (D2_U) | 12465.5 A. (D1-L) |  | - | - | - | - | - | - | - | - | - | - |

### Week 13 march 27  to 2 April  05

### [ not count  1214   ]

| From 27-03-2005 to 02-04-2005 |
| --- |
| - | Bench | Magnet Name | shafts | Start Date . | Connect Time | Wait +Cooling | Cold Tests | Wait +Warming | Disconnect Time | End Date . | Total | Q |
| 1 | TBF2 | MBBL2162 | False | 25-03-2005 | 50 h. | 24 h. | 17 h. | 15 h. | 17 h. | 30-03-2005 | 123 h. | 2 |
| 2 | TBF1 | MBAR3299 | False | 24-03-2005 | 37 h. | 22 h. | 30 h. | 15 h. | 12 h. | 29-03-2005 | 116 h. | 4 |
| 3 | TBF1 | MBAR3304 | False | 29-03-2005 | 27 h. | 30 h. | 14 h. | 13 h. | 10 h. | 02-04-2005 | 94 h. | 2 |
| 4 | TBE2 | MBAL3080 | False | 29-03-2005 | 26 h. | -19 h. | 58 h. | 13 h. | 8 h. | 02-04-2005 | 86 h. | 2 |
| 5 | TBE2 | MBBR3312 | False | 22-03-2005 | 27 h. | 32 h. | 31 h. | 15 h. | 15 h. | 29-03-2005 | 120 h. | 3 |
| 6 | TBE1 | MBBR3294 | False | 19-03-2005 | 26 h. | 97 h. | 20 h. | 11 h. | 35 h. | 29-03-2005 | 189 h. | 2 |
| 7 | TBD2 | SSS025 | False | 25-03-2005 | 40 h. | 14 h. | 199 h. | 0 h. | -135 h. | 31-03-2005 | 118 h. | 7 |
| 8 | TBD1 | SSS028 | False | 23-03-2005 | 22 h. | 20 h. | 17 h. | 9 h. | 77 h. | 29-03-2005 | 145 h. | 2 |
| 9 | TBC2 | MBAL2125 | False | 25-03-2005 | 32 h. | 29 h. | 14 h. | 14 h. | 17 h. | 30-03-2005 | 106 h. | 1 |
| 10 | TBC1 | MBAL1200 | False | 28-03-2005 | 18 h. | 24 h. | 22 h. | 14 h. | 8 h. | 01-04-2005 | 86 h. | 2 |
| 11 | TBC1 | MBAL2155 | False | 21-03-2005 | 31 h. | 30 h. | 26 h. | 14 h. | 34 h. | 28-03-2005 | 135 h. | 2 |
| 12 | TBC1 | MBBL1214 | False | 01-04-2005 | 25 h. | -25 h. | 27 h. | 1 h. | 4 h. | 02-04-2005 | 32 h. | 0 |
| 13 | TBB2 | MBBL2157 | False | 26-03-2005 | 34 h. | 30 h. | 28 h. | 17 h. | 12 h. | 31-03-2005 | 121 h. | 4 |
| 14 | TBB1 | SSS125 | False | 25-03-2005 | 33 h. | 23 h. | 19 h. | 10 h. | 22 h. | 30-03-2005 | 107 h. | 2 |
| 15 | TBA1 | MBBL1196 | True | 21-03-2005 | 34 h. | 31 h. | 85 h. | 16 h. | 17 h. | 29-03-2005 | 183 h. | 8 |

Back to Weekly resume

Magnets Tested per Month

Magnet Count till March 2005

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _29_A Tuesday _29_N | TBA2 | File C0503291927.a07 & The Quench File C0503300001.ka01 (for extended program test 1: Minimum energy at 6KA) not transferred to storage. | AB/CO | TEMA | Analysed in dacq-disk2 |
| Tuesday _29_N | TBA2 | SPA occurred during the cleansing cycle of the LHC-Cycle-long-inj-10000s test. | ACR |  | Test re-launched |
| Wednesday_30_M | TBA2 | Shaft 23, sector s0001, coil A: broken into the TRU | MTM |  | Checked By Maryline. Continue magnetic measurements with E coil. |
| Wednesday_30_A | ALL | Problem with siemens MOBIC T8: when starting it up it is asking for a USB driver name. All the network parameters were lost. |  |  | Richard advised to make a reset. Richard advised toset the network parameters by re-starting the home page of internet explorer to: https://edms.cern.ch/asbuilt/plsql/mtf.home . |
| Wednesday_30_N | TBA2 | AQA 6.0 could not run for the file C0503302304.ka01, due to the error - "MPdata: no datafile". The old version of AQA ran OK. | AB/CO | AQA |  |
| Wednesday_30_N | TBA2 | Quench in Vbus-ext during the pre-cycle of LHC cycle_PF_1800s. Cryo reported a temperature increase in the current leads. Restarted LHC Cycle, but cryo noticed that the Current Lead temperature was rising. Trying to increase the cooling resulted in a loss of Cryo OK. | ACR |  | TC and N.Sammut decided to re-start a cycle with along plateau at nominal. ACR team will profit to check working of the interlock system. |
| Thursday _31_M | TBD2 | TBD2: wrong label name for SSS022(D instead of F according to MTF) non conformity must be created. |  |  | TC called N.Bourcey. He confirmed that the SSS is a Focussing.NC opened. |
| Thursday _31_A | TBE2 | MB3080: PT5.2 showed an AQA alert on QH Tau Jmp V < -1.0. Informed Richard & TC. | MTM | AQA | TC and Richard checked the result and told to continue with PTE 3KA test after discussion with Pierre. Parasitic signal probably created the alarm |
| Thursday _31_A | TBE2 | MB3080 Test : Training Quench. Description : Some jump in Trigger and current view of hfoff. | MTM |  | TC has been informed and told us to check the earth current, to continue if OK but be careful with next one. |
| Thursday _31_A | TBE2 | MB3080:On instructions from Pierre to analyze PTE with new AQA version, warning observed on Diode Info Warning: VT_Ano_VT_Cat, Diode R > 15 microOhms Pierre told to generate NC in CDPT. | MTM |  |  |
| Friday _1_M | TBD1 | SSS117: dipole corrector IJ cannot be powered. |  |  | Richard investigated the problem Meanwhile the TC asked to proceed with the MQ training. Vtaps is the bottom of the Current lead are inverted. Richard managed to power the circuit up 60 A. Inversion will be confirmed with the continuity test at warm (CQW2). |
| Friday _1_M Friday_1_A | TBD1 | SSS117: Training quench 2. SPA do to a cryo problem Second loss of Cryo OK due cold compressor down . | ACR |  |  |

| Friday_1_A | TBB1 | unable to launch PT4: acquisition system doesn't arm | AB/CO | TEMA |  |
| --- | --- | --- | --- | --- | --- |
| Friday_1_N | TBD1 | During SSW@Cold(Field Strength: 760A), SPA occured...Checked with Cryo operators, No problem with "Cryo OK" signal | ? |  |  |
| Friday_1_N | TBC1 | leak on MRB side on the MB1214 |  |  | TC informed. He decided to replace the magnet on the bench. |
| Friday_1_N | TBA2 | Over-ranges in sector 01 in aperture 1 during LHC cycle. | MTM |  | TC informed. He asked Richard to check and if needed to replace the TRU. 2 last Tests will be repeated. Richard found that the problem came from the TRUTRU 5 installed instead of TRU 3. |
| Saturday_2_M | TBD1 | Survey cannot be done on D1, reflector support are missing. Taken from TBA2 | MTM |  |  |
| Saturday_2_M | TBA1 | As MMP program was launched on A2, results of HV program on A1cannot be seen. Even after closing the MMP program window of MMP was not closing. | MTM | TEMA | Salem solved the problem. |
| Saturday_2_A | TBD1 | FS @ 760Amp: equipment missing message for Leica,PC,conditioner crate | MTM |  | Equipment ID for Leica & PC was entered manually by Salem. But he could not locate "id" for conditioner crate. He has suggested to use PT4 test (for setting data aquisition) & ramp-up the current manually. But the problem is not solved, Maryline will look after this problem on Monday. |
| Saturday_2_A | TBE1 | During the de-ramp from 12850 A, SPA at 6500 A. | ACR |  | TC checked the quench file. He discussed with ACR team. Increase of one lead at the same time. He decided to proceed with the tests. |
| Saturday_2_N | TBD1 | HF crate is down. After trying to change to crate's power supply without success | AB/PO |  | Richard came took the chassis spare from cluster G. Now it is OK |
| Sunday_3_M | TBF2 | During MMP, motor 2 was not rotating, impossible to launch the test, we continue with the quenches. | MTM |  | Richard investigated. Start position of the shaft was too close to one end stop causing this intermittent problem. Angle of the shaft axe's with respect to the motor axe modified.. |
| Sunday_3_M | TBC1 | MB1212: HV MB vs ground failed at cold BP. | ICS |  | Richard and TC informed. Investigation through the IFS box. MB put in stand by at cold by the TC. It was found that the desiccant bag was inside the IFS box. TC asked to repeat the test. Active/Passive bus bar damaged. Richard advised to continue without this.Reparation foreseen Monday. |
| Sunday_3_M | TBB1 | Powering of the MB not possible. Impossible to start RMB. | MTM |  | Richard came and found that the inductance was not changed from SSS to dipole |
| Sunday_3_N | TBA1 | Stop at training number 2: No alarm but TC found abnormal parasitic noises in D1-D2 low gain and DVT. | MTM |  | TC asked to stop the test .Not to quench and warm the magnet. Pierre will prepare a special investigation. |
| Sunday_3_N | TBF1 | Alarm on anti-cryostat. | ICS |  | Heating cable ( blue cable) not connected. |
| Sunday_3_N | TBB2 | One training quench on MB1211 with SSS protection configuration (only half a circuit).MIITS ~ 37 | MTM |  |  |
| Monday_4_A | TBE1 | While launching HV test at cold before powering on MB1210, magnet name could not be entered. | AB/CO? | HV | Name entered manually. |