Vc/17-May-05

### Week 19  starting 9 May 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  09 May 05 @ 5/9/2005 1:40:30 PM

| 1 | A1 | MBAL1180 | PT 13.2 Resist. Meas. | 297.41 |
| --- | --- | --- | --- | --- |
| 7 | A2 | MBBR3274 | PT 5.1 1.5 KA Quench LF | 2.01 HV |
| 11 | B1 |  |  | 292.94 |
| 10 | B2 | SSS122 | Prep 5 Cool Down | 203.11 |
| 8 | C1 | MBBL1215 | Prep 5 Cool Down | 130.73 |
| 2 | C2 | MBAL2073 | PT 6.5 Training 5 | 4.55 HV |
| 4 | D1 | SSS009 | PT 1 IAP @ cold | 97.64 |
| 3 | D2 | SSS158 | PT 6.1 Training 1 | 1.91 |
| 9 | E1 | MBBR3265 | ICS 4 Final connection | 294.01 |
| 12 | E2 |  |  | 292.94 |
| 6 | F1 |  |  | 292.94 |
| 5 | F2 | MBBR3293 | Prep 5 Cool Down | 1.90 |

Temperature data are taken from TBSM18 site

Coordinator this week: stephane; MTM support on call :   16-3602

15  complete Magnet Tests cycles   ( 2173 ROBTC)

09  dipoles( 1 with MM )+ 5  SSS( 0  with MM) +  1  SSSS (with MM)

Disconnected or ROBTC after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Situation on Repair :   ?

- General Points

AOB

| Magnets under Test from 01-05-2005 to 07-05-2005 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBAL2126 |  |  | 26-04-2005 | 35 h. | 90 h. | 19 h. | 18 h. | 19 h. | 04-05-2005 | 181 h. | Stripping |  |  |
| 2 | TBA2 | MBAL2135 |  |  | 25-04-2005 | 38 h. | 31 h. | 68 h. | 19 h. | 7 h. | 02-05-2005 | 163 h. | Stripping |  |  |
| 3 | TBA2 | MBAR3322 |  |  | 02-05-2005 | 28 h. | 23 h. | 18 h. | 15 h. | 11 h. | 05-05-2005 | 95 h. | Stripping |  |  |
| 4 | TBB1 | SSS090 |  |  | 30-04-2005 | 25 h. | 15 h. | 16 h. | 8 h. | 18 h. | 04-05-2005 | 82 h. | Stripping |  |  |
| 5 | TBB2 | SSS120 |  |  | 03-05-2005 | 31 h. | 12 h. | 23 h. | 10 h. | 17 h. | 07-05-2005 | 93 h. | Stripping |  |  |
| 6 | TBB2 | SSS171 |  |  | 28-04-2005 | 38 h. | 14 h. | 39 h. | 10 h. | 10 h. | 03-05-2005 | 111 h. | Stripping |  |  |
| 7 | TBC1 | MBAL3158 |  |  | 27-04-2005 | 35 h. | 32 h. | 19 h. | 13 h. | 7 h. | 01-05-2005 | 106 h. | Stripping |  |  |
| 8 | TBC2 | MBAR3324 |  |  | 01-05-2005 | 21 h. | 41 h. | 21 h. | 16 h. | 9 h. | 05-05-2005 | 108 h. | Stand By |  |  |
| 9 | TBD1 | SSS126 |  |  | 28-04-2005 | 31 h. | 14 h. | 19 h. | 7 h. | 17 h. | 02-05-2005 | 88 h. | Stripping |  |  |
| 10 | TBD1 | SSS159 |  |  | 02-05-2005 | 39 h. | 15 h. | 17 h. | 14 h. | 8 h. | 06-05-2005 | 93 h. | Stripping |  |  |
| 11 | TBD2 | SSS615 |  |  | 28-04-2005 | 44 h. | 65 h. | 35 h. | 8 h. | 53 h. | 06-05-2005 | 205 h. | Stand By |  |  |
| 12 | TBE1 | MBAL2172 |  |  | 28-04-2005 | 27 h. | 23 h. | 28 h. | 16 h. | 7 h. | 03-05-2005 | 101 h. | Stripping |  |  |
| 13 | TBE2 | MBAL2173 |  |  | 30-04-2005 | 24 h. | 43 h. | 40 h. | 15 h. | 0 h. | 05-05-2005 | 122 h. | Not Found or Thermal cycle |  |  |
| 14 | TBF1 | MBAL1179 |  |  | 29-04-2005 | 31 h. | 24 h. | 36 h. | 14 h. | 9 h. | 04-05-2005 | 114 h. | Stripping |  |  |
| 15 | TBF2 | MBBL2164 |  |  | 29-04-2005 | 38 h. | 52 h. | 38 h. | 15 h. | 13 h. | 06-05-2005 | 156 h. | Stand By |  |  |
| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| 1 | MBAL2126 | 11723 A. (D2-U) | 12491 A. (D1- L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 2 | MBAL2135 | 11000 A. (No Data) | 12064 A. (No Data) | 12186 A. (No Data) | 12444 A. (No Data) | 12702 A. (No Data) | 12850 A. | - | - | - | - | - | - | - |
| 3 | MBAR3322 | 12139 A. (D1-L) | 12809 A. (D2-L) | 12851 A. | - | - | - | - | - | - | - | - | - | - |
| 4 | SSS090 | 12140 A. (QI-34) | 12224 A. (QE-34) | 12201 A. (QE-34) |  | - | - | - | - | - | - | - | - | - |
| 5 | SSS120 | 12039 A. (QE-34) | 12484 A. (QI-12) | 12732 A. (QE_34) |  | - | - | - | - | - | - | - | - | - |
| 6 | SSS171 | 12052 A. (QE_34) | 12039.6 A. (QE-34) | 12088.1 A. (QE-34) | 12242 A. (QI-34) | 12748 A. (QI_12) | 12850 A. | - | - | - | - | - | - | - |
| 7 | MBAL3158 | 12645 A. (D2 - L) | 12807 A. (D2-L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 8 | MBAR3324 | 12231.5 A. (D2-U) | 12715 A. (D2-U) | 0 A. (No Data) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 9 | SSS126 | 11850 A. (QI_12) | 12498 A. (QE-12) | 12672 A. (QI-34) |  | - | - | - | - | - | - | - | - | - |
| 10 | SSS159 | 12338.6 A. (QI_34) | 12538.4 A. (QE_12) | 12782 A. (QI-12) |  | - | - | - | - | - | - | - | - | - |
| 11 | SSS615 | 5489 A. (Q_I_12) | 5820 A. | - | - | - | - | - | - | - | - | - | - | - |
| 12 | MBAL2172 | 10314 A. (D1-L) | 11922.6 A. (D1_U) | 12509.8 A. (D2-L) | 12795.4 A. (D1_L) |  | - | - | - | - | - | - | - | - |
| 13 | MBAL2173 | 9408.8 A. (D2-L) | 11068.8 A. (D1-L) | 11941 A. (D1-L) | 12050 A. (D2-L) | 12259.4 A. (D2-L) | 11785 A. (D2_L) | 12423 A. (D2_L) | 12177 A. (D2-L) | 12699 A. (D2-L) |  | - | - | - |
| 14 | MBAL1179 | 9858 A. (D1-L) | 11507 A. (D2-L) | 11941 A. (D1-L) | 12020 A. (D1-U) | 12271 A. (D2-L) | 12575 A. (D2-L) | 12819 A. (D1-U) | 12860 A. | - | - | - | - | - |
| 15 | MBBL2164 | 12553 A. (D2-U-S02) | 12900 A. | 12850 A. | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |

### Week 17  :  01May to 07 may 05  (15 Magnet Tests)

### [    On bench rapid thermal cycle   ]

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Wednesday _04_A | TBD2 | Stand by of SIMCO survey team reserved till Wednesday 17h00. | MTM |  | TC called Guy who accepted to make the survey in the Aperture 2 of the SSS615 |
| Wednesday _04_N | TBA1 | Gamelle missing. ICS cannot close the magnet | MTM |  |  |
| Thursday _05_A |  | HV mobile rack 92 has a bad D21 connector. | MTM |  | Monday Richard replaced the wrong module mounted in connector D20. Rack is back in operation. |
| Thursday _05_N | TBB1 | Temperatures not correct on Video display for all the benches. | ACR |  | Jerome Axensalva found that a server was down Problem solved.. |
| Friday _06_M | TBA1 | Bad alignment of the shafts with the TRU. | MTM |  | Didier found that the table of the TRU had been moved laterally. He corrected the position of the table. ICS should act on the yellow trolley. |
| Friday _06_M | TBE1 | HV at cold : temperature sensor vs ground failed on MB3107. | MTM |  | Salem found that a pin was bent in the temp. sensor connector in the IFS box. He repaired the connector and the HV test passed OK |
| Friday _06_A | TBB | Tema is not working :No communication possible,LD_LIBRARY_PATH environment variable missing | AB/CO | TEMA | Salem contacted Adrian. PC next to the sun was moved and the network cable was unplugged |
| Friday _06_N | TBD2 | SSS158: No response for lf DAQ. | ICS? |  | Salem asked to trigger the DAQ using the card in the safety matrix: no trigger. The bench was in the special SSS configuration. |
| Saturday_07_M | TBD1 | While launching Prep 2.2 (IAP@warm), PLCSE indicated that the QH are not charged, But actually heaters were charged (approx. 840V). | MTM |  | Salem found a bad communication between PLCSE and SERVER application. He solved the problem by resetting server and PLC. |
| Saturday_07_M | TBA1 | Over-range in AP2 during MM2 | MTM |  | Salem found the origin of the problem: end cover touching the flat cable. |
| Saturday_07_M | TBA1 | Waterflow problem. PLC for water-flow control not responding, Quench could not be launched. | MTM |  | Salem resetted the PLCCU: OK |
| Saturday_07_N | TBA1 | Cryo loss during LHC cycle due to a rapid drop in helium flow to the B current lead. SPA during the pre-cycle. | MTM |  | LHC cycle to be repeated. |
| Saturday_07_N | TBF2 | Cryo team received a high temperature warning (65 deg C) on the anticryostat L1 when the cables were connected. Alarm on the anticryostat. | MTM MTM |  | Wrong parameter in the L1 anticryostat? Salem updated the parameters from MTF. |
| Sunday_08_M | TBC1 | Bad manipulation during C1 connection, General cut of electricity Ext trigger interlock on RMA | ICS |  | A1 was running with a LHC cycle, the cycle was stopped., RMA also switched off. TCR will send someone for repair. Salem solved it by switching on extraction switch |
| Sunday_08_M | TBA1 | Valve opened but no water. | MTM |  | Salem and TCR called. Salem solved the problem by resetting PLC-CU and |
| Sunday_08_A Mondat_09_M | TBD1 | IAP @cold on SSS09: spurious output trigger during AC compensation. | MTM |  | Investigation by Salem, Richard,Giorgio. Problem coming from the magnet, VTaps inverted inside. It should have been detected with IAP @warm. |
| Sunday_08_M | TBC1 | Login problem on cluster C due to an incorrect desktop (application not running) | AB/CO |  | M.Marchesotti solved the problem |