Ev/23-Jun-05

### Week 25  starting 20 June 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  20 June 05 @: 20/06/2005 07:53:29

LHC SM18 Magnet Test Facility

| 11 | A1 | MBAR3343 | PT 14 Magnet MAPS | 299.92 |
| --- | --- | --- | --- | --- |
| 10 | A2 | MBBL1235 | ICS 4 Final connection | 298.94 |
| 3 | B1 | SSS223 | PT 1 IAP @ cold | 1.90 |
| 9 | B2 | SSS219 | Prep 5 Cool Down | 73.55 |
| 4 | C1 | MBAR3252 | Prep 5 Cool Down | 4.66 |
| 2 | C2 | MBAL2191 | PT 6.8 Training 8 | 1.90 |
| 12 | D1 | SSS200 | PT 13.2 Resist. Meas. | 297.41 |
| 1 | D2 | SSS616 | Prep 5 Cool Down | 205.28 |
| 6 | E1 | MBAL3103 | PT 12 Warm Up | 135.88 |
| 7 | E2 | MBBR3362 | Prep 5 Cool Down | 84.54 |
| 8 | F1 | MBBR3361 | PT 12 Warm Up | 206.66 |
| 5 | F2 | MBAL2196 | Prep 5 Cool Down | 94.06 |

Temperature data are taken from TBSM18 site

Coordinator this week: S. Sanfilippo;

MTM support on call : R. Mompo  16-3602

16  complete Magnet Tests cycles   ( S277  ROBTC )

10  dipoles( 2 with MM )+ 5  SSS( 0  with MM) +  1 SSSS (with MM)

[SSSS639 put in standby & 1236 put in standby because Cryo Leak]

Disconnected or ROBTC after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- General Points

  - TV Network maintenance (Should inform SMCR next time 15-06-05).

  - Vtaps or current lead Wrongly connected by ICS (To enforce again connection quality)

  - Pumping down difficult for Cryo (Saturday problem).

  - 1236 leak inside magnet.

  - Mobile phone lost by ICS.

AOB

Tests Information

| Magnets under Test from 12-06-2005 to 18-06-2005 |
| --- |
| - | Bench | Magnet | Shafts | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |
| 1 | TBA1 | MBAR3305 |  | 08-06-2005 | 23 h. | 59 h. | 21 h. | 14 h. | 15 h. | 14-06-2005 | 132 h. | Stripping |
| 2 | TBA2 | MBAR3352 |  | 08-06-2005 | 24 h. | 28 h. | 39 h. | 18 h. | 8 h. | 13-06-2005 | 117 h. | Stripping |
| 3 | TBB1 | SSS283 |  | 14-06-2005 | 24 h. | 13 h. | 20 h. | 7 h. | 15 h. | 17-06-2005 | 79 h. | Stripping |
| 4 | TBB1 | SSS285 |  | 10-06-2005 | 29 h. | 12 h. | 16 h. | 6 h. | 14 h. | 13-06-2005 | 77 h. | Stripping |
| 5 | TBB2 | SSS213 |  | 13-06-2005 | 32 h. | 13 h. | 18 h. | 12 h. | 16 h. | 17-06-2005 | 91 h. | Stripping |
| 6 | TBB2 | SSS220 |  | 09-06-2005 | 29 h. | 14 h. | 15 h. | 14 h. | 6 h. | 12-06-2005 | 78 h. | Stripping |
| 7 | TBC1 | MBAL1224 |  | 13-06-2005 | 26 h. | 22 h. | 18 h. | 13 h. | 21 h. | 17-06-2005 | 100 h. | Stripping |
| 8 | TBC1 | MBAR3277 |  | 09-06-2005 | 22 h. | 22 h. | 15 h. | 15 h. | 8 h. | 13-06-2005 | 82 h. | Stripping |
| 9 | TBC2 | MBBL3100 |  | 11-06-2005 | 31 h. | 23 h. | 15 h. | 14 h. | 10 h. | 15-06-2005 | 93 h. | Stripping |
| 10 | TBD1 | SSS255 |  | 11-06-2005 | 35 h. | 31 h. | 28 h. | 15 h. | 27 h. | 16-06-2005 | 136 h. | Stripping |
| 11 | TBD2 | SSS639 |  | 07-06-2005 | 101 h. | 21 h. | 59 h. | 11 h. | 31 h. | 18-06-2005 | 223 h. | Stand By |
| 12 | TBE1 | MBBL1231 |  | 11-06-2005 | 38 h. | 21 h. | 17 h. | 18 h. | 8 h. | 16-06-2005 | 102 h. | Stripping |
| 13 | TBE2 | MBBL1236 |  | 10-06-2005 | 33 h. | 0 h. | 0 h. | 1 h. | 6 h. | 12-06-2005 | 40 h. | Stand By |
| 14 | TBE2 | MBBL3125 |  | 12-06-2005 | 26 h. | 23 h. | 22 h. | 13 h. | 17 h. | 16-06-2005 | 101 h. | Stripping |
| 15 | TBF1 | MBAR3347 |  | 09-06-2005 | 28 h. | 95 h. | 17 h. | 17 h. | 11 h. | 16-06-2005 | 168 h. | Stripping |
| 16 | TBF2 | MBAL1225 |  | 07-06-2005 | 26 h. | 32 h. | 157 h. | 13 h. | 25 h. | 18-06-2005 | 253 h. | Stripping |

Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAR3305 | 12385 A. (D2-U) | 12501 A. (D2-L) | 12710.9 A. (D2_U) |  | - | - | - | - | - | - |
| 2 | MBAR3352 | 12488 A. (D1-L) | 0 A. (No Data) | 12851 A. | - | - | - | - | - | - | - |
| 3 | SSS283 | 11974 A. (QE_12) | 12141 A. (QI_12) | 12220 A. (QI_34) |  | - | - | - | - | - | - |
| 4 | SSS285 | 11835 A. (QE-12) | 12342.2 A. (QE_34) | 12442.9 A. (QE-34) |  | - | - | - | - | - | - |
| 5 | SSS213 | 12833 A. (QI_12) | 12850 A. | - | - | - | - | - | - | - | - |
| 6 | SSS220 | 11250 A. (QI-12) | 12358 A. (QE-12) | 12539 A. (QE12) |  | - | - | - | - | - | - |
| 7 | MBAL1224 | 10845 A. (D2-L) | 12512 A. (D1-L) | 12620 A. (D1-L) |  | - | - | - | - | - | - |
| 8 | MBAR3277 | 11371 A. (D2-U) | 12394.6 A. (D2-U) | 12506 A. (D1-U) |  | - | - | - | - | - | - |
| 9 | MBBL3100 | 11527 A. (D1_U) | 12850 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - |
| 10 | SSS255 | 11260.1 A. (QI_12) | 12168 A. (QI_12) | 12356 A. (QI_34) |  | - | - | - | - | - | - |
| 11 | SSS639 | 5820 A. | - | - | - | - | - | - | - | - | - |
| 12 | MBBL1231 | 12050 A. (D2-U) | 12525 A. (D1-U) | 12857 A. (D2-U) |  | - | - | - | - | - | - |
| 13 | MBBL1236 |  | - | - | - | - | - | - | - | - | - |
| 14 | MBBL3125 | 11858.1 A. (D2_L) | 12680.1 A. (D2_U) | 12850 A. | - | - | - | - | - | - | - |
| 15 | MBAR3347 | 12719.9 A. (D1-U) | 9860 A. (D1-U) | 12850 A. | - | - | - | - | - | - | - |
| 16 | MBAL1225 | 11347 A. (D1-L-S01) | 12744 A. (No Data) | 12850 A. | - | - | - | - | - | - | - |

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _14_A | TBD1 | SSS255: Dummy resistor blown up when discharging quench heater capacitor | MTM |  | Salem replaced the dummy resistor. |
| Tuesday _14_A | TBE2 | No data in IAP file | MTM |  | Salem found that there was no trigger hf. Test repeated after “ reset”. |
| Tuesday _14_A | TBD1 | SSS255: Not able to power CD circuit During IAP at warm : error confirmed between the two circuits (MQT,sextupole) | ICS |  | Investigation by Salem. CD and RF circuit inverted during the connection (CD leads to EF bus-bar & EF leads to CD bus-bar) |
| Tuesday _14_N | TBA2 | Cryo leak in MB2524 (MRB side). | ICS |  | Joint forgotten. |
| Wednesday _15_M | TBF2 | SPA during machine cycle on MB1225. | ACR? AB/CO? |  | Temperature auxiliary fault ? Communication problem? |
| Wednesday _15_A | TBD2 | SPA during SSW measurements. | ? |  | Monitoring of SPA will be installed by Giorgio in the PLC. |
| Wednesday _15_N | TBD2 | MBBL3125: Resistance measurements in CDW2 show that this magnet has a diode. | ? |  | Information provided in the file http://lhcdipcoor.web.cern.ch/LHC-dipcoor/diodes/dioderesume.pdf for magnet MBBL3125 is wrong. Diode test was not performed. The diode test has to be performed in any case. |
| Wednesday _15_N | TBF2 | Machine cycle LHC_Pf_1800 (last test) was stopped during the measurement plateau because of a SPA | AB/PO? MTM |  | N.Sammut (contact person) was not informed. The test was not repeated. NB: This test is dedicated to check the influence of a long pre_injection porch and the data coming from it are very important |
| Thursday _16_M | TBF2 | The SSL quench at 4K is with a current of 10173A at 1350mBar higher than tolerance (<10100A), | MTM |  | Salem contacted. Decision was to warm up the magnet. |
| Thursday _16_A | TBE1 | Permanent Triggers present on VSupra_ext & VClamp_Ext for preparatory test PT2.1. | MTM |  | Salem has by-passed the Vtaps. Ok to continue. |
| Thursday _16_A | TBE1 | MBAL 3103, Prep2.1: the value of Quench Heater H4,TminV (=1741ms), is too high. | MTM |  | Salem checked the file and gave the OK to continue because the problem is coming from the QH power supply. This power supply was changed. |
| Thursday _16_N | TBA2 | MB2524: Data from AQA not available after quench. Lf config file not found in control room sun. Error displayed in local sun : Error 11 occurred at open file | AB/CO | lf | Data could not be validated in AQA immediately. But quench analysed locally. Then Salem transferred the data to the control room sun. |
| Friday_17_M | TBA2 | RMA cannot be restarted. | AB/PO |  | AB-PO experts called by Salem. Test in short circuit. Firing card changed. |
| Friday _17_N | TBA1 | AC offset for D1_L-U is very high. Equipment support will look after this problem. D1-D2 signal was flat. Equipment support has been contacted and Salem told that he will look for the problem and inform | MTM |  | Salem changed the potaim card Salem changed the potaim card. |
| Saturday _18_M |  | One analyser pumping unit is not working. Cooling down slow. | ACR |  | J.P.Lamboy gave some instruction. |
| Saturday _18_M |  | General stop due to water cut. Liquefactor stopped.. Magnets are warming up. | ? |  | Stop of all the tests at 1.9 K up to the afternoon |
| Saturday _18_M | TBB2 | IAP@warm 2.2C (corrector clamp check) on SSS219 could not be launched, triggers present on V_clamp E-F. | ICS |  | ISC team was asked to check the connections.Clamp re-done |
| Sunday _19_A | TBD2 | ICS found that one heating circuit of the anticryostat on SSS616 is opened. | MTM |  | Salem came and checked with ICS: no open circuit. But he found that the anticryostat parameters were wrong because of a new anticryostat. New parameters implemented by P.Legrand. |
| Sunday _19_A Monday_20_M | TBC2 TBD2 | Liquefactor problem: BP too high. Regulation difficult. Low capacity of pumping. Cooling of special SSS slow: Cryo team encountered several problems during the night causing very slow recovery rates on all benches. Another problem with BP too high | ACR |  |  |
| Sunday _19_N | TBD1 | SSS200: during resistance measurements (PT 13.2): Corrector Voltage Taps (E & F) and Sextupole Voltage Taps (C & D) seems to be inverted. | MTM |  | Salem advised to check with ICS during the magnet opening. SSS200 had a new configuration of connection: the values are OK. |
| Sunday _19_N | TBC2 | Difference in current reading between AQA and Voltmeter (more than 10 Amps) | MTM | AQA? | Checked Monday for the next quenches. The ADC card may be changed. |
| Monday _20_A | TBA1 | HV test could not be started. Voltage exceeded the upper limit. | MTM |  | Investigation by Richard. |