### CHOHAN /08-August-05

### Week 32  starting 08 Aug 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  08 August 05 : 8/8/2005 2:33:17 PM

| 8 | A1 | MBBL2194 | ICS 3 Traveller after warm test | 295.09 |
| --- | --- | --- | --- | --- |
| 2 | A2 | MBBL1190 | PT 13.1 IAP @ Warm | 299.11 |
| 11 | B1 | LE-C11R7 | Under Connection | 295 |
| 3 | B2 | MBAL1241 | PT 13.2 Resist. Meas. | 296.34 |
| 7 | C1 | MBAL2145 | Prep 5 Cool Down | 96.03 |
| 9 | C2 |  |  | 295.98 |
| 1 | D1 | SSS206 | PT 12 Warm Up | 122.71 |
| 6 | D2 | SSS607 | PT 6.7 Training 7 | 2.35 |
| 5 | E1 | MBBL2187 | PT 6.1 Training 1 | 4.89 |
| 10 | E2 | MBBL1191 | ICS 1 WP04 HV Test Warm | 295.8 |
| 4 | F1 | MBAL1248 | PT 12 Warm Up | 38.2 |
| 12 | F2 | MBBL2166 | Prep 0 Scan Bench | 295.09 |

Tempr. from TBSM18 site

Coordinator this week: Marta; MTM support 16-3602

12 complete Magnet Tests cycles

Only 10 dipoles ( 1 with MM )+ 2 SSS(without MM & SSW )

Disconnected or ROBTC after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- TWO Leaks on F cluster Magnets last week (3387 & 3187  )

Special  Points :

- Special SSS 607 on D2  (Kevin); SSW :Marco

- 1084  Lifetime tests : 200A/s cycles (Giorgio) + Instrumn.

- SSS206 on D1 : Life time tests : Cryo

- Connect. Cryostat Tests B1  (Louis Walckiers)

- Rotary Switch for Corrector powering (awaiting Giorgio?).

- Cluster F “INTERCOM Box” repair (awaiting Giorgio?).

| Magnets under Test from 31-07-2005 to 06-08-2005 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA2 | MBBR3381 |  |  | 30-07-2005 | 30 h. | 27 h. | 43 h. | 16 h. | 10 h. | 04-08-2005 | 126 h. | Stripping |  |  |
| 2 | TBB1 | SSS177 |  |  | 26-07-2005 | 26 h. | 26 h. | 19 h. | 8 h. | 11 h. | 31-07-2005 | 90 h. | Stripping |  |  |
| 3 | TBC1 | MBBR3311 |  |  | 01-08-2005 | 22 h. | 28 h. | 21 h. | 30 h. | 20 h. | 06-08-2005 | 121 h. | Stripping |  |  |
| 4 | TBC1 | MBBR3342 |  |  | 26-07-2005 | 46 h. | 23 h. | 18 h. | 20 h. | 33 h. | 01-08-2005 | 140 h. | Stripping |  |  |
| 5 | TBC2 | MBAL1227 |  |  | 29-07-2005 | 27 h. | 37 h. | 15 h. | 14 h. | 19 h. | 02-08-2005 | 112 h. | Stripping |  |  |
| 6 | TBD1 | SSS254 |  |  | 26-07-2005 | 47 h. | 14 h. | 21 h. | 12 h. | 21 h. | 31-07-2005 | 115 h. | Stripping |  |  |
| 7 | TBE1 | MBAL1178 |  |  | 01-08-2005 | 39 h. | 24 h. | 19 h. | 16 h. | 18 h. | 05-08-2005 | 116 h. | Stripping |  |  |
| 8 | TBE1 | MBBR3380 |  |  | 28-07-2005 | 26 h. | 20 h. | 15 h. | 16 h. | 13 h. | 31-07-2005 | 90 h. | Stripping |  |  |
| 9 | TBE2 | MBAR3368 |  |  | 29-07-2005 | 26 h. | 26 h. | 16 h. | 16 h. | 12 h. | 02-08-2005 | 96 h. | Stripping |  |  |
| 10 | TBE2 | MBBL3097 |  |  | 02-08-2005 | 18 h. | 29 h. | 17 h. | 16 h. | 21 h. | 06-08-2005 | 101 h. | Stripping |  |  |
| 11 | TBF1 | MBAR3375 |  |  | 30-07-2005 | 51 h. | 20 h. | 14 h. | 14 h. | 12 h. | 04-08-2005 | 111 h. | Stripping |  |  |
| 12 | TBF1 | MBBL3138 |  |  | 04-08-2005 | - | - | - | - | - | 05-08-2005 | 30 h. | Stand By |  |  |
| 13 | TBF2 | MBBR3379 |  |  | 28-07-2005 | 34 h. | 31 h. | 24 h. | 14 h. | 25 h. | 02-08-2005 | 128 h. | Stripping |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBR3381 | 10397.7 A. (D2_U_S01) | 12616.4 A. (D2_U_S13) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 2 | SSS177 | 12016 A. (QE-12) | 12414 A. (QI-12) | 12566 A. (QE-12) | 12528.5 A. (QE-34) |  | - | - | - | - | - | - | - | - |
| 3 | MBBR3311 | 12225 A. (D2_U) | 12548 A. (D1_L) | 12753 A. (D2-L) |  | - | - | - | - | - | - | - | - | - |
| 4 | MBBR3342 | 11868.5 A. (D1-L) | 12603 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 5 | MBAL1227 | 11349 A. (D2-U) | 12553.1 A. (D1-L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 6 | SSS254 | 11499.9 A. (QI-12) | 12024 A. (QE_12) | 12215 A. (QI-12) |  | - | - | - | - | - | - | - | - | - |
| 7 | MBAL1178 | 12097.8 A. (D2-L) | 12748 A. (D1-U) | 12857 A. | - | - | - | - | - | - | - | - | - | - |
| 8 | MBBR3380 | 11477.8 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 9 | MBAR3368 | 10537 A. (D2-L) | 12849 A. (D2-L) | 12858 A. | - | - | - | - | - | - | - | - | - | - |
| 10 | MBBL3097 | 9611 A. (D1_U) | 12749 A. (D2-L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 11 | MBAR3375 | 12453 A. (D2-U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 12 | MBBL3138 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 13 | MBBR3379 | 12337.9 A. (D2_U) | 12759.7 A. (D1_L) | 12484 A. (D2-U) | 12850 A. | - | - | - | - | - | - | - | - | - |

Statistics For end July 2005 :   We have tested More than Half of the ‘Expected to Test’  Magnets

Statistics For end July 2005 :   We have tested More than Half of the Expected to Test Magnets

Prediction on 4 Aug 05 For Termination by end 2006 :  Desired Magnets & Desired Bench Configurations

| Assumption : Testing Time (Upto may 05) : Dipole -131 hrs;SSS- 143 hrs, After May-05, Dipole - 120 Hrs, ARCSSS - 125 Hrs, IR SSS - 10 days |
| --- |
| Date | Magnets per Month Dipoles | Magnets per Month ARC SSS | Magnets per Month IR SSS | Cumulative Dipoles | Cumulative ARC SSS | Cumulative IR SSS | Cumulative Total | Magnets per Month | Configuration |
| Dec-03 | 95 | 2 |  | 95 | 2 | 0 | 97 |  | D-6/AS-1 |
| Jan-04 | 1 |  |  | 96 | 2 | 0 | 98 | 1 | D-8 |
| Feb-04 | 28 | 2 |  | 124 | 4 | 0 | 128 | 30 |  |
| Mar-04 | 37 | 1 |  | 161 | 5 | 0 | 166 | 38 |  |
| Apr-04 | 36 | 2 |  | 197 | 7 | 0 | 204 | 38 |  |
| May-04 | 30 | 1 |  | 227 | 8 | 0 | 235 | 31 | D-9/AS-1 |
| Jun-04 | 37 | 3 |  | 264 | 11 | 0 | 275 | 40 | D-10/AS-2 |
| Jul-04 | 38 | 6 |  | 302 | 17 | 0 | 319 | 44 | D-9/AS-3 |
| Aug-04 | 43 | 7 |  | 345 | 24 | 0 | 369 | 50 |  |
| Sep-04 | 41 | 6 | 1 | 386 | 30 | 1 | 417 | 48 | D-8/AS-3/IS-1 |
| Oct-04 | 47 | 10 | 1 | 433 | 40 | 2 | 475 | 58 |  |
| Nov-04 | 48 | 11 |  | 481 | 51 | 2 | 534 | 59 | D-9/AS-2/IS-1 |
| Dec-04 | 15 | 3 | 1 | 496 | 54 | 3 | 553 | 19 |  |
| Jan-05 | 6 | 1 | 0 | 502 | 55 | 3 | 560 | 7 | D-9/AS-2/IS-1 |
| Feb-05 | 43 | 13 | 2 | 545 | 68 | 5 | 618 | 58 |  |
| Mar-05 | 51 | 13 | 2 | 596 | 81 | 7 | 684 | 66 |  |
| Apr-05 | 51 | 21 | 0 | 647 | 102 | 7 | 756 | 72 | D-8/AS-4 |
| May-05 | 48 | 27 | 1 | 695 | 129 | 8 | 832 | 76 | D-8/AS-3/IS-1 |
| Jun-05 | 46 | 21 | 3 | 741 | 150 | 11 | 902 | 70 |  |
| Jul-05 | 48 | 15 | 4 | 789 | 165 | 15 | 969 | 67 | D-8/AS-2/IS-1 |
| Aug-05 | 40 | 7 | 4 | 829 | 172 | 19 | 1020 | 51 | DipoleSSS Special Tests D9/AS-1/IS-1 |
| Sep-05 | 46 | 7 | 4 | 875 | 179 | 23 | 1077 | 57 | D-8/AS-3/IS-1 |
| Oct-05 | 46 | 18 | 4 | 921 | 197 | 27 | 1145 | 68 |  |
| Nov-05 | 46 | 18 | 4 | 967 | 215 | 31 | 1213 | 68 |  |
| Dec-05 | 23 | 9 | 4 | 990 | 224 | 35 | 1249 | 36 |  |
| Jan-06 | 0 | 0 | 0 | 990 | 224 | 35 | 1249 | 0 | SHUTDOWN |
| Feb-06 | 36 | 18 | 4 | 1026 | 242 | 39 | 1307 | 58 |  |
| Mar-06 | 48 | 18 | 4 | 1074 | 260 | 43 | 1377 | 70 |  |
| Apr-06 | 48 | 18 | 4 | 1122 | 278 | 47 | 1447 | 70 |  |
| May-06 | 36 | 25 | 8 | 1158 | 303 | 55 | 1516 | 69 | D-6/AS-4/IS-2 |
| Jun-06 | 36 | 25 | 8 | 1194 | 328 | 63 | 1585 | 69 |  |
| Jul-06 | 36 | 25 | 8 | 1230 | 353 | 71 | 1654 | 69 |  |
| Aug-06 | 36 | 25 | 8 | 1266 | 378 | 79 | 1723 | 69 |  |
| Sep-06 | 36 | 25 | 8 | 1302 | 403 | 87 | 1792 | 69 |  |
| Oct-06 | 36 | 25 | 4 | 1338 | 428 | 91 | 1857 | 65 |  |
| Nov-06 | 30 | 7 |  | 1368 | 435 | 91 | 1894 | 37 |  |
| Dec-06 |  |  |  | 1368 | 435 | 91 | 1894 | 0 |  |
| Jan-07 |  |  |  | 1368 | 435 | 91 | 1894 | 0 |  |
| Feb-07 |  |  |  | 1368 | 435 | 91 | 1894 | 0 |  |
| Mar-07 |  |  |  | 1368 | 435 | 91 | 1894 | 0 |  |
| Apr-07 |  |  |  | 1368 | 435 | 91 | 1894 | 0 |  |
| May-07 |  |  |  | 1368 | 435 | 91 | 1894 | 0 |  |
| Basic Targets | 1232 | 360 | 114 | 114 (IR SSS)=50 (MS) + 32 (DS) + 32 (DS-MQ Type) |
| Re-Tested | 135.52 | 39.6 | 12.54 |  |  |  |  |  |  |
| Probable Target | 1367.52 | 399.6 | 126.54 |  |  |  |  |  |  |
| Total Magnets | 1893.66 |  |  |  |  |  |  |  |  |
|  |  |  |  |  | Basic Targets | 1232 | 392 | 82 |  |
|  |  |  |  |  | Re-Tested | 135.52 | 43.12 | 9.02 |  |
|  |  |  |  |  | Probable Target | 1367.52 | 435.12 | 91.02 |  |
|  |  |  |  |  | Total Magnets | 1893.66 |  |  |  |
| Estimates with Cryo Enchancement |  |  |  |  |  |  |
| Dipoles | ARC-SSS |  |  |  |  |  |  |
| 8-Benches | 48 | 3- Benches | 18.78261 |  |  |  |  |  |  |
| 7-Benches | 42 | 4-Benches | 25.04348 |  |  |  |  |  |  |
| 6-Benches | 36 | 6-Benches | 37.56522 |  |  |  |  |  |  |
| 9-Benches | 54 | 2- Benches | 12.52174 |  |  |  |  |  |  |
| NOTE |  |  |  |  |  |  |  |  |  |
| Average Time 2004 |  |  | Average Time 2005 | 04-Jul-05 |  |
| Item | Dipoles | SSS |  |  | Item | Dipoles | SSS |  |  |
| Connect Time | 30 | 82 |  |  | Connect Time | 30 | 30 |  |  |
| Cooling Time | 45 | 32 |  |  | Cooling Time | 30 | 20 |  |  |
| Cold Test Time | 34 | 75 |  |  | Cold Test Time | 30 | 35 |  |  |
| Warming Time | 20 | 12 |  |  | Warming Time | 15 | 10 |  |  |
| Disconnect Time | 8 | 8 |  |  | Disconnect Time | 15 | 20 |  |  |
| Total Time | 137 | 209 |  |  | Total Time | 120 | 115 |  |  |
| Average Time Due to Cyro Enchancement |  |  |  |  |  |  |  |
| Item | Dipoles | SSS |  |  |  |  |  |  |  |
| Connect Time | 30 | 82 |  |  |  |  |  |  |  |
| Cooling Time | 33 | 30 |  |  |  |  |  |  |  |
| Cold Test Time | 34 | 75 |  |  |  |  |  |  |  |
| Warming Time | 15 | 12 |  |  |  |  |  |  |  |
| Disconnect Time | 8 | 8 |  |  |  |  |  |  |  |
| Total Time | 120 | 207 |  |  |  |  |  |  |  |

CRYO    TIMES

ICS TIMES

MAGNETIC MEASUREMENT TIMES

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _02_A | TBA1 | MB2079: CDBP HV test failed on step Quench Heater Vs Gnd. | MTM |  | Special test: Failure at 3KV step for QH 212 (LF).TC and Federico contacted. Advise to perform a small quench 3 KA and then repeat HV test. Quench performed and again the failure. TC advised to proceed with the first quench and check afterwards the insulation. Federico opened the IFS box, blew the air to remove dust, perform the test manually, then OK after the first training quench. |
| Tuesday _02_A Wednesday _03_M | TBA2 | Over-range in sector 0001 of Ap2 during MM1 and in Ap1, S0809 (known problem). LHC Cycle & Load-line. Over-ranges in all measurements, Ap1 S0809, Ap2 S0001. | MTM |  | Federico contacted. This impeded to find the correct position of the shaft for ap2. Nikolay asked to continue the test. Sector 001 of shaft 29 sent for repair. |
| Tuesday _02_N | TBD2 | Gamelle at- 6 C(SSS607) | ICS |  | Federico contacted and explained that it may be due to bad thermal insulation. Confirmed: no thermal insulation installed. Gas flow N2 on. |
| Wednesday _03_M | TBA2 | SPA at 12850 A during training quench. | ACR? |  |  |
| Wednesday _03_A | TBC1 | MB3311: AQA Warnings on PT 5.1 - quench origin out of coil & Current HF-LF discrepancy; | MTM |  | Eq support contacted – He asked to repeat PT5.1 by triggering LF QH 3&4. Result OK. |
| Wednesday _03_A | TBC1 | MB3311: AQA Warnings on PT 5.2 –tmin<500ms; | MTM |  | Federico checked the quench: he advised to continue. |
| Wednesday _03_N | TBB2 | Communication error between the HV-PLC & Server. | AB/CO |  | Fault cleared after few attempts of reset. |
| Thursday _04_M | TBE1 | MB1178: Quench occurred at 7 kA during SPA test (PT8) | AB/PO |  | Federico explained that the power supply tripped and then Vsum detected it. |
| Thursday _04_A | TBD2 | XY measurements on SSS607: Movement in one direction (negative) along the X-axis not possible. Movement along Y direction inverted. | MTM |  | Nikolay advised to create a new DB file with Y direction inverted. OK (For system 3, script has to be run befor the test campaign) But problem remained in X direction. |
| Thursday _04_N | TBD1 | Interlock: no cryo OK (2 times): powering of corrector stopped (2x1.5 hours) | ACR |  |  |
| Thursday _04_N | TBD1 | Oracle problem: HWREC switched off. | IT |  |  |
| Friday_05_M | TBD1 | Loss of cryo OK | MTM |  | Good value of inductance entered. Test ran form local sun. |
| Friday_05_M | TBA2 | DB error in TEMA because the Tsensor is not attached to the magnet traveller of the MB1190. | CRI |  | Maryline opened a NC.. |
| Friday_05_M | TBF1 | During capillary tests, leak found on the heat exchanger in the magnet 3138 |  |  | Nikolay informed. Magnet removed from the bench for repair. |
| Friday_05_N | TBD1 | Cryo OK lost: training aborted. OK after 30 min. | ACR |  |  |
| Friday_05_N | TBF2 | Cryo leak in MB3387: leak on the screen. |  |  | TC informed. Magnet re-opened and removed. |
| Saturday_06_A | TBD1 | Cryo Ok has been lost due to a peak temperature. Cryo Ok will be available after supervision timer (3600s). | ACR |  | ACR support contacted. |
| Saturday_06_M | TBB2 | Unable to launch HV test due to HV cable connection problem PLCSE - PLC HV communication error. | MTM |  | Equip.support solved the problem temporarily. Cable connector of bench must be fixed (Action R.Mompo) |
| Sunday_07_A | TBA2 | Hf DAQ could not be armed for warm-up quench After quench, analysis shows that earth current (MB1190) exceeded 10 mA. | AB/CO | hf | Salem fixed the problem. Salem suggests to cool down to 4K and repeat CDAP HV test. |
| Sunday_07_N |  | e-traveller not working error message: Microsoft JET Database Engine error '80040e14' Syntax error (missing operator) in query expression 'CONDITION. |  |  | Web support contacted. |

# Text boxes

Green Connect + Pump down /Blue :  Wait + Cool down	Ivory: Tests at Cold (36 hr)Red :  Wait + Warm up (12 hr)	  Dark Green: Disconnect (8 hr)