### CHOHAN /22-August-05

### Week 34 starting 22 Aug 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  22 August 05 : 8/22/2005 4:19:42 PM

| 7 | A1 | MBBL2213 | PT 5.1 1.5 KA Quench LF | 1.90 |
| --- | --- | --- | --- | --- |
| 1 | A2 | MBAL1084 | PT 12 Warm Up | 11.08 |
| 3 | B1 | LE-C11R7 | Thermal cycle - No Tcoils | 145.75 |
| 10 | B2 | MBAR3374 | ICS 4 Final connection | 295.62 |
| 4 | C1 | MBAL3377 | PT 6.3 Training 3 | 12.74 |
| 9 | C2 | MBBR3384 | ICS 4 Final connection | 295.18 |
| 11 | D1 | SSS245 | PT 13.2 Resist. Meas. | 297.59 |
| 12 | D2 | SSS607 | PT 12 Warm Up | 294.1 |
| 8 | E1 | MBAL2205 | Prep 5 Cool Down | 103.58 |
| 2 | E2 | MBAL2072 | PT 12 Warm Up | 262.94 |
| 6 | F1 | MBBR3387 | Prep 5 Cool Down | 103.58 |
| 5 | F2 | MBBL2200 | PT 6.1 Training 1 | 1.93 HV |

Tempr. from TBSM18 site

Coordinator this week: marco; MTM support 16-3602

CC on B1 since 3rd Aug; 1084 on A2 since 10th Aug

Generally very good week for Dipoles…particularly end of last week with fast turn-round & good perf.     ICS ran out of dipoles Sunday afternoon!

12 complete Magnet Tests cycles ( +1 not counted last week SSS607 TC)

Only 11 dipoles ( 0 with MM )+ 1 SSS(without MM & SSW )+1 ssss (607)  Thermal cycle previous week was not counted

Disconnected or ROBTC after tests last week ( Sat Midnight to Sat midnight) Last week: Many Thermal cycles (Ansaldo) ; SSS607 (Q4) on D2 MM tests last week; CC on B1 long cooling down… 1084 on A2

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Special  Points :

- Special SSS 606(Q5) on D2  (J Halik set up after Tuesday this week)

- 1084  Lifetime tests : 200A/s  4000 cycles  finished  & Thermal cycle

- Connect. Cryostat Tests B1  : not arrived at 1.9 K ? (JPLamboy, JP Tock ?)

## Tests

| Magnets under Test from 14-08-2005 to 20-08-2005 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBBL3138 |  |  | 11-08-2005 | 27 h. | 23 h. | 15 h. | 29 h. | 46 h. | 17-08-2005 | 140 h. | Stand By |  |  |
| 2 | TBB2 | MBAL2190 |  |  | 09-08-2005 | 53 h. | 36 h. | 18 h. | 22 h. | 38 h. | 16-08-2005 | 167 h. | Stand By |  |  |
| 3 | TBC1 | MBAL2532 |  |  | 11-08-2005 | 25 h. | 23 h. | 16 h. | 17 h. | 19 h. | 15-08-2005 | 100 h. | Stripping |  |  |
| 4 | TBC1 | MBAR3353 |  |  | 15-08-2005 | 37 h. | 23 h. | 20 h. | 15 h. | 6 h. | 19-08-2005 | 101 h. | Stripping |  |  |
| 5 | TBC2 | MBAL2182 |  |  | 13-08-2005 | 1 h. | 18 h. | 15 h. | 16 h. | 35 h. | 17-08-2005 | 85 h. | Stripping |  |  |
| 6 | TBD1 | SSS124 |  |  | 12-08-2005 | 37 h. | 14 h. | 24 h. | 10 h. | 27 h. | 17-08-2005 | 112 h. | Stand By |  |  |
| 7 | TBE1 | MBAL3106 |  |  | 13-08-2005 | 52 h. | 62 h. | 35 h. | 10 h. | 16 h. | 20-08-2005 | 175 h. | Stripping |  |  |
| 8 | TBE2 | MBAL2208 |  |  | 14-08-2005 | 37 h. | 20 h. | 26 h. | 14 h. | 12 h. | 19-08-2005 | 109 h. | Stripping |  |  |
| 9 | TBE2 | MBBL1191 |  |  | 08-08-2005 | 41 h. | 39 h. | 26 h. | 17 h. | 32 h. | 14-08-2005 | 155 h. | Stripping |  |  |
| 10 | TBF1 | MBBL3075 |  |  | 09-08-2005 | 41 h. | 33 h. | 19 h. | 12 h. | 15 h. | 14-08-2005 | 120 h. | Stripping |  |  |
| 11 | TBF1 | MBBR2220 |  |  | 14-08-2005 | 52 h. | 39 h. | 20 h. | 15 h. | 15 h. | 20-08-2005 | 141 h. | Stand By |  |  |
| 12 | TBF2 | MBAR3348 |  |  | 12-08-2005 | 48 h. | 21 h. | 76 h. | 16 h. | 8 h. | 19-08-2005 | 169 h. | Stripping |  |  |

### Also Finished 1153 on C2 & 1237 on B2  on Sunday

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBL3138 | 12272 A. (D1-L-S13) | 12716.7 A. (D2_L_S01) |  | - | - | - | - | - | - | - | - | - | - |
| 2 | MBAL2190 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 3 | MBAL2532 | 11348 A. (D1-L) | 12541 A. (D1-L) | 12739 A. (D1-L) | 0 A. (No Data) |  | - | - | - | - | - | - | - | - |
| 4 | MBAR3353 | 12171 A. (D2_L) | 12851 A. | - | - | - | - | - | - | - | - | - | - | - |
| 5 | MBAL2182 | 11794 A. (D1-L) | 12306 A. (D1-L) |  | - | - | - | - | - | - | - | - | - | - |
| 6 | SSS124 | 11716.5 A. (QI_12) | 12448 A. (QI_12) | 12603.7 A. (QI_34) |  | - | - | - | - | - | - | - | - | - |
| 7 | MBAL3106 | 9732 A. (D1_U) | 11988 A. (D1_U) | 12800 A. (D1-U) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 8 | MBAL2208 | 10982 A. (D1-L) | 11827 A. (D1_U) | 12392 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 9 | MBBL1191 | 12421 A. (D1-U) | 12617.8 A. (D2-L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 10 | MBBL3075 | 10850 A. (D1-U) | 12721.6 A. (D2_U) | 12850 A. (D2-L) |  | - | - | - | - | - | - | - | - | - |
| 11 | MBBR2220 | 11815 A. (D2-U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 12 | MBAR3348 | 11368 A. (D1_U_S13) | 11714 A. (D1_U_S01) | 12809 A. (D1-L-S01) | 12850 A. | 12850 A. | - | - | - | - | - | - | - | - | - |

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _16_M | TBF2 | Unable to launch MMP (shaft position error) while trying to perform MM1 on MB3348. | MTM |  | Maryline informed. She finds the sliding ruler for the position check (MM or Qloc) is wrongly adjusted |
| Tuesday _16_A | TBF2 | Over-range in sector0203 ap2 during MM1. | MTM |  | Guy asked to replace TRU2 by TRU4. |
| Tuesday _16_N | TBF2 | Over-range in sector 0001 ap2 during MM2. | MTM |  | Stephane informed Guy. TRU4 changed with TRU5. |
| Wednesday _17_N | All | Cryo OK lost in all cluster. | ACR |  |  |
| Wednesday _17_N | TBA2 | Cryo OK lost caused by Tleads fault. | ACR |  |  |
| Wednesday _17_N | TBC1 | Broken connector in the cable HVrack -PLCSE | MTM |  | HV test. delayed. Repaired by F.Flamand |
| Thursday_18_M | TBF2 | Panel on the top of F2 bench is not showing light when current is on | MTM |  |  |
| Thursday_18_N | TBD2 | SSS607: quench during load-line | AB/PO |  | Quemch occurred during a SPA because the threshold of the potaim card wasnot adapted for the Q4 (higher inductance). Threshold changed on the potaim card. |
| Thursday_18_N | TBE1 | MXI error : 3rd training on MB3106 could not be launched. | AB/CO |  | LF and HF crates were reset and tried launching again but same error Local server was logged out and logged in Hf and lf could arm. |
| Thursday_18_N | TBE1 | 3rd training on MB3106 performed: No hf file.. | AB/CO |  | Quench analysed using lf DAQ. |
| Thursday_18_N | TBE1 | E traveller was signed by ICS for standby/stripping without generating report for WP07 HV test. | ICS |  | Process stopped waiting for HV test result. |
| Friday_19_M | TBE2 | HF arming error on PT 6.4 in the MB3106. Magnet reached ultimate, but could not generate the u file. | AB/CO |  | Guy asked to perform the SPA without u file. He contacted A.Raimondo. He checked the crate and solved the problem. |
| Friday_19_M | TBF1 | MB2220: PT9 failed for the step QH versus Ground for 3050 V. | MTM |  | Investigation sequence launched. Then OK. |
| Friday_19_A | TBF | Broken HV connector between PLC HV and PLCSE Same problem Sunday night. | MTM |  | 2 hours to find a good contact. Repaired by F.Flamand. |
| Friday_19_A | TBA2 | Quench during special cycles in the 1084 magnet. | AB/CO |  | Giorgio found the mechanical switch open, he closed it |
| Friday_19_A | TBC2 | Application stuck in the bench. | AB/CO |  | Stop all and restart applications.. |
| Friday_19_N | TBD2 | Full load line on SSS607: Error due to power supply or cryo OK: test aborted after two positions. | ? |  | Guy came. Measurement was started again by launching the test for the remaining 4 position after modifying the run type. |
| Friday_19_N | TBD2 | Full load line on SSS607: test stop because of an error due to power supply or cryo OK. ( 2 times) | ? |  | Run recovery launched. |
| Friday_19_N | TBC2 | PT9 HV test was failing in the dipole-Quench- heater vs ground step at 3050V several times for MB1153. | ? |  | The test was successful after re-launching the test 2-3 times. |
| Friday_19_N | TBC2 | PT6.3 on MB1153: Hf acquisition file generated but not lf file | AB/CO | lf |  |
| Saturday_20_M | TBE2 | MB2072: leak in the manchette. | CRI |  | . |
| Saturday_20_A | TD1 | HV failed in the step Orbit Dipoles Vs GND. | MTM |  | Condensation on orbital dipole connectors. Heater fan is missing. Action P.Legrand? |
| Sunday _21_M | TBF | PLCSE cable contacts loose | MTM |  |  |
| Sunday _21_N | TBB2 | IAP @ 300K: lf heater discharge not OK. | MTM | lf | No magnet (MB3374) in the DB because HWREC was off for the special test on TBB1. Guy called. HWREC switched on. File could be generated. |
| Monday _22_M | TBA2 | HV test failed on dipole vs ground .Moisture was seen on top of CFB. | ACR |  | Action taken by ACR team |
| Monday _22_M | TBD1 | PT13 (resistance and voltage measurements) on SSS245: errors on -voltage of six Vtaps -voltage of correctors (EF circuit) -voltage of correctors (CD circuit) | MTM |  | Values different and correct as per template. Discrepancies coming from connection of power supply (-ve) to the meter -ve terminal which should be disconnected for the differential measurement reverse polarity connection (EF and CD) |

# Text boxes

Green Connect + Pump down /Blue :  Wait + Cool down	Ivory: Tests at Cold (36 hr)Red :  Wait + Warm up (12 hr)	  Dark Green: Disconnect (8 hr)