### G. CROCKFORD and E. VEYRUNES/19 sept 05

### Week 38 starting 19 Sept 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  19 Sept 05 : 9/19/2005 2:05:50 AM

| 11 | A1 | MBAL3399 | PT 13.2 Resist. Meas. | 298.31 |
| --- | --- | --- | --- | --- |
| 1 | A2 | MBBR1017 | PT 10 Min. NRJ 11850 A | 1.89 |
| 12 | B1 |  |  | 293.65 |
| 2 | B2 | SSS343 | PT 12 Warm Up | 53.15 |
| 10 | C1 | MBBL2214 | PT 13.2 Resist. Meas. | 299.11 |
| 6 | C2 | MBAR2230 | Prep 5 Cool Down | 93.26 |
| 7 | D1 | SSS180 | ICS 2 Connect Magnet (ICS) | 294.73 |
| 5 | D2 | SSS606 | PT 6.14 Training 14 | 2.35 |
| 4 | E1 | MBAL1201 | Prep 5 Cool Down | 91.4 |
| 9 | E2 | MBBL2201 | ICS 1 WP04 HV Test Warm | 295.8 |
| 8 | F1 | MBAL1250 | ICS 4 Final connection | 293.83 |
| 3 | F2 | MBAL3401 | PT 6.1 Training 1 | 1.90 |

Tempr. from TBSM18 site

Coordinator this week: John Miles;

MTM support Giorgio 16-3602

15 complete Magnet Tests cycles

Only 15 dipoles ( 0 with MM )+ 1 Thermal cycle

Disconnected or ROBTC after tests last week ( Sat Midnight to Sat midnight)

Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- B1 back to SSS? B2 (first SSS has been tested without any problem).

- Situation of Special SSS 606(Q5) on D2

- 1017 special test need equipment support help, because of test out of to do list.

# From 9/11/2005 to 9/17/2005

Tests Information

| Magnets under Test from 11-09-2005 to 17-09-2005 |
| --- |
| - | Bench | Magnet | Shafts | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |
| 1 | TBA1 | MBAL3391 |  | 10-09-2005 | 26 h. | 26 h. | 27 h. | 18 h. | 13 h. | 14-09-2005 | 110 h. | Stripping |
| 2 | TBA2 | MBAL3402 |  | 11-09-2005 | 22 h. | 26 h. | 21 h. | 15 h. | 11 h. | 15-09-2005 | 95 h. | Stripping |
| 3 | TBA2 | MBBR1256 |  | 07-09-2005 | 26 h. | 21 h. | 16 h. | 17 h. | 20 h. | 11-09-2005 | 100 h. | Stand By |
| 4 | TBB1 | MBAL1230(2) |  | 07-09-2005 | 40 h. | 22 h. | 33 h. | 18 h. | 0 h. | 12-09-2005 | 113 h. | Thermal cycle |
| 5 | TBB1 | MBAL1230(2) |  | 12-09-2005 | 0 h. | 16 h. | 12 h. | 20 h. | 20 h. | 15-09-2005 | 68 h. | Stripping |
| 6 | TBB2 | MBBL1197 |  | 10-09-2005 | 29 h. | 37 h. | 16 h. | 16 h. | 15 h. | 15-09-2005 | 113 h. | Stripping |
| 7 | TBC1 | MBBL1217 |  | 09-09-2005 | 54 h. | 25 h. | 15 h. | 16 h. | 15 h. | 14-09-2005 | 125 h. | Stripping |
| 8 | TBC2 | MBAL2138 |  | 10-09-2005 | 35 h. | 48 h. | 24 h. | 18 h. | 10 h. | 16-09-2005 | 135 h. | Stripping |
| 9 | TBE1 | MBAL1087 |  | 13-09-2005 | 24 h. | 21 h. | 16 h. | 16 h. | 11 h. | 16-09-2005 | 88 h. | Stripping |
| 10 | TBE1 | MBAL3392 |  | 08-09-2005 | 35 h. | 26 h. | 20 h. | 14 h. | 10 h. | 13-09-2005 | 105 h. | Stripping |
| 11 | TBE2 | MBBL2188 |  | 08-09-2005 | 46 h. | 26 h. | 29 h. | 17 h. | 10 h. | 13-09-2005 | 128 h. | Stripping |
| 12 | TBF1 | MBAL1242 |  | 13-09-2005 | 32 h. | 25 h. | 15 h. | 14 h. | 7 h. | 17-09-2005 | 93 h. | Stripping |
| 13 | TBF1 | MBAL2177 |  | 09-09-2005 | 29 h. | 21 h. | 20 h. | 16 h. | 10 h. | 13-09-2005 | 96 h. | Stand By |
| 14 | TBF2 | MBBL1239 |  | 07-09-2005 | 31 h. | 31 h. | 16 h. | 14 h. | 18 h. | 11-09-2005 | 110 h. | Stripping |
| 15 | TBF2 | MBBL2143(2) |  | 11-09-2005 | 29 h. | 39 h. | 11 h. | 15 h. | 13 h. | 16-09-2005 | 107 h. | Stripping |

Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAL3391 | 9841.9 A. (D1_L) | 11763.6 A. (D1_L) | 11756 A. (D1-U) | 12801 A. (D2-U) | 12850 A. | - | - | - | - | - |
| 2 | MBAL3402 | 12014 A. (D1 - U) | 12185 A. (D2-U) | 12467 A. (D2_U) |  | - | - | - | - | - | - |
| 3 | MBBR1256 | 12538 A. (D1_L) | 12764 A. (D1-L) | 12670 A. (D1-L) |  | - | - | - | - | - | - |
| 4 | MBAL1230 | 10706 A. (D1 - U) | 11414.3 A. (D2_L) | 12050 A. (D1-L) | 12102 A. (D1-L) | 12350 A. (D1-L) | 12459 A. (D2-L) | 12616.9 A. (D2_U) | 12690.4 A. (D2_L) | 12710 A. (D2-U) |  |
| 5 | MBAL1230 | 12645 A. (D2-L) | 12836 A. (D2_L) |  | - | - | - | - | - | - | - |
| 6 | MBBL1197 | 11849 A. (D1-L) | 11933.5 A. (D1_U) | 12301 A. (D1-L) | 12687 A. (D2-L) |  | - | - | - | - | - |
| 7 | MBBL1217 | 12208 A. (D1_U) | 12368.8 A. (D2_U) | 12623.3 A. (D1_U) |  | - | - | - | - | - | - |
| 8 | MBAL2138 | 11319.4 A. (D1_U) | 12230 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - |
| 9 | MBAL1087 | 12547 A. (D2-U) | 12408 A. (D1-L) | 12850 A. | - | - | - | - | - | - | - |
| 10 | MBAL3392 | 8430 A. (D2-U) | 11460 A. (D2_U) | 12262 A. (D2-U) | 12777 A. (D1-U) |  | - | - | - | - | - |
| 11 | MBBL2188 | 10064 A. (D2_U) | 12225 A. (D1-L) | 12754 A. (D1-L) |  | - | - | - | - | - | - |
| 12 | MBAL1242 | 11613 A. (D2-U) | 12257 A. (D2-U) | 12574.4 A. (D1_L) |  | - | - | - | - | - | - |
| 13 | MBAL2177 | 10986 A. (D2-U) | 12241.2 A. (D1-U) | 11985.6 A. (D2-U) | 12633.6 A. (D2_U) |  | - | - | - | - | - |
| 14 | MBBL1239 | 11786.1 A. (D1-L) | 12128.5 A. (D1-L) | 12485.7 A. (D2_U) |  | - | - | - | - | - | - |
| 15 | MBBL2143 | 12295 A. (D2-L) | 12850 A. | - | - | - | - | - | - | - | - |

| Date | Bench | Problem | Type | Application/soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_13_A | TBB2 | On MB1197 SPA occurred after 10 min at 12000 A caused by a loss of cryo OK. | ACR |  | De-excitation test repeated. |
| Wednesday _14_M | TBA2 | Quench trigger occurred when the rotary switch was moved to CC position without stopping HF & LF acquisitions | MTM |  |  |
| Thursday_15_A | TBE1 | After the Min NRJ on 1087 no SMS was received . Magnet was cooled down again. | MTM/ACR |  | Magnet warm up with 4.5 K quench |
| Thursday_15_N | TBE2 | In TEMA the magnet name for the MB 2217 is not displayed “error” in place. | TS/CE? | MTF? | Solved by Richard Friday morning. |
| Thursday_15_N | TBE2 | Loss of cryo OK. | ACR |  |  |
| Friday_16_N | TBB | HWREC switched off by Richard. | MTM | HWREC |  |
| Friday_16_N | TBE2 | AQA warnings for V-taps in training quenches (MB2217) | MTM | AQA | Richard advised to keep the magnet cold for further analysis. |
| Sunday_18_M | TBF | On F cluster problem with profibus cable for PLC-HV still persists we used XMT.MY95 | MTM | HV | Change the rack |
| Sunday_18_A | TBA2 | Problem was that the test was configured for the previous version of MMP. | MTM | MMP | Richard has corrected both "miniloadline" and "miniloadline+Joint" measurements with an extension "_MMP660". |