### Chohan /24-Oct-05

### Week 43 starting 24Oct 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  24Oct 05 :  10/24/2005 1:08:58 PM

| 4 | A1 | MBAR1275 | PTE Min. NRJ 3KA Diode Test | 4.19 HV |
| --- | --- | --- | --- | --- |
| 1 | A2 | MBAL1154 | ICS 5 Traveller Last | 296.43 |
| 6 | B1 | SSS638 | PT 1 IAP @ cold | 1.90 |
| 10 | B2 | SSS162 | Prep 3 CDW1 HV | 295.8 |
| 7 | C1 | MBAL3139 | New Magnet arrived | 294.64 |
| 8 | C2 | MBBL3405 | Prep 5 Cool Down | 95.7 |
| 3 | D1 | SSS292 | PT 8 Fast De-excitation | 1.90 |
| 11 | D2 | SSS604 | PT 12 Warm Up | 295 |
| 2 | E1 | MBBL3404 | PT 12 Warm Up | 280.18 |
| 12 | E2 | CFB_LEAK | Cryo leak repair | 298.04 |
| 5 | F1 | MBAR2236 | PT 13.2 Resist. Meas. | 300.82 |
| 9 | F2 | MBAL2210 | Prep 5 Cool Down | 280.36 |

Coordinator this week: stephane; MTM support: 16-3602

Only 14 complete Magnet Tests cycles  from 11 benches ( E2 still under CFB repair)   10 Dipoles + 3  SSS  + 1 SSSS

Disconnected or ROBTC after tests last week ( Sat Midnight to Sat midnight)

Situation From Cryogenics:  J-P Lamboy or ACR Rep.

No new magnet tested on E2 since 29 Sept (already 3 weeks !)

E2 CFB (leak) Repaired  & restart usage ?

Special  Points :

- Special SSS on B1 :  Roy

- Special SSS631  on D2 :   MM Chaconsa or not?

- 2236 HV failure  & investigation ?

AOB

| Magnets under Test from 16-10-2005 to 22-10-2005 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBAL2207 |  |  | 12-10-2005 | 30 h. | 28 h. | 17 h. | 15 h. | 19 h. | 16-10-2005 | 109 h. | Stripping |  |  |
| 2 | TBA1 | MBBR1259 |  |  | 16-10-2005 | 37 h. | 34 h. | 13 h. | 12 h. | 23 h. | 21-10-2005 | 119 h. | Stripping |  |  |
| 3 | TBA2 | MBAR1265 |  |  | 13-10-2005 | 32 h. | 21 h. | 54 h. | 16 h. | 6 h. | 19-10-2005 | 129 h. | Stripping |  |  |
| 4 | TBB1 | SSS654 |  |  | 15-10-2005 | 74 h. | 29 h. | 10 h. | 16 h. | 14 h. | 20-10-2005 | 143 h. | Stand By |  |  |
| 5 | TBB2 | SSS049 |  |  | 15-10-2005 | 59 h. | 13 h. | 16 h. | 14 h. | 12 h. | 20-10-2005 | 114 h. | Stripping |  |  |
| 6 | TBC1 | MBAL2192 |  |  | 16-10-2005 | 36 h. | 24 h. | 34 h. | 28 h. | 0 h. | 21-10-2005 | 122 h. | Not Found or Thermal cycle |  |  |
| 7 | TBC1 | MBBR1278 |  |  | 11-10-2005 | 43 h. | 22 h. | 28 h. | 15 h. | 16 h. | 16-10-2005 | 124 h. | Stand By |  |  |
| 8 | TBC2 | MBAL1070 |  |  | 18-10-2005 | 32 h. | 8 h. | 27 h. | 17 h. | 10 h. | 22-10-2005 | 94 h. | Stripping |  |  |
| 9 | TBC2 | MBAR2222 |  |  | 14-10-2005 | 38 h. | 26 h. | 16 h. | 17 h. | 8 h. | 18-10-2005 | 105 h. | Stripping |  |  |
| 10 | TBD1 | SSS163 |  |  | 12-10-2005 | 34 h. | 17 h. | 45 h. | 7 h. | 10 h. | 16-10-2005 | 113 h. | Stripping |  |  |
| 11 | TBD1 | SSS292 |  |  | 16-10-2005 | 49 h. | 26 h. | 60 h. | - | - | 22-10-2005 | 135 h. | Not Found or Thermal cycle |  |  |
| 12 | TBE1 | MBBL3411 |  |  | 15-10-2005 | 32 h. | 22 h. | 18 h. | 17 h. | 10 h. | 19-10-2005 | 99 h. | Stripping |  |  |
| 13 | TBF1 | MBAR1268 |  |  | 14-10-2005 | 26 h. | 37 h. | 33 h. | 18 h. | 13 h. | 19-10-2005 | 127 h. | Stripping |  |  |
| 14 | TBF2 | MBBL2055 |  |  | 14-10-2005 | 41 h. | 21 h. | 31 h. | 15 h. | 7 h. | 18-10-2005 | 115 h. | Stripping |  |  |

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAL2207 | 10855 A. (D2_L) | 12572.9 A. (D1_U) | 12830.7 A. (D2_U) |  | - | - | - | - | - | - | - | - | - |
| 2 | MBBR1259 | 12640.7 A. (D2-U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 3 | MBAR1265 | 11318.8 A. (D2_U_S01) | 11842.8 A. (D1_L_S01) | 12461.7 A. (D2_L) | 12786.2 A. (D2-L-S13) | 12850 A. | - | - | - | - | - | - | - | - |
| 4 | SSS654 | 5820 A. | - | - | - | - | - | - | - | - | - | - | - | - |
| 5 | SSS049 | 12223.2 A. (QE_12) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 6 | MBAL2192 | 10734.4 A. (D2_L) | 11761.9 A. (D1_L) | 12103.7 A. (D2-U) | 12276.7 A. (D1-U) | 12329.9 A. (D1-L) | 12461.8 A. (D1_L) | 12568.8 A. (D2_U) | 11971 A. (D2_U) | 12742.9 A. (D1_L) |  | - | - | - |
| 7 | MBBR1278 | 11906 A. (D1_l) | 12578 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 8 | MBAL1070 | 11959 A. (D2_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 9 | MBAR2222 | 12466.1 A. (D1-L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 10 | SSS163 | 12244.7 A. (QE_12) | 0 A. (No Data) | 12366 A. (QI-12) | 12710 A. (QI-12) |  | - | - | - | - | - | - | - | - |
| 11 | SSS292 | 11490.2 A. (QI_12) | 11914.2 A. (QE_34) | 12162 A. (QI_12) | 12263 A. (QE_34) | 12358.3 A. (QI-34) | 12506 A. (QI-12) | 12566 A. (QE-12) | 12678.9 A. (QE_12) | 12795.4 A. (QI-12) |  | - | - | - |
| 12 | MBBL3411 | 11834.1 A. (D2-U) | 12727 A. (D1-L) | 12857 A. | - | - | - | - | - | - | - | - | - | - |
| 13 | MBAR1268 | 11715.6 A. (D2_U) | 12108.8 A. (D1-L) | 12692.1 A. (D2_L) |  | - | - | - | - | - | - | - | - | - |
| 14 | MBBL2055 | 9912 A. (D2_U) | 11765.6 A. (D2-U) | 12267.7 A. (D1-U) | 12684.2 A. (D2_U) | 12847.7 A. (D1_L) |  | - | - | - | - | - | - | - |

Shows  SSS 163 and 292(both on   D1) and long COLD TESTS Times due to Cohabitation with SpecialSSS on D2

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_18_N | TBC2 | Mobic 2 displays a message “MTF temporarily unavailable”. Scan of the MBAL1298 performed correctly with MOBIC 1 | MTF |  | Network not available last week? |
| Friday_21_A | TBB2 | New version of IAP installed in local server did not show the VT readings. It was working properly through TEMA in the control room. | AB/CO | IAP | Giorgio was called. He davised to perform the analysis from control room. He sent the info to Boris Khomenkho Monday. Problem solved. |
| Friday_21_N | TBD1 | SSS292: Alarm on anticryostat ap 2. The short anticryostat of AP2 (aperture 2) is at 6 C. SSW return wire was trapped inside the anticryostat. | MTM |  | Giorgio informed. Salem increased the flow of N2. After investigation Giorgio concluded that ICS did not install the gamelle and MLI |
| Friday_21_N | All | Unable to cool down or warm up magnets in the gaseous phase because the supervision of the HP compressor is not working (software blocked) | ACR |  | 11 hours lost. |
| Saturday_22_A | TBB1 | Warm up quench on SSS150 : ALERTS : QH1 (Imax<50, TauIniV>117, TauIniI>120, Rinit>16.5) | MTM | AQA | Giorgio checked the results and asked QH resistance measurements: OK Magnet cooled down again and warm up quench repeated. When quench analysed from control room it is OK. |
| Saturday_22_N | TBA2 | Cryo loss during HV test on MB1154. HV test (CDAP) failing at external and internal bus bar vs ground | MTM |  | 7 hours of stop. Related? |
| Sunday_23_M | TBF1 | HV test dipole vs ground failed on MB2236. | MTM |  | Pierre will ask the mechanics to disconnect the current leads to measure the resistance magnet/ground. Resistance between GND and current lead after disconnection of the leads is 0.1 ohm. Pierre advised to warm up the magnet slowly. Problem confirmed monday after disconnection. It came from the magnet. |
| Monday_24_M |  | MOBIC1 not working . USB driver error. | MTM |  | Rest done by Richard then OK. |
| Monday_24_M | TBD1 | Values of TF measured with SSW are wrong. | MTM |  | Richard called. SSW The connector from the return wire was damaged and it was repaired by Fabienne |
| Monday_24_A | TBD1 | XY values not correct because there is a cold point in the short anti-cryostat. | MTM |  | Nikolay asked to dry the ttwo aperture with a blower and repeat the measurements fast. XY measurements are suppresses by Stephane. Only transfer function at one current remained. |