### Week 11 starting 13 march 06

### SM18 Operation Meeting (Tuesdays) 14 March 06

Bench status  on Monday  13-03- 2006 at: 3/13/2006 9:03:27 AM

| 12 | A1 | MBBR2286 | Prep 5 Cool Down | 195.52 |
| --- | --- | --- | --- | --- |
| 8 | A2 | MBBR1310 | PT 14 Magnet MAPS | 298.04 |
| 7 | B1 | SSS038 | Extra min. Energy test | 2.35 |
| 5 | B2 | SSS248 | PT 13.2 Resist. Meas. | 298.04 |
| 4 | C1 | SSS227 | PT 13.2 Resist. Meas. | 296.88 |
| 10 | C2 | SSS325 | PT 1 IAP @ cold | 2.01 HV |
| 2 | D1 | SSS621 | Stretched Wire Cold SC11 | 2.35 |
| 9 | D2 | SSS637 | PT 2 HV Cold Test | 1.90 |
| 6 | E1 | MBAR1325 | PT 12 Warm Up | 256.36 |
| 1 | E2 | MBAR2245 | Prep 3 CDW1 HV | 297.86 |
| 3 | F1 | MBBR2256 | PT 14 Magnet MAPS | 297.41 |
| 11 | F2 | MBAR1327 | Prep 5 Cool Down | 88.41 |

Tempr. from TBSM18 site

Supervisor:Vittorio;           Eq support : 16-3602 :

18 (17)  Magnet Tests completed :  9 Dipoles (2 with MM both  were  REPEATS)

: 6 Arc-SSS,

: 3 IR-SSS  (610 2nd run, 656  cryo leak but took 4 days on bench)

General Points :

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- Situation from (SSS, SSSS)  :  J.Halik

- PC’s in Control Room & Usage

- Special SSSS issues if any:

A.O.B.

| Magnets under Test from 05-03-2006 to 11-03-2006 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBAL1092 |  |  | 08-03-2006 | 19 h. | 20 h. | 12 h. | 18 h. | 8 h. | 11-03-2006 | 77 h. | Stand By |  | HV@w |
| 2 | TBA1 | MBBR2299 |  |  | 03-03-2006 | 31 h. | 21 h. | 31 h. | 17 h. | 10 h. | 08-03-2006 | 110 h. | Stripping |  |  |
| 3 | TBA2 | MBBL2142 |  |  | 25-02-2006 | 25 h. | 22 h. | 113 h. | 16 h. | 39 h. | 06-03-2006 | 215 h. | Stripping |  |  |
| 4 | TBA2 | MBBR2297 |  |  | 06-03-2006 | 33 h. | 23 h. | 15 h. | 15 h. | 12 h. | 10-03-2006 | 98 h. | Stripping |  |  |
| 5 | TBB1 | SSS216 |  |  | 03-03-2006 | 21 h. | 14 h. | 37 h. | 6 h. | 22 h. | 07-03-2006 | 100 h. | Stripping |  |  |
| 6 | TBB2 | SSS085 |  |  | 04-03-2006 | 57 h. | 13 h. | 19 h. | 9 h. | 16 h. | 09-03-2006 | 114 h. | Stripping |  |  |
| 7 | TBC1 | SSS187 |  |  | 02-03-2006 | 23 h. | 25 h. | 19 h. | 7 h. | 16 h. | 06-03-2006 | 90 h. | Stripping |  |  |
| 8 | TBC1 | SSS251 |  |  | 06-03-2006 | 26 h. | 21 h. | 17 h. | 8 h. | 11 h. | 09-03-2006 | 83 h. | Stripping |  |  |
| 9 | TBC2 | SSS168 |  |  | 01-03-2006 | 19 h. | 12 h. | 34 h. | 7 h. | 17 h. | 06-03-2006 | 89 h. | Stripping |  |  |
| 10 | TBC2 | SSS237 |  |  | 06-03-2006 | 35 h. | 23 h. | 19 h. | 13 h. | 26 h. | 11-03-2006 | 116 h. | Stand By |  | Wp14 |
| 11 | TBD1 | SSS610 |  |  | 02-03-2006 | 8 h. | 53 h. | 4 h. | 11 h. | 65 h. | 08-03-2006 | 141 h. | Stand By |  |  |
| 12 | TBD2 | SSS656 |  |  | 06-03-2006 | 28 h. | - | - | - | - | 10-03-2006 | 96 h. | Stand By |  | Leak |
| 13 | TBD2 | SSS691 |  |  | 28-02-2006 | 49 h. | 24 h. | 15 h. | 13 h. | 46 h. | 06-03-2006 | 147 h. | Stand By |  |  |
| 14 | TBE1 | MBAR1321 |  |  | 01-03-2006 | 38 h. | 22 h. | 16 h. | 14 h. | 19 h. | 06-03-2006 | 109 h. | Stripping |  |  |
| 15 | TBE1 | MBBR2292 |  |  | 06-03-2006 | 23 h. | 22 h. | 15 h. | 16 h. | 19 h. | 10-03-2006 | 95 h. | Stripping |  |  |
| 16 | TBE2 | MBBR1309 |  |  | 04-03-2006 | 30 h. | 24 h. | 18 h. | 17 h. | 8 h. | 08-03-2006 | 97 h. | Stripping |  |  |
| 17 | TBF1 | MBAR2234 |  |  | 04-03-2006 | 25 h. | 22 h. | 22 h. | 14 h. | 12 h. | 08-03-2006 | 95 h. | Stand By |  | ? |
| 18 | TBF2 | MBAR2247 |  |  | 04-03-2006 | 30 h. | 37 h. | 54 h. | 21 h. | 18 h. | 11-03-2006 | 160 h. | Stripping |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAL1092 | 12264 A. (D1-L) | 11456 A. (No Data) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 2 | MBBR2299 | 10879.7 A. (D2-L) | 11711.4 A. (D1-L) | 11890 A. (D2-U) | 12314.9 A. (D2-U) | 12411.4 A. (D2_U) | 12601.7 A. (D2_U) | 12850 A. | - | - | - | - | - | - |
| 3 | MBBL2142 | 12476 A. (D1-L-S07) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 4 | MBBR2297 | 10320 A. (D2-U) | 12403 A. (D1-L) | 12783.1 A. (D2_U) |  | - | - | - | - | - | - | - | - | - |
| 5 | SSS216 | 11711.1 A. (QE_34) | 12177.7 A. (QI_12) | 0 A. (QE_12) |  | - | - | - | - | - | - | - | - | - |
| 6 | SSS085 | 12198.9 A. (QI_12) | 12293 A. (QE_12) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 7 | SSS187 | 11365 A. (QI-34) | 12507 A. (QE-12) | 12650 A. (QE_34) | 12791 A. (QI-12) |  | - | - | - | - | - | - | - | - |
| 8 | SSS251 | 11422.1 A. (QE_12) | 12027 A. (QI_12) | 12132.6 A. (QE_12) |  | - | - | - | - | - | - | - | - | - |
| 9 | SSS168 | 11687.2 A. (QI_34) | 11928.7 A. (QI_34) | 12079 A. (QI_34) | 12178 A. (QE_12) | 12431.3 A. (QE_34) | 12714.3 A. (QI_34) | 12837.5 A. (QE_12) | 12850 A. | - | - | - | - | - |
| 10 | SSS237 | 12166 A. (QE-12) | 12665.4 A. (QI_34) |  | - | - | - | - | - | - | - | - | - | - |
| 11 | SSS610 | 3998.6 A. (Int_MQM) | 4259.1 A. (Ext_MQM) | 4514.3 A. (Ext-MQM) | 4571.9 A. (Ext-MQM) | 4609.2 A. (Int_MQM) | 4645.7 A. (Ext_MQM) | 4650 A. | 4650 A. | - | - | - | - | - | - |
| 12 | SSS656 |  | - | - | - | - | - | - | - | - | - | - | - | - |  |
| 13 | SSS691 | 5779 A. (MQM-I) | 5820 A. | - | - | - | - | - | - | - | - | - | - | - |  |
| 14 | MBAR1321 | 11494 A. (D1- U) | 12469.8 A. (D2_U) | 12611.9 A. (D1_L) |  | - | - | - | - | - | - | - | - | - |  |
| 15 | MBBR2292 | 11578.7 A. (D1_U) | 12316.9 A. (D1_U) | 12747.4 A. (D2_U) |  | - | - | - | - | - | - | - | - | - |  |
| 16 | MBBR1309 | 11714 A. (D2-U) | 12209 A. (D2-U) | 12680 A. (D2_U) |  | - | - | - | - | - | - | - | - | - |  |
| 17 | MBAR2234 | 11088.9 A. (D1_U) | 11772 A. (D1_L) | 12830.7 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - | - | - |  |
| 18 | MBAR2247 | 12529 A. (D2_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |  |

MM on 2142 , 2247      +      Usual problems of  D  Cluster in Disconnect ( or connect !) : 610, 691 ;   Cryo Leak on  656