### Week 13 starting 27 march 06

### SM18 Operation Meeting (Tuesdays) 28 March 06

Bench status  on Monday  27-03-2006 at: 3/27/2006 7:27:09 PM

| 3 | A1 | MBAR2251 | PT 3 QH Measure | 1.90 |
| --- | --- | --- | --- | --- |
| 12 | A2 | MBAR1334 | PT 5.2 1.5 KA Quench HF | 1.90 |
| 5 | B1 | SSS657 | PT 13.2 Resist. Meas. | 296.88 |
| 1 | B2 | SSS191 | PT 13.2 Resist. Meas. | 299.65 |
| 9 | C1 | SSS239 | PT 14 Magnet MAPS | 298.58 |
| 11 | C2 | SSS328 | Prep 3 CDW1 HV | 299.11 |
| 4 | D1 | SSS620 | Temporary test and preparation | 2.35 |
| 10 | D2 |  |  | 299.65 |
| 7 | E1 | MBBR2240 | ICS 1 WP04 HV Test Warm | 301.35 |
| 8 | E2 | MBAR2237 | PT 6.2 Training 2 | 1.90 |
| 6 | F1 | MBBL1097 | Prep 5 Cool Down | 4.38 HV |
| 2 | F2 | MBAL1162 | PT 12 Warm Up | 86.47 |

Tempr. from TBSM18 site

Supervisor: M. BUZIO;           Eq support : 16-3602 : G. DEFERNE

15(16) Magnet Tests completed :  9 Dipoles (1310  REPEAT +2278 ROBTC)

: 4 Arc-SSS (191  REPEAT)

: 2 IR-SSS

General Points :

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- Situation from (SSS, SSSS)  :  J.Halik

- QH error (Imax, Vinit) on during (IAP, Training Quench).

- Water leak (ICS kitchen, cluster C) CALL “Giorgio” or “77777”?

A.O.B.

## Tests

| Magnets under Test from 19-03-2006 to 25-03-2006 |
| --- |
| - | Bench | Magnet | Shafts | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |
| 1 | TBA1 | MBBL1109 |  | 15-03-2006 | 25 h. | 42 h. | 45 h. | 20 h. | 11 h. | 21-03-2006 | 143 h. | Stripping |
| 2 | TBA2 | MBBR1284 |  | 19-03-2006 | 18 h. | 20 h. | 25 h. | 18 h. | 12 h. | 23-03-2006 | 93 h. | Stripping |
| 3 | TBA2 | MBBR1310(2) |  | 13-03-2006 | 23 h. | 27 h. | 69 h. | 16 h. | 10 h. | 19-03-2006 | 145 h. | Stripping |
| 4 | TBB1 | SSS504 |  | 17-03-2006 | 3 h. | 43 h. | 23 h. | 12 h. | 26 h. | 22-03-2006 | 107 h. | Stand By |
| 5 | TBB2 | SSS094 |  | 18-03-2006 | 27 h. | 15 h. | 18 h. | 13 h. | 9 h. | 21-03-2006 | 82 h. | Stripping |
| 6 | TBB2 | SSS191(2) |  | 21-03-2006 | 27 h. | 17 h. | 56 h. | 3 h. | 0 h. | 25-03-2006 | 103 h. | Thermal cycle |
| 7 | TBC1 | SSS233 |  | 18-03-2006 | 26 h. | 18 h. | 55 h. | 5 h. | 18 h. | 23-03-2006 | 122 h. | Stripping |
| 8 | TBC2 | SSS183 |  | 17-03-2006 | 24 h. | 13 h. | 59 h. | 8 h. | 23 h. | 22-03-2006 | 127 h. | Stripping |
| 9 | TBD2 | SSS611 |  | 15-03-2006 | 35 h. | 23 h. | 20 h. | 11 h. | 40 h. | 21-03-2006 | 129 h. | Stand By |
| 10 | TBE1 | MBBR2288 |  | 18-03-2006 | 19 h. | 24 h. | 41 h. | 16 h. | 10 h. | 23-03-2006 | 110 h. | Stripping |
| 11 | TBE2 | MBBR2278(2) |  | 17-03-2006 | 20 h. | 19 h. | 58 h. | 16 h. | 0 h. | 22-03-2006 | 113 h. | Thermal cycle |
| 12 | TBE2 | MBBR2278(2) |  | 22-03-2006 | 0 h. | 18 h. | 19 h. | 15 h. | 14 h. | 25-03-2006 | 66 h. | Stripping |
| 13 | TBF1 | MBAR1319 |  | 21-03-2006 | 26 h. | 19 h. | 32 h. | 17 h. | 7 h. | 25-03-2006 | 101 h. | Stripping |
| 14 | TBF1 | MBBR1279 |  | 17-03-2006 | 24 h. | 23 h. | 21 h. | 15 h. | 10 h. | 21-03-2006 | 93 h. | Stripping |
| 15 | TBF2 | MBAL1064 |  | 20-03-2006 | 23 h. | 23 h. | 18 h. | 18 h. | 17 h. | 24-03-2006 | 99 h. | Stand By |
| 16 | TBF2 | MBAR1328 |  | 15-03-2006 | 23 h. | 23 h. | 33 h. | 18 h. | 14 h. | 20-03-2006 | 111 h. | Stripping |

Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBL1109 | 10874.6 A. (D1-U) | 11661.5 A. (D2-L) | 12033.7 A. (D2_L) | 12320.7 A. (D2_L) | 12526.7 A. (D2-U) | 12591.4 A. (D1-L) | 12836.1 A. (D2-L) | 12850 A. | - | - |
| 2 | MBBR1284 | 12725.4 A. (D2_L) | 12515.7 A. (D2-L) | 12850 A. | - | - | - | - | - | - | - |
| 3 | MBBR1310 | 11698.2 A. (D2-U-S01) | 12196.2 A. (D1_U) | 12670 A. (D1_U) |  | - | - | - | - | - | - |
| 4 | SSS504 | 39.8 A. (N.a.) | 12189.8 A. (QI_12) | 12850 A. | - | - | - | - | - | - | - |
| 5 | SSS094 | 12850 A. | - | - | - | - | - | - | - | - | - |
| 6 | SSS191 | 11318.2 A. (QE_34) | 11860 A. (QI_12) | 12050 A. (QI_34) | 12216.3 A. (QE-34) | 12388.8 A. (QE_34) | 12472.7 A. (QI_34) | 12573 A. (QI_12) | 12638 A. (QE-12) | 12720 A. (QE_34) |  |
| 7 | SSS233 | 12012.8 A. (QI_12) | 12851.3 A. (QE-12) | 12850 A. | - | - | - | - | - | - | - |
| 8 | SSS183 | 12094 A. (QE_34) | 12501.5 A. (QI_12) | 12608.2 A. (QI_12) |  | - | - | - | - | - | - |
| 9 | SSS611 | 5110.8 A. (EXT_MQM2) | 5820 A. | - | - | - | - | - | - | - | - |
| 10 | MBBR2288 | 11596.1 A. (D2_U) | 12131.4 A. (D2_L) | 12386.2 A. (D2-L) | 12739.6 A. (D2-U) | 12850 A. | - | - | - | - | - |
| 11 | MBBR2278 | 9749 A. (D2_U) | 10390.8 A. (D1-U) | 11290.7 A. (D2-U) | 11641.4 A. (D1-L) | 11885.8 A. (D2_L) | 12318 A. (D1_U) | 12473.6 A. (D2_L) | 12681.9 A. (D2_U) | 12851 A. (D1-U) |  |
| 12 | MBBR2278 | 11473 A. (D2_L) | 12181.6 A. (D1_L) | 12337.6 A. (D2_L) |  | - | - | - | - | - | - |
| 13 | MBAR1319 | 12247.5 A. (D1_L) | 12832.5 A. (D1_L) | 12806.4 A. (D1_U) | 12378.1 A. (D1_U) | 12850 A. | - | - | - | - | - |
| 14 | MBBR1279 | 12164.2 A. (D1_U) | 12388.3 A. (D2_U) | 12554 A. (D1_L) |  | - | - | - | - | - | - |
| 15 | MBAL1064 | 12638 A. (D2_U) | 12701.9 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - |
| 16 | MBAR1328 | 10275.9 A. (D2_U) | 11735.4 A. (D1-L) | 12209 A. (D1_L) | 12387.6 A. (D1_L) | 12662.4 A. (D2_U) | 12845 A. (D1_L) | 12850 A. | - | - | - |

Long training for 	Dipole :2278, 1109.

SSS:191

MM for Dipole 1310

Rotary Switch problem on Cluster C for S183 (<4Q), S233 (<3 Q) around cold test 58 hours.