### Week 20 starting 15 May 06

### SM18 Operation Meeting (Tuesdays) 16 May 06

Bench status  on Monday  15-05-2006 at: 5/15/2006 10:54:44 AM

| 1 | A1 | MBAR2236 | PT 7 Magnet Trained | 4.95 |
| --- | --- | --- | --- | --- |
| 7 | A2 |  |  | 298.76 |
| 4 | B1 | SSS334 | PT 14 Magnet MAPS | 298.58 |
| 2 | B2 | SSS524 | PT 8 Fast De-excitation | 4.30 HV |
| 8 | C1 | MBAR1274 | ICS 4 Final connection | 296.97 |
| 5 | C2 | MBBR1282 | PT 5.2 1.5 KA Quench HF | 4.63 HV |
| 11 | D1 | SSS635 | PT 7 Magnet Trained | 4.55 HV |
| 3 | D2 | SSS658 | PT 12 Warm Up | 174.94 |
| 9 | E1 | MBBR2318 | Prep 5 Cool Down | 262.41 |
| 6 | E2 | MBAR1293 | PT 14 Magnet MAPS | 300.1 |
| 10 | F1 | MBBR2261 | Prep 5 Cool Down | 269.6 |
| 12 | F2 | MBBR2257 | PT 14 Magnet MAPS | 298.58 |

CONFIG :

D8- AS2 –IRSSS2;

Supervisor: gijs-gilbert/        Eq support : 16-3602 :

16 Magnet Tests completed :  11 Dipoles (),    2 Arc-SSS  ( no   REPEAT)

(15 Magnets in all)		1  500-series ( 504 : 2 runs /pressure tests/leaks)

1  600-series ( 609 , , 609 finished prev week  but not disconnected &

removed till this week : total 307 hrs )

{ Long MM on 2254 ( 241 hrs = twice w.o. MM) }

General Points :

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- News /Situation from (SSS, SSSS)  :  J Halik

- Situation SSSS  :OP issues :

- A.O.B.

| Magnets under Test from 07-05-2006 to 13-05-2006 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBBL1251 |  |  | 01-05-2006 | 61 h. | 20 h. | 20 h. | 14 h. | 16 h. | 07-05-2006 | 131 h. | Stripping |  |  |
| 2 | TBA1 | MBBL1391 |  |  | 07-05-2006 | 25 h. | 21 h. | 20 h. | 15 h. | 12 h. | 11-05-2006 | 93 h. | Stripping |  |  |
| 3 | TBA2 | MBAL1385 |  |  | 06-05-2006 | 30 h. | 25 h. | 18 h. | 18 h. | 9 h. | 10-05-2006 | 100 h. | Stripping |  |  |
| 4 | TBB1 | SSS229 |  |  | 07-05-2006 | 29 h. | 18 h. | 23 h. | 9 h. | 19 h. | 11-05-2006 | 98 h. | Stripping |  |  |
| 5 | TBB1 | SSS333 |  |  | 02-05-2006 | 44 h. | 13 h. | 16 h. | 8 h. | 17 h. | 07-05-2006 | 98 h. | Stripping |  |  |
| 6 | TBB2 | SSS504 |  |  | 05-05-2006 | 32 h. | 18 h. | 39 h. | - | - | 10-05-2006 | 89 h. | Not Found or Thermal cycle |  |  |
| 7 | TBB2 | SSS504 |  |  | 10-05-2006 | 0 h. | 18 h. | 5 h. | 8 h. | 11 h. | 12-05-2006 | 42 h. | Stand By |  |  |
| 8 | TBC1 | MBAR2322 |  |  | 04-05-2006 | 18 h. | 25 h. | 18 h. | 15 h. | 7 h. | 08-05-2006 | 83 h. | Stripping |  |  |
| 9 | TBC2 | MBBR1276 |  |  | 08-05-2006 | 22 h. | 21 h. | 21 h. | 16 h. | 19 h. | 13-05-2006 | 99 h. | Stripping |  |  |
| 10 | TBC2 | MBBR1280 |  |  | 03-05-2006 | 31 h. | 23 h. | 21 h. | 17 h. | 10 h. | 08-05-2006 | 102 h. | Stripping |  |  |
| 11 | TBD2 | SSS609 |  |  | 27-04-2006 | 83 h. | 25 h. | 69 h. | 15 h. | 115 h. | 09-05-2006 | 307 h. | Stand By |  |  |
| 12 | TBE1 | MBBR2221 |  |  | 05-05-2006 | 30 h. | 21 h. | 19 h. | 17 h. | 8 h. | 09-05-2006 | 95 h. | Stripping |  |  |
| 13 | TBE1 | MBBR2260 |  |  | 09-05-2006 | 24 h. | 23 h. | 18 h. | 19 h. | 10 h. | 13-05-2006 | 94 h. | Stripping |  |  |
| 14 | TBE2 | MBBR1350 |  |  | 06-05-2006 | 23 h. | 33 h. | 17 h. | 15 h. | 6 h. | 10-05-2006 | 94 h. | Stripping |  |  |
| 15 | TBF1 | MBBR2254 |  |  | 03-05-2006 | 26 h. | 19 h. | 171 h. | 17 h. | 8 h. | 13-05-2006 | 241 h. | Stripping |  |  |
| 16 | TBF2 | MBBR2264 |  |  | 04-05-2006 | 23 h. | 22 h. | 16 h. | 19 h. | 7 h. | 09-05-2006 | 87 h. | Stripping |  |  |

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBL1251 | 10380.3 A. (D1_L) | 11653.3 A. (D2_L) | 12325.8 A. (D2_U) | 12793.6 A. (D2_L) |  | - | - | - | - | - | - | - | - |
| 2 | MBBL1391 | 12604.1 A. (D1-U) | 12604.5 A. (D1-L) | 12811 A. (D2-U) |  | - | - | - | - | - | - | - | - | - |
| 3 | MBAL1385 | 11030.7 A. (D1-U) | 12083.6 A. (D1-L) | 12495.1 A. (D1-U) |  | - | - | - | - | - | - | - | - | - |
| 4 | SSS229 | 11574.1 A. (QE_12) | 12602.7 A. (QI_12) | 12780.3 A. (QE_34) |  | - | - | - | - | - | - | - | - | - |
| 5 | SSS333 | 12226.9 A. (QE_12) | 12496.8 A. (QE_12) | 12761.6 A. (QE-34) |  | - | - | - | - | - | - | - | - | - |
| 6 | SSS504 | 12849.4 A. (QE-34) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 7 | SSS504 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 8 | MBAR2322 | 12182.5 A. (D2-U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 9 | MBBR1276 | 10744.2 A. (D2_U) | 12420 A. (D2_L) | 12698.9 A. (D2_L) |  | - | - | - | - | - | - | - | - | - |
| 10 | MBBR1280 | 12706.1 A. (D1_L) | 12735 A. (D1-L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 11 | SSS609 | 3678 A. (ExtMQY_1) | 3728 A. (ExtMQY_1) | 3794 A. (Ext_MQY2) | 3828 A. (ExtMQY_2) | 3844.8 A. (Ex_MQY_1) | 3896.9 A. (In_MQY_2) | 3900 A. | - | - | - | - | - | - |
| 12 | MBBR2221 | 11993.3 A. (D1_L) | 12572.9 A. (D2_L) | 12853 A. | - | - | - | - | - | - | - | - | - | - |
| 13 | MBBR2260 | 10340 A. (D2_L) | 11779.9 A. (D1_L) | 12568.7 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 14 | MBBR1350 | 12052.5 A. (D1_L) | 11968.7 A. (D1-L) | 12853 A. | - | - | - | - | - | - | - | - | - | - |
| 15 | MBBR2254 | 11331 A. (D1_L) | 12741.6 A. (D2_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 16 | MBBR2264 | 11874.9 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |

609 done prev week  but not disconnected & removed this week; 504 has DCF leaks etc/pressure tests & re-cooldown etc

2254 had Long MM…

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_09_A | TBE1 | MBBR2221: during PT13.2 (resistance measurements), a difference on the connector was noticed. Pin-4 of (female) PT10 was not making contact with Jackie box male pin during measurement. | MTM |  | PT10 connector repaired by P. Legrand |
| Wednesday_10_A Thursday_11_M | TB all | No water is available for any of the test benches. TCR was contacted at 1900 h for rectification. AGAIN no water...TCR + Giorgio informed | AB/CO? |  | Giorgio came after discussing with people from demineralised water: water plant is OK, pressure OK, but the PLCCU is in Safe mode: all valve closed. Giorgio reset the PLCCU, and then it was OK. Water OK (Giorgio) Off/On on PLCCU + investigation... |
| Wednesday_10_A | TBD1 | SSS635 : No column was found for signature in Paper traveler after carrying out HV by OP. |  |  |  |
| Wednesday_10_A |  | Battery on Mobic 2 is very weak. Needs replacement. |  |  |  |
| Thursday_11_M | TBE2 | MBAR1293: Permanent trigger on potaim card starting IAP at warm. Not possible to charge quench heaters. Contacted ICS for connection check | ICS |  | ICS rectified the connection |
| Thursday_11_M | TBC2 | MBBR2255: Quench during PT8 on flat top after 3 mn at 12000 A (general cut of water). Unfortunately, the ub extension was already changed by error! Shift leader will ask Giorgio to change it again in training extension. | MTM |  | OK name of previous file (.ub) changed by Giorgio |
| Thursday_11_M | TBD2 | SSS658 : No magnet name label on the cold mass.. | CRI |  | Giorgio is informed. He discussed with the responsibles. |
| Thursday_11_M | TBB1 | ICS sent an SMS that new magnet (SSS038) had arrived on TBB1. It was scanned and MTR template printed and name added in SMTMS. Later on ICS informed that a new magnet (SSS 334) arrived in the same bench. | ICS |  | The SSS 038 was removed from SMTMS. NB: Why this sudden change? |
| Thursday_11_A | TBE1 | HV remote screen on the TEMA goes off in between the test | ? |  | OK from local station. |
| Thursday_11_N | TBB1 | SSS334:Starting CD connection circuit check no trigger was present. | MTM |  | Safety matrix configured for SSS504 (no sextupole). Back to configuration for arc SSS. |
| Thursday_11_N | TBB2 | The magnet CDR template indicates an FB type magnet, but the CQW2 template does not correspond to the measured values (it indicates type DA) | MTM |  | Changes uploaded in MTF. |
| Saturday_13_M | TBF1 | MB2255: Traveler is missing. | MTM |  | ICS informed. Chef the banc is hill, ICS phone is not responding |
| Saturday_13_M | TBF2 | HV test with XMT.MY92: Voltage is 4000V instead of 180V! | AB/CO | HV | Stop immediately all the tests with this rack. Why used for test in MB2236? |
| Saturday_13_A | TBC2 | MB1282: HV failed in step Dipole-quench heaters vs gnd failed, other steps Ok. | MTM | HV | Giorgio repeated the test. The error message is still the same. He advised to go on as the result of resistance is good (10^9 Ohms). Giorgio will discuss with Richard on monday. |
| Saturday_13_A | TBB1 | SSS334: Powering of CD and EF circuit not possible. Interlock that cannot be cleared. Application stuck in bad state? | AB/CO? |  | Giorgio called. He switched off/on the AC main power and could clear the interlock. |
| Saturday_13_N | TBE2 | MB1293: AQA alert during the warm-up quench.”DirVT exceeded the threshold (200V)” on the poles. | MTM | AQA | Action Giorgio or TC called. Giorgio advised to repeat the ramp to ultimate and the warm-up quench. Done with no Alert. Magnet signed for warm-up |
| Saturday_13_N | TBB1 | SSS334: PT 6.3 could not be launched. TEMA error during arming LF "Crate 2, Slot 21, MXI out error". Test re-launched first with data acquisition reset then with Sunmta33 workstation reset. No display on monitor of sunmta33 upon restart. Check application unsuccessful from TEMA. | AB/CO |  | Giorgio called. Problem with the LF crate No2, the power supply has been changed and the fan system has been changed (took from G cluster). |
| Saturday_13_N | TBD2 | SSS658: PT5.1 could not be launched from TEMA. Error message "faults present on PLC". OP team checked the test preparation and attempted several times unsuccessfully. QH_power OK fault seen in PLC however QHeaters are charged | AB/CO | TEMA (Power) | Test launched in local? GHH reset TEMA, quit all applications and re-launched, then OK. |
| Sunday_14_M | TBB | HF and LF Sun show black screen | AB/CO | LF | Giorgio called. He asked us to try with telnet remote control. But LF acquisition is not getting arm, The video card of the sun of cluster B (sunmta33) is not working. Giorgio launched hf and lf from sunmta34 via telnet to sunmta33. Then everything is OK. Giorgio left the message that hf and lf are visible on sunmta34 instead of sunmta33, and that both application must not be quit! AB-CO will fix the problem on monday. |
| Sunday_14_M | TBA1 | MB2236: During shaft insertion force exceeded 16kG. ICS stopped and informed OP team. | MTM |  | Giorgio called. He asked to proceed carefully. |
| Sunday_14_A | TBA1 | MB2236: HV failed with Rack MY92 on the step Cryoheaters vs. ground test. Megger stopped several time during all measurement. | MTM |  | Test repeated with rack MY93: same message. |
| Sunday_14_A | TBA1 | MB2236: Error during IAP at cold. Error in DB while executing hr_confgen. Error occurred while accessing to the DB. | IT |  | Giorgio asked to switch off HWREC. He came to check the files and everything was OK. |
| Monday_15_M | All | General Power cut. Power back but no Helium available. |  |  | OK at the beginning of the afternoon. |
| Monday_15_M | TBA1 | MB2236: TRU removed from the bench but MM2 and MM3 data cannot be found. In the white board it is written that MM are executed. | MTM |  | Stephane asked to reconnect the TRU and perform MM2 and MM3. |