### Week 21 starting 22 May 06

### SM18 Operation Meeting (Tuesdays) 23 May 06

Bench status  on Monday  22-05-2006 at: 5/22/2006 9:16:31 AM

| 1 | A1 | MBBR2304 | PT 9 CDAP HV | 1.90 |
| --- | --- | --- | --- | --- |
| 6 | A2 | MBBR2277 | PT 3 QH Measure | 4.48 |
| 7 | B1 | SSS263 | Prep 3 CDW1 HV | 298.58 |
| 11 | B2 | SSS628 | PT 14 Magnet MAPS next 505 | 299.11 |
| 5 | C1 | MBAL1393 | PT 12 Warm Up | 235.47 |
| 8 | C2 | MBBR2291 | ICS 3 Traveller after warm test | 298.04 |
| 9 | D1 | SSS647 | Under connection (for Tues ready?) | 298.67 |
| 12 | D2 | SSS663 | PT 14 Magnet MAPS | 299.11 |
| 10 | E1 | MBAL1388 | PT 13.2 Resist. Meas. | 300.82 |
| 3 | E2 | MBAR2275 | Prep 5 Cool Down | 4.19 HV |
| 2 | F1 | MBBR2321 | PT 12 Warm Up | 296.34 |
| 4 | F2 | MBBR2305 | Prep 5 Cool Down | 97.48 |

Tempr. from TBSM18 site

CONFIG :

D8- AS2 –IRSSS2;

Supervisor: marco/        Eq support : 16-3602 :

16 Magnet Tests completed :  12 Dipoles (),    1 Arc-SSS  ( no   REPEAT)

(16 Magnets in all)		1  500-series ( 524 :)

2  600-series [ 635 (343 hrs),  658 (186 hrs) ]

{  MM on 2236 ( 181 hrs)  }

General Points :

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- News /Situation from (SSS, SSSS)  :  J Halik

- Situation SSSS  :OP issues :

- A.O.B.

| Magnets under Test from 14-05-2006 to 20-05-2006 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBAR2236 |  |  | 11-05-2006 | 39 h. | 26 h. | 89 h. | 15 h. | 12 h. | 19-05-2006 | 181 h. | Stand By |  |  |
| 2 | TBA2 | MBAR1296 |  |  | 10-05-2006 | 30 h. | 24 h. | 20 h. | 15 h. | 14 h. | 15-05-2006 | 103 h. | Stripping |  |  |
| 3 | TBA2 | MBAR1298 |  |  | 15-05-2006 | 34 h. | 28 h. | 12 h. | 18 h. | 7 h. | 19-05-2006 | 99 h. | Stripping |  |  |
| 4 | TBB1 | SSS038 |  |  | 15-05-2006 | - | - | - | - | - | 17-05-2006 | 34 h. | Not Found or Thermal cycle |  |  |
| 5 | TBB1 | SSS334 |  |  | 11-05-2006 | 24 h. | 13 h. | 30 h. | 11 h. | 11 h. | 15-05-2006 | 89 h. | Stripping |  |  |
| 6 | TBB2 | SSS524 |  |  | 12-05-2006 | 27 h. | 18 h. | 33 h. | 8 h. | 10 h. | 16-05-2006 | 96 h. | Stripping |  |  |
| 7 | TBC1 | MBAR1274 |  |  | 14-05-2006 | 35 h. | 28 h. | 14 h. | 18 h. | 8 h. | 18-05-2006 | 103 h. | Stripping |  |  |
| 8 | TBC1 | MBBR2255 |  |  | 08-05-2006 | 27 h. | 47 h. | 15 h. | 17 h. | 10 h. | 14-05-2006 | 116 h. | Stripping |  |  |
| 9 | TBC2 | MBBR1282 |  |  | 13-05-2006 | 28 h. | 21 h. | 23 h. | 22 h. | 15 h. | 17-05-2006 | 109 h. | Stripping |  |  |
| 10 | TBD1 | SSS635 |  |  | 05-05-2006 | 151 h. | 18 h. | 120 h. | 6 h. | 48 h. | 18-05-2006 | 343 h. | Stand By |  |  |
| 11 | TBD2 | SSS658 |  |  | 09-05-2006 | 80 h. | 31 h. | 21 h. | 13 h. | 41 h. | 17-05-2006 | 186 h. | Stand By |  |  |
| 12 | TBE1 | MBBR2318 |  |  | 13-05-2006 | 28 h. | 29 h. | 24 h. | 17 h. | 10 h. | 18-05-2006 | 108 h. | Stripping |  |  |
| 13 | TBE2 | MBAL2549 |  |  | 15-05-2006 | 31 h. | 21 h. | 27 h. | 18 h. | 16 h. | 20-05-2006 | 113 h. | Stripping |  |  |
| 14 | TBE2 | MBAR1293 |  |  | 10-05-2006 | 32 h. | 22 h. | 25 h. | 15 h. | 24 h. | 15-05-2006 | 118 h. | Stand By |  |  |
| 15 | TBF1 | MBBR2261 |  |  | 13-05-2006 | 30 h. | 47 h. | 23 h. | 16 h. | 14 h. | 19-05-2006 | 130 h. | Stripping |  |  |
| 16 | TBF2 | MBBR1311 |  |  | 15-05-2006 | 30 h. | 43 h. | 16 h. | 15 h. | 8 h. | 20-05-2006 | 112 h. | Stripping |  |  |
| 17 | TBF2 | MBBR2257 |  |  | 09-05-2006 | 21 h. | 32 h. | 36 h. | 16 h. | 43 h. | 15-05-2006 | 148 h. | Stripping |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAR2236 | 12474.7 A. (D2-L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 2 | MBAR1296 | 11782.8 A. (D1-U) | 12605.9 A. (D1_L) | 12842.2 A. (D1_L) |  | - | - | - | - | - | - | - | - | - |
| 3 | MBAR1298 | 12853 A. | 12853 A. | - | - | - | - | - | - | - | - | - | - | - | - |
| 4 | SSS038 |  | - | - | - | - | - | - | - | - | - | - | - | - |  |
| 5 | SSS334 | 12214 A. (QE_34) | 12328 A. (QI_12) | 12554.6 A. (QE-12) |  | - | - | - | - | - | - | - | - | - |  |
| 6 | SSS524 | 11888.1 A. (QE_12) | 12246.1 A. (QI_34) | 12487.2 A. (QI-34) |  | - | - | - | - | - | - | - | - | - |  |
| 7 | MBAR1274 | 12464 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |  |
| 8 | MBBR2255 | 10372.5 A. (D2_U) | 12066.1 A. (D2_L) | 12290 A. (D1-L) |  | - | - | - | - | - | - | - | - | - |  |
| 9 | MBBR1282 | 12169.7 A. (D1_U) | 12517.6 A. (D2_U) | 12851.3 A. (D1_L) |  | - | - | - | - | - | - | - | - | - |  |
| 10 | SSS635 | 3385 A. (INT-12) | 3546.8 A. (ExtMQY12) | 3636 A. (ExtMQY34) | 3712 A. (ExtMQY34) | 3774.3 A. (Ex_MQY34) | 3845.7 A. (I_MQY12) | 3858.5 A. (1.9) | 3900 A. | - | - | - | - | - |  |
| 11 | SSS658 | 3686.5 A. (ExtMQY34) | 3764 A. (E-M-P12) | 3745 A. (I-M-34) | 3849 A. (E-M-34) | 3880.7 A. (Ex_MQ_34) | 3900 A. | - | - | - | - | - | - | - |  |
| 12 | MBBR2318 | 10659 A. (D2_L) | 12274.5 A. (D2_U) | 12798 A. (D1-L) |  | - | - | - | - | - | - | - | - | - |  |
| 13 | MBAL2549 | 10434.6 A. (D2_U) | 11760.6 A. (D1_L) | 11989.9 A. (D2_L) | 12418.2 A. (D1_U) | 12725.2 A. (D1-U) | 12852.4 A. | - | - | - | - | - | - | - |  |
| 14 | MBAR1293 | 12654 A. (D1_U) | 12819.2 A. (D1_L) | 12850 A. | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 15 | MBBR2261 | 11053.6 A. (D1_U) | 12593.1 A. (D1_L) | 12853 A. | - | - | - | - | - | - | - | - | - | - |  |
| 16 | MBBR1311 | 12281.1 A. (D1_L) | 12645.7 A. (D1_U) | 12544.8 A. (D1_L) |  | - | - | - | - | - | - | - | - | - |  |
| 17 | MBBR2257 | 11850.3 A. (D2_U) | 12758.1 A. (D2_L) | 12852 A. | - | - | - | - | - | - | - | - | - | - |  |

2236  had Long MM…  ;      sss38 was taken off at start;       635   & 658  .. explanations are well known why

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_16_M | TBE1 | TT821 is not working on E1 MBBR2318 and also files are not present in storage | MTM |  | TT821 non conformity has been opened by Federico, and telnet to local sun station (No data validation can be done). |
| Tuesday_16_A monday_15_M | TB all | Power cut @ CERN. First magnet @ cold. Waiting 3 H for first magnet @1.9K. Problem to restart water circuit after power cut. Waiting 8 H first magnet @ 1.9K | TCR |  | Federico informed us about water problem, and TCR will inform us when the problem is solved. |
| Tuesday_16_A |  | HV racks MY92 and MY95: I have replaced both Meggers in the 2 racks by new ones, because lately some communication problems were observed causing a wrong voltage to be applied during a test which can be potentially harmful to the magnet. Now it should back to normal. Richard. | MTM |  |  |
| Tuesday_16_A | TB All | New battery is installed into MOBIC 2. | MTM |  | Please inform Giorgio if the problem with Mobis battery persists. Giorgio |
| Tuesday_16_A | TBD1 | value 0 A during LHC cycle, Guy has been called and ask us to abort measurment. Launch Quench @ nominal current, then launch MM. Federico can be called as support. Federico is on call if the config file are not working properly. | MTM |  | OP team requested a procedure to equipment support because no information was given for this measurement on S4. |
| Tuesday_16_A | TBA1 | Equipment support has been called without success, about a problem with power converter "Ext2" error. | MTM |  | On cluster A, a rack near top of PC needs a reset in case of power cut. Action done by OP, because N. SAMUT waiting for measurement since power cut and Equipment support not responding. |
| Wednesday_17_N | Video | Video Display SMTMS and Traveller are fixed without change. | OP |  | Video server was hanging, restart has been made. |
| Thursday_18_A | TBD2 | SSS635: Magnet was removed from bench without any information. No final HV was done. It has not come for signing stripping/standby as yet. | MECA |  | MECA should signed like ICS for each step of electrionic traveler, to not miss test. Test has been perform Friday 19 16:00. |
| Thursday_18_A | TBB1 | There is no place in the paper traveller to sign for 'CFB Electrically safe to be disconnected' after the warm final resistance measurement. Hence paper traveller not signed after PT13.2. | ICS |  | ICS has been informed. |
| Friday_19_A Saturday 20_M | TBC & TBE TBE | 18:36 : Tema problem on C and E cluster No name from the DB (Hw rec) Many restart of server and Tema Finally works for C but not for E Waiting PIquet for E TBE control error | MTM |  | 20:54 : Test done in local for E cluster. Sunias08 stopped and restarted - OK |
| Saturday 20_M | TBB | TBB problem of LF acquisition : SW OFF & Reboots but still problem | MTM |  | Calling TS |
| Saturday 20_M | TBD2 | TC called to inform him about D2 Alarm (MIITS 3.5) | MTM |  | TC inform us : “Going ahead with training” |
| Saturday 20_A | TBB2 | Unable to lunch Training (PT6.1). FB loop fault shows. Contacted Equipment support. | MTM |  | Federico came and solved the problem. |
| Sunday 21_N | TBB | HV panel not visible on TEMA | AB/CO | HV |  |