### Week 22 starting 29 May 06

### SM18 Operation Meeting (Tuesdays) 30 May 06

Bench status  on Monday  29-05-2006 at: 5/29/2006 10:22:26 AM

| 7 | A1 | MBBL1236 | Min Energy @ 6500 | 5.47 |
| --- | --- | --- | --- | --- |
| 9 | A2 | MBAR1354 | PTE Min. NRJ 3KA Diode Test | 1.90 |
| 6 | B1 | SSS264 | PT 13.2 Resist. Meas. | 298.94 |
| 12 | B2 | SSS505 | PT 14 Magnet MAPS | 300.82 |
| 10 | C1 | MBAR2310 | Prep 5 Cool Down | 98.46 |
| 8 | C2 | MBAR2272 | PT 6.4 Training 4 | 1.90 |
| 4 | D1 | SSS647 | MM - Stretched wire at 4.4K | 2.35 |
| 1 | D2 | SSS677 | PT 14 Magnet MAPS | 299.83 |
| 11 | E1 | MBAR2276 | ICS 4 Final connection | 300.37 |
| 5 | E2 | MBAR1353 | PT 9 CDAP HV | 2.35 |
| 3 | F1 | MBBR2279 | PT 14 Magnet MAPS | 299.83 |
| 2 | F2 | MBAR2327 | ICS 3 Traveller after warm test | 299.92 |

CONFIG :

D8- AS2 –IRSSS2;

Supervisor: Pierre/        Eq support : 16-3602 :

14 Magnet Tests completed :  10 Dipoles (),    2 Arc-SSS  ( no   REPEAT)

(14 Magnets in all)		0  500-series ( 524 :)

(2255 was a  mistake)             2  600-series [ 628 (134 hrs),  663 (114 hrs) ]

General Points :

- Last weekend problems  ( (Eric) : F2 2279/

  - A & F clusters both cold due to problems on 2279 & 1236 (MM)

  - D1 problems powering ( IAP adjust) SSW, Shafts delay etc

  - D2 Special M’ments & switching back & forth?

  - 505 could not be removed since last Friday ? (Rocla, Transport etc)

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- News /Situation from (SSS, SSSS)  :  J Halik

- Situation SSSS  :OP issues :

- A.O.B.

SITUATION ON SATURDAY EVENING   at 20.45 hrs  27th May 2006;

TOO Many Things all cold same time;   A1, A2;   F1, F2;   D1 ,D2

## Tests

| Magnets under Test from 21-05-2006 to 27-05-2006 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBBR2304 |  |  | 19-05-2006 | 24 h. | 20 h. | 24 h. | 15 h. | 16 h. | 23-05-2006 | 99 h. | Stripping |  |  |
| 2 | TBA2 | MBBR2277 |  |  | 19-05-2006 | 21 h. | 28 h. | 37 h. | 22 h. | 10 h. | 24-05-2006 | 118 h. | Stripping |  |  |
| 3 | TBB1 | SSS263 |  |  | 21-05-2006 | 31 h. | 15 h. | 23 h. | 9 h. | 12 h. | 25-05-2006 | 90 h. | Stripping |  |  |
| 4 | TBB1 | SSS265 |  |  | 17-05-2006 | 29 h. | 19 h. | 18 h. | 10 h. | 32 h. | 21-05-2006 | 108 h. | Stripping |  |  |
| 5 | TBB2 | SSS628 |  |  | 16-05-2006 | 31 h. | 18 h. | 51 h. | 11 h. | 23 h. | 22-05-2006 | 134 h. | Stand By |  |  |
| 6 | TBC1 | MBAL1393 |  |  | 18-05-2006 | 29 h. | 20 h. | 23 h. | 18 h. | 27 h. | 23-05-2006 | 117 h. | Stripping |  |  |
| 7 | TBC1 | MBBR2315 |  |  | 23-05-2006 | 32 h. | 23 h. | 19 h. | 14 h. | 7 h. | 27-05-2006 | 95 h. | Stripping |  |  |
| 8 | TBC2 | MBBR2287 |  |  | 17-05-2006 | 24 h. | 24 h. | 21 h. | 18 h. | 8 h. | 21-05-2006 | 95 h. | Stripping |  |  |
| 9 | TBC2 | MBBR2291 |  |  | 21-05-2006 | 29 h. | 20 h. | 17 h. | 15 h. | 23 h. | 26-05-2006 | 104 h. | Stripping |  |  |
| 10 | TBD2 | SSS663 |  |  | 17-05-2006 | 33 h. | 18 h. | 25 h. | 10 h. | 28 h. | 22-05-2006 | 114 h. | Stand By |  |  |
| 11 | TBE1 | MBAL1388 |  |  | 18-05-2006 | 24 h. | 20 h. | 19 h. | 16 h. | 20 h. | 22-05-2006 | 99 h. | Stripping |  |  |
| 12 | TBE1 | MBBR2255 |  |  | 22-05-2006 | 9 h. | 0 h. | 0 h. | 0 h. | 16 h. | 23-05-2006 | 25 h. | Stand By |  |  |
| 13 | TBE2 | MBAR2275 |  |  | 20-05-2006 | 24 h. | 21 h. | 16 h. | 15 h. | 34 h. | 25-05-2006 | 110 h. | Stripping |  |  |
| 14 | TBF1 | MBBR2321 |  |  | 19-05-2006 | 22 h. | 22 h. | 14 h. | 15 h. | 9 h. | 22-05-2006 | 82 h. | Stripping |  |  |
| 15 | TBF2 | MBBR2305 |  |  | 20-05-2006 | 28 h. | 22 h. | 16 h. | 15 h. | 16 h. | 24-05-2006 | 97 h. | Stripping |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBR2304 | 10294.9 A. (D1_L) | 11575.3 A. (D2_U) | 12026.8 A. (D2_L) | 12427.8 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - |
| 2 | MBBR2277 | 10223.2 A. (D1_U) | 11814.4 A. (D2_U) | 12296.8 A. (D2_L) | 12442 A. (D1_U) |  | - | - | - | - | - | - | - | - |
| 3 | SSS263 | 11833.7 A. (QE_12) | 11982.5 A. (QE-34) | 12282.3 A. (QE_12) | 12000 A. (No Data) | 12000 A. (No Data) | 12504 A. (QE-34) |  | - | - | - | - | - | - |
| 4 | SSS265 | 11121.8 A. (QI-34) | 12082.7 A. (QI_12) | 12234 A. (QE_12) |  | - | - | - | - | - | - | - | - | - |
| 5 | SSS628 | 5692.1 A. (E_MQML12) | 5820 A. | - | - | - | - | - | - | - | - | - | - | - |
| 6 | MBAL1393 | 10553.3 A. (D1_L) | 11615 A. (D2_U) | 12315.7 A. (D2-U) | 12533.1 A. (D1-U) |  | - | - | - | - | - | - | - | - |
| 7 | MBBR2315 | 12374.3 A. (D1_L) | 12256.6 A. (D2-U) | 12800.5 A. (D1_L) |  | - | - | - | - | - | - | - | - | - |
| 8 | MBBR2287 | 9830 A. (D1_L) | 11859.3 A. (D1_U) | 12268.5 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 9 | MBBR2291 | 11198.6 A. (D1_L) | 11851.3 A. (D2_L) | 12772.1 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 10 | SSS663 | 5582 A. (MQMC_EXT) | 5850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 11 | MBAL1388 | 11529.2 A. (D1_L) | 12204 A. (D1-L) | 12544 A. (D1-U) |  | - | - | - | - | - | - | - | - | - |
| 12 | MBBR2255 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 13 | MBAR2275 | 10698.9 A. (D!-U) | 12221 A. (D1_U) | 12714 A. (D2_L) |  | - | - | - | - | - | - | - | - | - |
| 14 | MBBR2321 | 12491 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 15 | MBBR2305 | 11112.7 A. (D2_L) | 12219.1 A. (D2_U) | 12557.4 A. (D1_U) |  | - | - | - | - | - | - | - | - | - |

2255 brought by mistake ;     so only 14 magnets in whole week due to problems & holiday w’end (ascension) :

( SSSS & powering +SSW,  special tests .. etc)

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_23_M | TBB1 | SSS263 : PT2 HV-CQBP fails (Quadrupole vs Ground shows high current > 10A). | MTM |  | Richard advised to re-launch the test and check. |
| Tuesday_23_M | TBD2 | SSS677: Lyre test problem: current 2A at 15V not possible. | MTM |  | Richard advised to conduct test at 15V,1A. Final test was OK Action : Define a global rule for S4: i.e. 1A and 15 V |
| Tuesday_23_M | TBB1 | SSS263: Magnet name not appearing in TEMA. So proper test sequence could not be selected. | MTM | TEMA | Richard advised to re-launch TEMA application. The problem was solved. |
| Tuesday_23_A | TBB1 | SSS263 : Communication error with PLC & LF. | AB/CO |  | Closing all lab-view applications and re-launching them was not sufficient to get it restarted. Communication error was solved by shutting down the Sun machine. |
| Tuesday_23_N | TBF2 | MB2305 : IAP@warm. LF DAQ did not trigger. | AB/CO | lf | OP team reset all the crates and Suns, TEMA application is stuck with "Error receiving info from udp socket". IAP test was executed locally. |
| Tuesday_23_N | TBB1 | SSS263: MQ quenched twice on the 12kA plateau during PT8 test. Cryo will increase the flow in the bus-bars for the next attempt. | ACR |  |  |
| Wednesday_24_M Thursday_25_A Saturaday_27_M | TBB1 TBB2 TBB1 | HV panel is not opening at control room. Same problem during HV test on SSS505. Same problem for HV test on SSS264 | IT | HV | Network traffic increased? HV panel aperture is set in low priority. |
| Wednesday_24_M | TBF | Communication problem coming from UDP socket. | IT? |  | Richard contacted and had to reboot TEMA in Sun machine. Traffic in the network? |
| Wednesday_24_M | TBB2 | SSS505 : PT4 could not be done because of cryo instability | ACR |  | Why? |
| Wednesday_24_A | TBD1 | SSS647: Error in HV while doing step related to orbit dipole. Errors observed in Temperature sensor and Cryo heater steps also in the report but not while performing tests. | MTM |  | Richard asked to skip step. Richard will correct the report and generate NC as required. |
| Wednesday_24_A | TBB2 | SSS505: Corrector powering EF could not be launched due to trigger condition. Origin : NO CD circuit in 5XX series. | MTM |  | Richard cleared the triggers and restarted the system. Corrector powering could be carried out subsequently. |
| Wednesday_24_A | TBD1 | SSS647: Discrepancy between the values read by TT821-PT147 sensors. | ACR |  | Richard asked to ignore PT147 as this sensor is not calibrated down to 4.5K Action : TT821 will be compared with PT148. |
| Wednesday_24_A | TBB2 | SSS505: Corrector EF powering: 10 quenches done but still the corrector is not trained. Detraining observed in 9th and 10 quenches (from 590 to 540A) Files not generated for the 9th and 10th quench. | MTM |  | TC informed. He asked to perform one more training. Test was repeated multiple times in the afternoon shift. Marco was contacted and confirmed that test can be considered successful as current is greater than 400 A. |
| Thursday_25_M | TBB2 | SSS505: Auxiliary Current lead temperature problem, OP team is not able to launch PT5.8 corrector powering (MN). Cryo has been informed. | ACR |  | Solved by ACR. |
| Thursday_25_A | TBD1 | SSS647: SSW could not be initiated because of an interlock. Power converter could not be put on "ON STANDBY". | MTM |  | Richard contacted. He found out that the interlock came from the offsets of some cards for the other magnet in the cluster (MQM). Origin : Wrong sequence selected. With the correct sequence the problem persists. Richard had also redone the AC compensation, on the first 2 Potaim cards, in order to be able to switch ON the power converter. Origin non symmetrical inductance in one magnet of the test bench. Action: For S4, AC compensations have to be r-done when there is a switch for one bench to another. |
| Thursday_25_N | TBB1 | SSS264 is shown on the e-traveller in place of SSS 269. Magnet entered in SMTMS properly as SSS 264. But on CRYO screens it is reported as SSS 269. | ACR? ICS? |  |  |
| Thursday_25_N | TBD1 | SSS647: Strength measurement and roll angle tests carried out at 2000A with SSW. The results were bad and were communicated to Marco (TC). | MTM |  | TC asked to send the result files by mail. Files have been sent. Tried to run roll angle thrice but the results were bad. Stages were misaligned by 6 mm wr MQY. SSW adjusted Friday morning. |
| Friday_26_M | TBA1 | Display on TBA1 is showing the previous dipole name, MB2235 instead of MB1236. | ACR |  | Stephane contacted TC for action with J.Axensalva. |
| Friday_26_M | TBD1 | SSS647: SSW measurements of strength faulty quench outside the magnet (Ext-Vbus trigger). | ACR |  | A. Siemko asked cryo_team to increase the flow of cool down, because a problem of N line. Cryo team has been informed. |
| Friday_26_N | TBF1 | MB2279:Second training quench : alarm was reported in AQA (D1_L DirVT voltage exceeded 200V) | MTM |  | TC contacted. He advised to proceed with HV test and check with the warm _up quench. OK, no alarm only warning. |
| Friday_26_M | TBA1 | MB1236: Extended test planned by Stephane after MM3. Launching the tests by TEMA are not working. | MTM |  | TC and Richard contacted. Congig files of the cycles must be modified with new versions of MMP. Cleaning of the cycles will be performed (Action Guy, Stephane). |
| Saturday_27_A | TBE1 | MB2284: Problem in launching HV Test due to communication failure between PLCHV and PLCSE. | AB/CO |  | PLCSE needed RESET to launch CDBP. |
| Saturday_27_N | all | Video display in control room stuck. |  |  |  |
| Saturday_27_N | TBD2 | SSS677: Communication problem between D.Bocian and OP team. Darius performed separately the warm_up quench. Change of shift occurred and details of the special test were not known. |  |  | CQPT was not uploaded. |
| Saturday_27_N | TBB1 | SSS264: PT5.4 failed due to TCLAux_x1 signal problem. BP problem., oscillation. | ACR? |  | Test launched after 3 hours and many attempt. |
| Sunday_28_M | TBD1 | SSS647: SSW measurements, same problem as Friday. Quench in Ext-Vbus after few minutes at high current. |  |  | A. Siemko advised. TC will be contacted but according A. Siemko this magnet have specifique cable and it is not possible to make long measurement with current applied. Magnet put in stand_by up to Monday for investigation |
| Monday_29_A | All | Video display is not working in Bellegarde side. |  |  |  |