### VC- /Week 45 starting 13 Nov  06

### SM18 Operation Meeting (Tuesdays) 14 Nov 06

Bench status  on Monday 13-Nov-2006 at: 11/13/2006 9:12:11 AM

| 4 | A1 | MBBR2395 | Prep 5 Cool Down | 95.62 |
| --- | --- | --- | --- | --- |
| 12 | A2 | MBAR2403 | Prep 3 CDW1 HV | 295.62 |
| 10 | B1 | SSS366 | PT 14 Magnet MAPS | 295 |
| 2 | B2 |  |  | 295.54 |
| 11 | C1 | SSS064 | Test coordinator intervention | 4.60 HV |
| 5 | C2 | SSS371 | PT 14 Magnet MAPS | 295.18 |
| 6 | D1 | SSS671 | PT 1 IAP @ cold | 2.35 |
| 9 | D2 | SSS670 | PT 12 Warm Up | 215.15 |
| 1 | E1 |  |  | 295.89 |
| 7 | E2 | MBAR2406 | PT 12 Warm Up | 29.57 |
| 3 | F1 | MBAL1415 | PT 14 Magnet MAPS | 293.57 |
| 8 | F2 | MBAR1361 | PT 14 Magnet MAPS | 294.19 |

CONFIG :  D 6 - AS 2 -  IRSSS 4 ;

Supervisor: ;                        Eq support 16-3602 :

11 Magnet Tests completed ( 8 Magnet Stripped )

8 Dipoles tests	( 2405 stndby)

2 Arc-SSS (  047 ok in HV )

1 500-series (521 )

0  600-series []

General Points:

- S064  c1 waiting since  Sat evening 18.00 hrs for no obvious reason…!

- Situation from Cryogenics:  J-P Lamboy or ACR Rep :

- Situation SSSS  :D1 & D2 issues J Halik

- A.O.B

| Magnets under Test from 05-11-2006 to 11-11-2006 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBBR2398 |  |  | 28-10-2006 | 21 h. | 26 h. | 63 h. | 16 h. | 53 h. | 05-11-2006 | 179 h. | Stripping |  |  |
| 2 | TBA1 | MBBR2400 |  |  | 05-11-2006 | 36 h. | 23 h. | 15 h. | 16 h. | 40 h. | 11-11-2006 | 130 h. | Stripping |  |  |
| 3 | TBA2 | MBBL1405 |  |  | 03-11-2006 | 39 h. | 25 h. | 36 h. | 15 h. | 39 h. | 10-11-2006 | 154 h. | Stripping |  |  |
| 4 | TBC1 | SSS362 |  |  | 04-11-2006 | 21 h. | 12 h. | 12 h. | 6 h. | 35 h. | 09-11-2006 | 86 h. | Stripping |  |  |
| 5 | TBC2 | SSS047 |  |  | 31-10-2006 | 38 h. | 29 h. | 14 h. | 7 h. | 76 h. | 07-11-2006 | 164 h. | Stripping |  |  |
| 6 | TBD1 | SSS521 |  |  | 23-10-2006 | 92 h. | 17 h. | 142 h. | 12 h. | 91 h. | 07-11-2006 | 354 h. | Stripping |  |  |
| 7 | TBE1 | MBAR2405 |  |  | 02-11-2006 | 37 h. | 21 h. | 20 h. | 17 h. | 11 h. | 08-11-2006 | 106 h. | Stand By |  |  |
| 8 | TBE2 | MBBL2355 |  |  | 30-10-2006 | 34 h. | 19 h. | 23 h. | 20 h. | 38 h. | 05-11-2006 | 134 h. | Stripping |  |  |
| 9 | TBE2 | MBBR2401 |  |  | 05-11-2006 | 29 h. | 21 h. | 17 h. | 16 h. | 26 h. | 10-11-2006 | 109 h. | Stripping |  |  |
| 10 | TBF1 | MBBL1407 |  |  | 31-10-2006 | 30 h. | 24 h. | 17 h. | 19 h. | 65 h. | 07-11-2006 | 155 h. | Stripping |  |  |
| 11 | TBF2 | MBAR1412 |  |  | 01-11-2006 | 36 h. | 20 h. | 26 h. | 17 h. | 29 h. | 07-11-2006 | 128 h. | Stripping |  |  |

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBR2398 | 10364 A. (D1_U) | 10428.3 A. (D2_L) | 11526.8 A. (D2_U) | 11963 A. (D2-L) | 12281 A. (D1-U) | 12611 A. (D1_U) | 12727 A. (D2_L) | 12850 A. | - | - | - | - | - |
| 2 | MBBR2400 | 10968.9 A. (D2_L) | 12003.9 A. (D1_U) | 12372.9 A. (D1_U) |  | - | - | - | - | - | - | - | - | - |
| 3 | MBBL1405 | 11477.8 A. (D2_L) | 12541.4 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 4 | SSS362 | 12020.9 A. (QE_12) | 12562 A. (QI_34) | 12732.7 A. (QE_12) |  | - | - | - | - | - | - | - | - | - |
| 5 | SSS047 | 12374 A. (QE_34) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 6 | SSS521 | 11506 A. (QI_34) | 12175 A. (QI_12) | 12325.8 A. (QI-34) | 12578.9 A. (QE_12) | 12636.1 A. (QE_34) | 12764.9 A. (QE_12) | 12808 A. (QI_34) | 12847.3 A. (QI_34) | 12850 A. | - | - | - | - |
| 7 | MBAR2405 | 12244.3 A. (D1_L) | 12858 A. | - | - | - | - | - | - | - | - | - | - | - |
| 8 | MBBL2355 | 12031.5 A. (D2_U) | 12514.8 A. (D1_U) | 12783.8 A. (D2_U) |  | - | - | - | - | - | - | - | - | - |
| 9 | MBBR2401 | 12195.3 A. (D2_L) | 12771.1 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 10 | MBBL1407 | 11874 A. (D2_U) | 11852.1 A. (D2_L) | 12407.5 A. (D2_U) | 12575.2 A. (D2_U) |  | - | - | - | - | - | - | - | - |
| 11 | MBAR1412 | 10942.7 A. (D1_U) | 11379.3 A. (D1_L) | 11856.5 A. (D2_U) | 12351 A. (D1-L) | 12642 A. (D2-L) | 12850 A. | - | - | - | - | - | - | - |

COMPARISON WITH 11 magnets p.week vs 18 magnets per week from earlier this year; Vertical scale too is different!

| Date | Bench | Problem | Type | Appli. software | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_07_A | TBA2 | MB1405: During MM3 (load-line) power supply tripped. | MTM |  | Stéphane asked to repeat the test. |
| Tuesday_07_A | TBA2 | MB1405: In the result of MM3 full load line, Angle difference error & Std. deviation error is found | MTM |  | Stéphane contacted. He validated the measurements. The errors are coming from the misalignment in height of the anticryostat short in the CFB with respect to the long ones in the magnet. Readjustment done Thursday by MTM mechanic team. |
| Tuesday_07_N | TBD2 | SSS670: Scan has been performed again because identification on MRB was not available during connection. | MTM |  | NOACXX has been scan on MRB and temperature became OK for cryo. |
| Tuesday_07_N | TBE | DAQ applications are very slow hfoff/AQA, TBA HF and LF take same time around 4-5 mins to be loaded. | IT? | Hf, lf |  |
| Tuesday_07_N | TBB | TBB1: SSS366 Communication problem on the main power converter during auto test. A test launch with MMP does not communicate with the power converter. | AB/CO |  | Manual communication check of power converter need to be done, before launch a MM. |
| Wednesday_08_A | TBD2 | SSS670: QH Resistance measurements @ cold. The resistances were well out of tolerance when measured via the P10 connector. | MTM |  | Richard measured the resistances via the IFS box. OK. Additional contribution came from the length of the measurements cables. |
| Wednesday_08_N | TBD2 | SSS670: Second training quench at 33 A for the MQY. Third at 40 A! ACR team cannot feed the main bus bar with He liquid. | MTM/ACR? |  | L. Herblin adjusted the cryogenic parameters in order to obtain enough pressure difference and manage to cool the bus bar. |
| Thursday_09_M | TBB1 | SSS666: Cryo OK lost during special machine cycle. | ACR |  |  |
| Thursday_09_A | TBC1 | SS064: LF not getting triggered during Prep2.2C. | MTM |  | Richard found that two potaim cards are blocked in the trigger mode. He changed the two cards on the crate 2. |
| Thursday_09_N |  | Video screen control is not updating. |  |  | . |
| Friday_10_M | TBD1 | SSS671: Paper Traveller not present on the bench. | ICS? ACR? |  |  |
| Saturday_11_M | TBD1 | SSS671: T_anticryostat rising during the cool down phase. | MTM |  | Cryo team informed. |
| Saturday_11_M | TBB1 TBC2 | SSS366: During CQW2 resistance template for this magnet is wrong in SMTMS, there is first entry missing for Diode voltage taps measurements. SSS:371 NC was generated in CQW2 due to wrong CQR template | MTM |  | This template has been modified and updated in SMTMS. Corrected version of file for this magnet has been uploaded in MTF. Template was corrected and NC closed. |
| Saturday_11_A | TBC1 | SSS064: HV CQBP was successful | MTM |  | TC asked to maintain the magnet in stand_by for further tests Monday. Why? |
| Saturday_11_N | TBD1 | SSS671: HV cryo Lost twice during HV (CDBP) | ACR |  | Cryo team said that HV cryo phase not stable. No stability Sunday morning. |
| Sunday_12_A | TBD2 | SSS670: Long training of MQY: it took 14 quenches to go to 3794 A. | MTM |  | TC consulted. He advised to end_up with the test program. And warm up the magnet without ROBTC) |
| Monday_13_M | TBD1 | SSS671: QH3 I max too low (~49.6 A). Red alert on AQA | MTM |  | Giorgio checked: OK to proceed. |