### SM18 Operation Meeting (Tuesday 20 Feb  07)

Bench status  on Monday 19-Feb-2007

Situation on the test benches ;

-MB1334 warmed up. 1 MB for the machine in cold test MB2394.

-All SSS tested. SSS529 : Good quench performance after the thermal cycle.

-Special tests on MB2598 (spare magnet, test with laser, MM program by Nicholas Sammut).

Situation from Cryogenics.

Situation of mechanics.

General Points:

- Last SM18 OP meeting.

- Test of spare magnets. All with Anti-cryostats. Standard tests + MM.

- Some spare magnets will come with no diode: No PTE test for them.

- Cryo team from 7h30 to 20 h up to the 15 March.

4  Magnet Tests completed () + 1 run on SSS529;

2 Dipoles tests

2 SSS 500 series.

### Quench Information

Many Thanks To All the Teams

For These Five Years of

Successful Collaboration!

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_13_M | TBE1 | MBBR 2424: WP07 was done in morning. The test was failed in Dipole Vs MCS Correctors. | ICS |  | Trying to contact ICS repeatedly but not responding. Hence magnet waiting for stripping or standby. Richard executed the test again: OK. Ground connection was not fixed properly. |
| Tuesday_13_A | TBE2 | MB2426 : Cyro phase lost while launching PT9 (HV after powering) and hence test could not be launched. | ACR |  |  |
| Tuesday_13_A | TBD1 | SSS529: Load-line on MQTL. The 600 A power supply could not stand more than 100 A in negative polarity. | MTM |  | Guy called AB/PO. Check of the loop OK. Ramp rate 5A/s to high for the threshold voltage of potaim card. Threshold increased and OK. |
| Wednesday_14_M | TBB1 | SSS374 : MM3 could not be launched from TEMA. | AB/CO | MMP | Guy called. He reset the DIGIO card. OK. |
| Wednesday_14_M Wednesday_14_A | TBA1 | MB1334: Problem in the Main power converter (RMA), AB/PO contacted. PT4 could not be launched. | AB/PO |  | Richard contacted ABPO on inferring that the circuit breakers are faulty. RMA power converter is back in operation after that AB/PO had replaced the circuit breaker (MCB). |
| Wednesday_14_A | TBD1 | SSS529: Current automatically dropped to zero (without quench heaters being discharged, and no disturbance in cryo phase) half-way during ramp_660 (special MM). | MTM |  | Guy instructed to stop measurement for the day. He will check the result tomorrow morning, and will do the remaining tests. |
| Monday_19_M | TBE1 | MB2394: HV application could not be launched due to communication error with PLCSE, even after changing HV rack. | MTM |  | Problem rectified by resetting the PLC –HV. |
| Monday_19_M | TBE1 | MB2394: Step 'Cold temperature sensor vs ground' failed in CDBP. High leakage current detected. | MTM |  | Richard detected that a pin of the temperature sensor connector inside the IFS box is bent. NC generated and proceeding further, as suggested by Richard. |