Tss22-Mar-04

### Week 12  starting 15 Mar 04

### SM18 Operation SCHEDULE Meeting (Tuesdays)

Bench status  and tests for this week :

- Shafts in 2069 (F1-QLOC); 3010 (TBD2-MM); 2066 (TBC1-MM).

Test Coordinator this week:  Nicolay (164674); MTM support on call Guy:  (16-3602).

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Leaks: 3105 (leak in the magnet); 3056 (joint in the flange of the long anticryostat).

- Process: Lost of cryo OK in TBC2: Level in pressure detectors  changed. Solved. Lost of cryo OK after small quenches : level in the probe in the current leads corrected.

- Other General Issues

Request of F.Flamand : After a MM, please plug the cover of the connector cables coming from the shaft.

| Date | Problem | Type | Application/soft. | Action/advise |
| --- | --- | --- | --- | --- |
| Thursday 11_A Friday 12_M | TBD2: The shaft could not rotate well in aperture 1 when it is closed to the MRB. TBD2: Rotation in aperture 1 not good (3%). The “flat cables” is scratching the the protection cover. LHC cycle: Over-ranges in aperture 1 at high currents: | ICS |  | Pierre, Stephane informed. Common decision was to use as it is because the situation looks like better in MM and quench position. Origin of problem found by mecanician team . J.Halik will inform ICS Maryline informed Stephane. OK to continue. Maryline informed Stephane. He advises to reduce the speed for the load line. |
| Thursday 11_A | TBD2: Crate 2 slot 8 Tcoil out of range. Investigation by Maryline :ADC card not working.. | AB-CO |  | Card changed Monday. |
| Tuesday 09_M Friday 12_ M | All benches with MM : Problem in Sun MTA24. Display is not working. Vi missing Problem appeared again. | AB-CO |  | Repaired by A.Raimondo. Investigation on the run. |
| Thursday 11_A, N | Cluster E, D,F: HWREC switched off because MTF is not running (upgrade) . When MTF operational : no access to the DB for cluster D, E. Quit all application, restart TEMA, Re-launched again : Error was : error updating DB unable to access hr_chklst” | TS/CSE | MTF HWREC | . Marlyne’s instructions:HWREC stopped again . Names entered manually in TEMA for cluster D and E. Maryline generated the settings manually. Maryline informed T.Ladzinski. He solved the problem. An immediate check is now added in the procedure after each upgrade. |
| Thursday 11_N | TBD2: SPA test. Interlock water. Electro valve was closed (low water pressure). TEMA was launched but hf application hung. | TS/EL AB-CO | TEMA | Other valves closed and interlock disappears (Florence). Re-launch and OK (Florence) |
| Friday 12_M | TBE1: Task phase not signed (break vacuum) because L.Herblin asked to make a leak test before. ICS activity stopped. J.Axensalva contacted to change the status: stripping to stand by. |  |  |  |
| Friday 12_A | TBC1, TBD2: HV tests: Some new serial ports installed by AB-Co in fixed rack are not working properly. | AB-CO |  | Richard consulted. HV Tests performed using other serial ports. M.Duron is investigating to solve the problem (Monday). |
| Friday 12_A | TBE1: Problem in the Task Tracking System. ICS did not see the changes in the task because there is no automatic refresh, | ACR |  | J.Axensalva is solving this problem. |
| Sat 13_N | TBD2: HV initial at warm . The meeger stopped four times during the test. Run-recovery then “error 15530” | ACR | HV | OK after shaking ??? communication cables. (Jackie) |
| Sunday 14_N | TBD2: Alarm anti-cryostat during the night. | MTM |  | Support not called ?? |