### CHOHAN /10--Oct- 05

### Week 41 starting 10 oct 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  10 Oct 05 : 10/10/2005 8:47:15 PM

| 1 | A1 | MBBL2216 | PT 10 Min. NRJ 11850 A | 1.90 |
| --- | --- | --- | --- | --- |
| 7 | A2 | MBAL1249 | ICS 4 Final connection | 294.28 |
| 3 | B1 | SSS629 | Prep 5 Cool Down | 18.88 |
| 8 | B2 | SSS349 | Prep 3 CDW1 HV | 295.8 |
| 11 | C1 | MBBR1263 | PT 13.2 Resist. Meas. | 299.2 |
| 4 | C2 | MBBL2071 | Prep 5 Cool Down | 98.13 |
| 2 | D1 | SSS249 | PT 12 Warm Up | 273.06 |
| 10 | D2 | SSS630 | Finished Ready to go Final scan | 295.18 |
| 6 | E1 | MBAR2229 | Prep 5 Cool Down | 267.56 |
| 12 | E2 | CRYO | Cryo leak repair | 295.62 |
| 9 | F1 | MBBL1189 | Prep 3 CDW1 HV | 294.1 |
| 5 | F2 | MBAR1253 | Prep 5 Cool Down | 245.47 |

Tempr. from TBSM18 site

Coordinator this week: Marta; MTM support: 16-3602

15 complete Magnet Tests cycles  from 11 benches ( E2 under CFB repair)

10 dipoles + 3  SSS  + 2  SSSS

Disconnected or ROBTC after tests last week ( Sat Midnight to Sat midnight)

Situation From Cryogenics:  J-P Lamboy or ACR Rep.

[ CRYO PROBLEMS SAT & SUN ? Cooldown of sss629 on B1 ?]

E2 CFB (leak) Repair whole Week this week again?  ( 11 benches usage only)

Special  Points :

- Special SSS on B1 : ( )

- Special SSS on D2 :Tests Plans ?

- AOB

15 Magnets this week. Only 11 Benches.( E 2 under CFB leak repair) 2 benches with special SSSS

+ at least 2 magnets more were finished on Saturday(sss160 & sss630)  BUT ICS had STAFF problems ( only 1 person on shift)

Status Sun Morning09/10/2005 11:40:08

| 3 | A1 | MBBL2216 | PT 2 HV Cold Test | 4.39 |
| --- | --- | --- | --- | --- |
| 8 | A2 | MBBR1261 | ICS 6 Disconnect Magnet WP07 | 294.73 |
| 11 | B1 | SSS629 | CR 1 Leak Test | 294.55 |
| 9 | B2 | SSS160 | PT 14 Magnet MAPS | 295.8 |
| 2 | C1 | MBBR1263 | PT 6.6 Training 6 | 4.95 |
| 4 | C2 | MBBL2071 | ICS 4 Final connection | 294.64 |
| 5 | D1 | SSS249 | Prep 5 Cool Down | 5.02 |
| 7 | D2 | SSS630 | PT 14 Magnet MAPS | 295.18 |
| 10 | E1 | MBAR2229 | Prep 3 CDW1 HV | 295.8 |
| 12 | E2 | CRYO | Cryo leak repair | 295.09 |
| 1 | F1 | MBAL3643 | PT 13.2 Resist. Meas. | 296.88 |
| 6 | F2 | MBAR1253 | ICS 4 Final connection | 293.12 |
| Magnets under Test from 02-10-2005 to 08-10-2005 |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBBR3389 |  |  | 01-10-2005 | 49 h. | 23 h. | 47 h. | 14 h. | 12 h. | 07-10-2005 | 145 h. | Stripping |  |  |
| 2 | TBA2 | MBAL3400 |  |  | 29-09-2005 | 26 h. | 21 h. | 20 h. | 16 h. | 11 h. | 03-10-2005 | 94 h. | Stripping |  |  |
| 3 | TBB1 | SSS642 |  |  | 20-09-2005 | 124 h. | 14 h. | 144 h. | 24 h. | 25 h. | 05-10-2005 | 331 h. | Stand By |  |  |
| 4 | TBB2 | SSS167 |  |  | 01-10-2005 | 34 h. | 18 h. | 17 h. | 9 h. | 8 h. | 05-10-2005 | 86 h. | Stripping |  |  |
| 5 | TBC1 | MBAL3734 |  |  | 02-10-2005 | 31 h. | 31 h. | 18 h. | 15 h. | 15 h. | 06-10-2005 | 110 h. | Stripping |  |  |
| 6 | TBC2 | MBAL2064 |  |  | 02-10-2005 | 61 h. | 20 h. | 15 h. | 15 h. | 7 h. | 07-10-2005 | 118 h. | Stand By |  |  |
| 7 | TBC2 | MBBL1214 |  |  | 30-09-2005 | 39 h. | 3 h. | 0 h. | 8 h. | 0 h. | 02-10-2005 | 50 h. | Stand By |  |  |
| 8 | TBD1 | SSS091 |  |  | 30-09-2005 | 35 h. | 12 h. | 37 h. | 9 h. | 29 h. | 05-10-2005 | 122 h. | Stripping |  |  |
| 9 | TBE1 | MBBL2203 |  |  | 01-10-2005 | 34 h. | 27 h. | 4 h. | 17 h. | 11 h. | 04-10-2005 | 93 h. | Stripping |  |  |
| 10 | TBE1 | MBBL3403 |  |  | 04-10-2005 | 29 h. | 23 h. | 16 h. | 17 h. | 9 h. | 08-10-2005 | 94 h. | Stripping |  |  |
| 11 | TBF1 | MBBL2204 |  |  | 01-10-2005 | 33 h. | 23 h. | 15 h. | 18 h. | 17 h. | 05-10-2005 | 106 h. | Stripping |  |  |
| 12 | TBF2 | MBAL3396 |  |  | 04-10-2005 | 24 h. | 24 h. | 18 h. | 16 h. | 9 h. | 08-10-2005 | 91 h. | Stripping |  |  |
| 13 | TBF2 | MBBR1257 |  |  | 30-09-2005 | 32 h. | 25 h. | 14 h. | 16 h. | 18 h. | 04-10-2005 | 105 h. | Stripping |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBR3389 | 9066 A. (D2_U) | 10221 A. (D1_U) | 11978 A. (D1_U) | 12350.4 A. (D1-L) | 12734 A. (D2-U) | 12850 A. | - | - | - | - | - | - | - |
| 2 | MBAL3400 | 11098.5 A. (D2-L) | 11928.9 A. (D1-U) | 12680.6 A. (D2-U) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 3 | SSS642 | 5799 A. (MQMC_Int) | 5800 A. | 4650 A. | - | - | - | - | - | - | - | - | - | - | - |
| 4 | SSS167 | 12179.3 A. (QE12) | 12429.3 A. (QE34) | 12541.8 A. (QE_34) |  | - | - | - | - | - | - | - | - | - |  |
| 5 | MBAL3734 | 12119.3 A. (D2_L) | 12818.5 A. (D2_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |  |
| 6 | MBAL2064 | 11795.5 A. (D1-L) | 12157 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |  |
| 7 | MBBL1214 |  | - | - | - | - | - | - | - | - | - | - | - | - |  |
| 8 | SSS091 | 11319 A. (QI_34) | 11459 A. (QI_34) | 11850 A. (QE_12) | 11923.3 A. (QI_12) | 11986.8 A. (QI_12) | 12084.1 A. (QE-34) | 12551.2 A. (QE-12) | 12747 A. (QE-12) | 12850 A. | - | - | - | - |  |
| 9 | MBBL2203 |  | - | - | - | - | - | - | - | - | - | - | - | - |  |
| 10 | MBBL3403 | 12569.3 A. (D2-L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |  |
| 11 | MBBL2204 | 11973 A. (D1-L) | 12472 A. (D1-L) | 12818 A. (D1_L) |  | - | - | - | - | - | - | - | - | - |  |
| 12 | MBAL3396 | 10003.1 A. (D2-U) | 12325.1 A. (D2-L) | 11996 A. (D2-U) | 12159 A. (D2-L) | 12303 A. (D1-U) |  | - | - | - | - | - | - | - |  |
| 13 | MBBR1257 | 10613.7 A. (D1_U) | 12171 A. (D1_U) | 12493.8 A. (D1-U) |  | - | - | - | - | - | - | - | - | - |  |

Shows  SSSS work on B1  including painful 10 wire breaks in SSW ; But also long period in connection before cooldown

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_27_A | TBF2 | MB3396: wrong TT821 value. | CRI |  | NC opened. HWREC put off. Missing data in MTF. It will be the case for the next Noell magnet |
| Tuesday_04_N | TBC1 | MB3734 , first training quench. The hfoff analysis show an inverted signal for D1-D2 | MTM | HWREC | Maryline called. Configuration with HWREC was not properly done. |
| Tuesday_04_N | TBD2 | SSL test at 4.55 K for SSS630: analysis of first quench shows a peak voltage out of limit. | MTM |  | Salem contacted. He advised to continue. |
| Tuesday_04_N | TBE1 | MB3403: wrong TT821 value. | MTM |  | Test (IAP@warm) stopped till HWREC is put off. |
| Tuesday_04_N | TBA1 | MB3389: Shaft in aperture 1 not possible to insert. | ICS |  | Stephane asked to proceed with the MM only in aperture 2 to Juan and Jurek to investigate the origin of the problem From Leica measurements it was found that the MRB is laterally displaced by 5 mm. |
| Wednesday_05_A | TBB1 | SSS631 not conform: one corrector dipole circuit was found opened at warm. | CRI |  | Connection stopped. Magnet removed from the bench and sent back to 904. |
| Thursday_06_M | TBD2 | SSW measurements values not correct. | MTM |  | Guy found that the current reading in not correct (10 times lower). Config. Corrected in the Data Base then OK. |
| Thursday_06_N | TBC1 | Megger in MY93 rack doen't work properly. It delivers around 700V instead of 600V asked. | MTM |  | To be repaired. Action Richard. |
| Thursday_06_N | TBC1 | Bad communication between control room sun-station and local sun station during HV test and IAP@warm . | AB/CO |  | Logout on the both sides and re-launch TEMA |
| Thursday_06_N | TBE | IAP @ cold on MB3403: the following potaim cards could not be adjusted: -Vgas_int_lead -D1_L-U. | MTM |  | Potentiometer has no effect. Potaim card Vgas_int_lead changed by Maryline. |
| Friday_07_M | TBC2 | Resistance value of the cryo_heater in MB2064 indicated that the circuit is opened. | MTM |  | NC opened? |
| Friday_07_M | TBB2 | LF could not be armed while launching the powering of the correctors of the SSS160. | ACR/MTM |  | Maryline informed. E_clamp was found broken in the CFB durin the cool-down. A short has been done on the rear of the POTAIM crate. |
| Friday_07_N | TBD1 | LF DAQ could not be triggered automatically during prep2.2 C (test of correctors).. | MTM |  | Maryline advised to trigger the DAQ manually. |
| Sunday_09_M | ALL | Not possible to cool down magnet below 80 K. Communication problem in the process, | ACR |  | Cold test stopped for 4 hours. |
| Sunday_09_N | TBD1 | Quenches (5 times) on the corrector CD circuit. No file because card disabled. | MTM |  | Maryline contacted. She enables the card. |
| Sunday_09_N | TBB1 | Alarm on anticrysotat of SSS629. | MTM |  | Maryline contacted. She asked to load the corresponding parameters for the new AC. |
| Sunday_09_N | TBB2 | Trigger on V_ E_clamp during IAP@warm for SSS349 | ICS? |  | Giorgio checked with a bridge and electronics is o.k. ICS was asked to check the connections. |
| Sunday_09_N | TBA1 | LHC cycle could not be lauched. Missing files in HWREC. | MTM |  | The wrong cycle was selected in the extended test program. Correct cycle was chosen and then it worked. |
| Sunday_09_N | TBA1 | LHC cycle on MB2216 stopped because of a cryo loss. | ACR |  | Flow of current lead went down below the threshold. Protection were activated.. |