### Week 27 starting 3 July 06

### SM18 Operation Meeting (Tuesdays) 04 July 06

Bench status  on Monday  03-07-2006 at:  7/3/2006 11:14:48 AM

2 A1 MBBR2253 PT 12 Warm Up 5.59 8 A2 MBAR1295 Prep 5 Cool Down 93.98 9 B1 SSS270 ICS 4 Final connection 301.17 1 B2 SSS527 Prep 5 Cool Down 1.90 3 C1 SSS271 PT 12 Warm Up 292.94 7 C2 SSS337 PT 12 Warm Up 16.34 12 D1 SSS634 ICS 5 Traveller Last 299.65 6 D2 SSS682 PT5.4: Powering Corrector EF 4.51 10 E1 MBBR2319 Prep 5 Cool Down 228.39 4 E2 MBAR1299 PT 9 CDAP HV 1.90 11 F1 MBAR2271 Prep 5 Cool Down 291.33 5 F2 MBAR1360 PT 5.1 1.5 KA Quench LF 1.91

CONFIG :  D6- AS3 –IRSSS3;

Supervisor: REMONDINO Vittorio/

Eq support (Salem till Friday and Giorgio Sat/Sun) : 16-3602 :

17 Magnet Tests completed :

9 Dipoles (2303 (61h.), 1292 (72 h.) with Magnetic Mesurment),

6 Arc-SSS  (271 2 times thermal cycle, 319 2nd run Stripping)

1  500-series [ 502 (167 hrs) ]

1  600-series[ 656 1° run in March “He leak of magnet/cfb circuits”  2nd run (114 hrs) ]

(13 Fresh Magnets sign for “Stripping”)

General Points :

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- News /Situation from (SSS, SSSS)  :  J Halik

- Situation SSSS  :OP issues :

- A.O.B..

# From 6/25/2006 to 7/1/2006

Tests Information

| Magnets under Test from 25-06-2006 to 01-07-2006 |
| --- |
| - | Bench | Magnet | Shafts | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |  |
| 1 | TBA1 | MBBR2303 |  | 22-06-2006 | 35 h. | 28 h. | 61 h. | 15 h. | 6 h. | 28-06-2006 | 145 h. | Stripping |  |  |  |
| 2 | TBA2 | MBBL1390 |  | 24-06-2006 | 29 h. | 38 h. | 34 h. | 21 h. | 17 h. | 30-06-2006 | 139 h. | Stand By |  |  |  |
| 3 | TBB1 | SSS295 |  | 27-06-2006 | 29 h. | 16 h. | 29 h. | 8 h. | 11 h. | 01-07-2006 | 93 h. | Stand By |  |  |  |
| 4 | TBB1 | SSS319(2) |  | 24-06-2006 | 0 h. | 0 h. | 47 h. | 8 h. | 14 h. | 27-06-2006 | 69 h. | Stripping |  |  |  |
| 5 | TBB2 | SSS502 |  | 21-06-2006 | 69 h. | 14 h. | 62 h. | 10 h. | 12 h. | 28-06-2006 | 167 h. | Stripping |  |  |  |
| 6 | TBC1 | SSS271(3) |  | 24-06-2006 | 25 h. | 23 h. | 30 h. | 8 h. | 0 h. | 28-06-2006 | 86 h. | Thermal cycle |  |  |  |
| 7 | TBC1 | SSS271(3) |  | 28-06-2006 | 0 h. | 12 h. | 51 h. | 11 h. | 0 h. | 01-07-2006 | 74 h. | Thermal cycle |  |  |  |
| 8 | TBC2 | SSS097 |  | 21-06-2006 | 24 h. | 31 h. | 23 h. | 8 h. | 13 h. | 25-06-2006 | 99 h. | Stripping |  |  |  |
| 9 | TBC2 | SSS273 |  | 25-06-2006 | 37 h. | 13 h. | 18 h. | 7 h. | 10 h. | 29-06-2006 | 85 h. | Stripping |  |  |  |
| 10 | TBD2 | SSS656(2) |  | 23-06-2006 | 38 h. | 21 h. | 29 h. | 12 h. | 14 h. | 28-06-2006 | 114 h. | Stand By |  |  |  |
| 11 | TBE1 | MBAR1389 |  | 21-06-2006 | 23 h. | 24 h. | 28 h. | 14 h. | 22 h. | 26-06-2006 | 111 h. | Stripping |  |  |  |
| 12 | TBE1 | MBBR2324 |  | 26-06-2006 | 30 h. | 35 h. | 20 h. | 13 h. | 15 h. | 01-07-2006 | 113 h. | Stand By |  |  |  |
| 13 | TBE2 | MBBR1277 |  | 21-06-2006 | 39 h. | 25 h. | 18 h. | 18 h. | 104 h. | 30-06-2006 | 204 h. | Stripping |  |  |  |
| 14 | TBF1 | MBAR1352 |  | 19-06-2006 | 36 h. | 34 h. | 51 h. | 15 h. | 31 h. | 26-06-2006 | 167 h. | Stripping |  |  |  |
| 15 | TBF1 | MBBR2334 |  | 26-06-2006 | 29 h. | 43 h. | 19 h. | 17 h. | 11 h. | 01-07-2006 | 119 h. | Stripping |  |  |  |
| 16 | TBF2 | MBAL1395 |  | 25-06-2006 | 42 h. | 26 h. | 28 h. | 20 h. | 8 h. | 30-06-2006 | 124 h. | Stripping |  |  |  |
| 17 | TBF2 | MBAR1292 |  | 18-06-2006 | 27 h. | 27 h. | 74 h. | 16 h. | 23 h. | 25-06-2006 | 167 h. | Stripping |  |  |  |

Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBR2303 | 9853.4 A. (D1_U) | 10047.9 A. (D1_U) | 11682.6 A. (D2_L) | 10949.3 A. (D2_L) | 0 A. (No Data) | 0 A. (No Data) | 12695 A. (D2-L) | 12850 A. | - | - |
| 2 | MBBL1390 | 11969.6 A. (D2-L) | 12405 A. (D1_U) | 12653.5 A. (D2-L) |  | - | - | - | - | - | - |
| 3 | SSS295 | 12105 A. (QI_34) | 12698.8 A. (QE_12) | 12820 A. (QE-34) |  | - | - | - | - | - | - |
| 4 | SSS319 | 11784.2 A. (QE-12) | 12058.1 A. (QE-34) | 12200 A. (QE_12) | 12501.2 A. (QE-34) |  | - | - | - | - | - |
| 5 | SSS502 | 12025 A. (No Data) | 12346.4 A. (QE_34) | 12435 A. (QE-24) | 12620.7 A. (QE_34) |  | - | - | - | - | - |
| 6 | SSS271 | 10922.9 A. (QI_12) | 11145.2 A. (QI - 12) | 11737.5 A. (QI-34) | 11855.2 A. (QI-12) | 12156 A. (QI-12) | 12239 A. (QI_34) | 12336 A. (QI_34) | 12396 A. (QI_12) | 12430 A. (QI_12) | 12625.6 A. (QE-12) |  |
| 7 | SSS271 | 11429.9 A. (QI_12) | 11548.2 A. (QI_12) | 12203.3 A. (QE_34) | 12318.9 A. (QI_34) | 12415 A. (QI-12) | 12407 A. (QI-34) | 12550 A. (QE_34) | 12663.1 A. (QI- 12) | 12700.7 A. (QI- 12) | 12737.9 A. (QI_12) |  |
| 8 | SSS097 | 12319.8 A. (QI_12) | 12650.3 A. (QI_34) | 12850 A. | - | - | - | - | - | - | - |  |
| 9 | SSS273 | 12647.6 A. (QE-12) | 12838.4 A. (QE-34) | 12850 A. | - | - | - | - | - | - | - |  |
| 10 | SSS656 | 5714.3 A. (Ext_MQMC) | 5850 A. | - | - | - | - | - | - | - | - |  |
| 11 | MBAR1389 | 12037.8 A. (D1_U) | 12427.7 A. (D1_U) | 12738 A. (D2-U) |  | - | - | - | - | - | - |  |
| 12 | MBBR2324 | 11679 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - | - |  |
| 13 | MBBR1277 | 12258 A. (D2_U) | 12330.6 A. (D2_U) | 12380.2 A. (D1_L) |  | - | - | - | - | - | - |  |
| 14 | MBAR1352 | 12482.3 A. (D2_L) | 12463.1 A. (D2_L) | 12661.8 A. (D2_L) |  | - | - | - | - | - | - |  |
| 15 | MBBR2334 | 10678 A. (D2-L) | 12263 A. (D1-U) | 12503.8 A. (D1_L) |  | - | - | - | - | - | - |  |
| 16 | MBAL1395 | 12684.2 A. (D1-U) | 12828.6 A. (D2_U) | 12771.2 A. (D1_L) | 12619.2 A. (D1_L) | 12850 A. | - | - | - | - | - |  |
| 17 | MBAR1292 | 12078.7 A. (D2-L) | 12827.1 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - |  |

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_27_M | TBB2 | SSS502:problem HV test at cold on MQTL : condensation water on the current leads | ACR |  | TC asked cryo team for action. |
| Tuesday_27_M | TBF1 | MB2334 : Prep2.1 &Prep2.2 could not be performed. | AB/CO | Lf, hf | Tests were successful only after resetting & restarting PLCSE, LF &HF Acquisition and TEMA before both the tests. |
| Tuesday_27_A | TBA1 | MB2303: Overr-ange error appearing in results of MM3, in sector 0203 and 0809 in aperture 2. | MTM |  | Equipment support is contacted as per recommendation by TC, and the problem is being investigated. Guy suggested that MM3 can be repeated after the electronics is checked, and to quit MM and proceed with power tests if it fails in that run also. Guy tuned the electronics, and MM3 is being repeated. |
| Tuesday_27_A | TBD1 | SSS634 : MM3 Fault on coil head. |  |  | Guy rectified the fault on coil head and performed MM1. MM3 is being repeated on his recommendation. |
| Tuesday_27_A | TBA1 | MB2303 : Shaft position error along with over-range errors in MM3; in s001, s0203 and s0809. TC contacted. | MTM |  | Advised to proceed, treating the results as OK. |
| Tuesday_27_A | TBD1 | SSS634 : Training quench: AQA Alert TauInitV>117 during PT6.11 (resistance of one quench heater high). | MTM | AQA | TC came and checked the signals. He recommended proceeding with MM2. |
| Tuesday_27_N | TBD2 | SSS65 : problem to launch HV WP27 with the mobile rack 93: PLC_HV and PLCSE doesn't allow to start the test. | MTM |  | Mobile HV rack was changed by rack 93 and it was ok. |
| Wednesday_28_M | TBF2 | MB1395 : Sensor problem on F2, no Cryo OK | ACR |  | Cryo problem solved after 30’ |
| Wednesday_28_N | TBC1 | SSS271 : PT-1,part-III, patch cable used for the AC signal injection is insufficient (in length?).. | MTM |  | Equip.Support is asked to replace the cable |
| Wednesday_28_N | TBF2 | MB1395: Magnet crossed 12500A in first attempt itself. However, during PT6.3 which was supposed to be the warm-up quench, it showed detraining by more than 50A. Hence another quench is being performed. Magnet further detrained. | MTM |  | TC contacted after a second detraining quench. He asked to perform another one to confirm or infirm the trend. If the magnet will train again it will be warmed up. |
| Wednesday_28_N | TBB2 | SSS527: Clamp connection checking of correctors could not be done as there was permanent trigger on the safety matrix. | MTM |  | All connections checked and repeatedly tried to carry out the test but it could not be done. Equipment support will be called. In the morning. |
| Wednesday_28_N | TBE1 | MB2324: Magnet cooled down only for leak test. CRYO team could not signed the e-traveler because such a configuration is not foreseen. | ACR |  |  |
| Wednesday_28_N | TBE1 | MB2324 : PT1 could not be launched as the DAQ is getting a permanent trigger. | MTM |  | Connections were checked and found OK. Equipment support will be contacted. |
| Thursday_29_M | TBB2 | SSS527 : Permanent triggers existing on potaim cards 9 and 11. | ICS |  | Richard investigated and observed that ICS connected the new VT wire on both current leads. Connection seems to be OK. But the problem persists. Giorgio found that the wires were connected to the old VT in the plate-form. |
| Thursday_29_M | TBC1 | SSS271: Magnet under Thermal Cycle. Three training quenches done in last shift. Thermal cycle rules were not followed. As per training rules, after second quench, magnet stands NOT ACCEPTED. But a third training quench was done in the Night shift. NB: The rules of a thermal cycle allow a maximum of 5 quenches. In any-case the test coordinator has the possibility to extend the training depending on the results. | MTM |  | TC contacted in the morning. He came and checked the results. He advised to make fourth quench as anyway. TC asked to continue the training up to 9 quenches or ultimate. Decision to make a second thermal cycle will be taken by TC after that. |
| Thursday_29_M | TBE1 | MB2324: Test IAP at cold PT1 - AC compensation at cold not possible. Adjustment on potaim card 1 and 3. not possible. | MTM |  | One voltage tap ( EE211) was found open. Richard connected redundant tap. He will open NC. Further tests being continued. |
| Thursady_29_N | TBB1 | SSS295: MQT (EF circuit) quenched 4 times and was still well below the ultimate. | MTM |  | TC advised to make 9 quenches and stop the training in EF circuit. NC has to be opened. Then other tests can be performed. |
| Thursady_29_N | TBC1 | SSS271: Problem in starting the server. Hard disk of server is full and TEMA is unable to launch any application. | MTM |  | Solved by Maryline. |
| Thursady_29_N | TBA1 | MB2253: Difficulty to place the plexi-panel (for P10,C10,C11 connectors at the bottom of bench). | MTM |  |  |
| Friday_30_M | TBA2 | MB1390: PTE (diode test) forgotten. Magnet warm. | MTM |  | TC informed. Magnet put in stand_by. TC informed Pierre for treating the non conformity.. |
| Saturday_01_N | TBA1 | MB2253 : Magnet name, temperature and Coil number were not correctly read in MMP program run from local server | MTM | MMP | Data validated again in MOBIC and MMP repeated with correct configuration. |
| Saturday_01_N | TBA1 | MB2253 : Error in PT1, Part II for channel D1_C_S03_S05 (chassis2 card3 channel 1): Potentiometer could not be adjusted. | MTM |  |  |
| Sunday_02_M | TBB2 TBB1 | SSS527 : Q heater 3 shows overcharged (1200V) in PLC reading. Further training could not be launched due to PLCSE fault. SSS270 : Warm tests are not possible because the dummy quench heater adapter of B1 is used on B2 bench | MTM |  | Guy informed. Guy replaced the discharge resistor adaptor with the one from B1, as well the 3d QH power supply. The spare QH power supply was not working. Guy replaced that QH power supply 3) by a QH power supply of the D cluster. It will be replaced Monday. |
| Sunday_02_A | TBA1 | MBB2253 : During the training, slow PA occurred. Origin? No fault reported from the cryo | MTM |  |  |
| Sunday_02_N | TBD2 | SSS682 : OP team sent the message to cryo team to cool down the SSS at 1.9 K. Why?. In the board it is clearly written that the SSS worked at 4.55 K. Cryo team tried to cool down the magnet at 1.9 K without exchanger? | MTM ACR |  | Time lost. |
| Monday_03_M | TBD2 | SSS682: Leak found at cold on the capillary. Magnet could not be qualified. | ACR |  | TC informed. Giorgio will make some electrical checks and the SSS will be warmed up for leak investigations. |