Tss23-Feb-04

### Week 10  starting 23 Feb 04

### SM18 Operation SCHEDULE Meeting (Tuesdays)

Bench status  and tests for this week :

Tested : 3031 (TBE2);3019 (TBE1);3087 (TBD1);3064 (TBF2); 3090 (run1)

- Priority list of magnets to be tested available in the cryodipole web page. To come : 2027 retested and 1047 (w.o anticryostat).

- Shafts in 3090 (C1); 206 (TBF1); 2050 (TBD2) then in 2046 (TBF2).

Test Coordinator this week:  Marco 160621 ;  MTM  Tech. on call Ewald:  (16-3602) .

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Leaks: 3090 (TBC1) bellows; 3019 (TBE1) Joint or detection system,  3031 (TBE2) Joint or detection system.

- Wednesday night: Lost of communication in PLC of the CFB. CFB not controlled. Restart Thursday morning.

- Friday : stop of the CWS for reconditioning ( loss to high)

- TBF2 (3064) : Discrepancy between TT821 (302 K ) and Tcryo (170 K).

- Other General Issues

  - TBD1 (1037) : Insulation  dipole vs. ground, good at warm, failed below 250 K. End of the story.

  - SSS06: Short between sextupole and octupole correctors

  - General water cut : Wednesday afternoon.

  - 3090: Thermal cycle executed without going to 9 T or executing  9 ramps.

| Date | Problem | Type | Application/soft. | Action/advise |
| --- | --- | --- | --- | --- |
| Wednesday_18_A | TBD1 (3087): IAP@warm : energy of 1 LF QH less than the limit | MTM | IAP | Giorgio informed. Fit in IAP to be improved. Done by Boris. |
| Wednesday_18_N | TBD2: No magnet names in TEMA | MTM | TEMA | Solved by Giorgo. Scan and run of test . |
| Wednesday_18_N | TBD2: No Temperature data for TT821 IAP not working properly in Sun MTA37 | AB/CO | IAP | Advise of Giorgio: LF DAQ slot 6 checked : T=279 K. Boris informed. Use : another machine for analysis |
| Thursday 19_N | TBE2: change of Iearth during the quenches. | MTM |  | Magnet maintained at 1.9 K. Pierre’s instruction: HV test performed : OK Additional quench performed. OK then warm up |
| Friday 20_A | TBC1: DAQ not getting armed for IAP Abort and re-launched. No success. | AB/CO |  | Solution : Reset VME crates then DAQ got armed. |
| Friday 20_N Sunday 22_A | TBF2: Slow PA launched. Failed. QH discharged. TBD2: same problem. | MTM |  | Giorgio advised to check whether the 2 A cable was connected to the magnet. Check all the connection: Found that D20 not connected. |
| Saturday 21_A | TBF2 (3064): Discrepancy between reading of TT821 (302 K) and Tcryo (170 K) | MTM/ACR |  |  |
| Saturday 21_A | TBD1: Problem with AC compensation of potaim cards D1-L-U,D1-D2 | MTM | IAP |  |
| Saturday 21_A | TBE2 (3077): no anticryostat stickers “no Anti Cryostat”is not recognised by the system during the scan. | MTM | HWREC | Solution proposed by Giorgio: leave the field empty. |
| Saturday 21_N | Problem with warning boxes”high currents” in particular for D1. | MTM |  | Action launched by Giorgio: Improvement of the fixation of these boxes. |
| Sunday 22_A | TBE2 (3077): Name in the sticker on the magnet wrong : impossible to start TEMA. | MTM |  | Giorgio came and corrected the name. |
| Sunday 22_N | TBF1: Sun MTA 28 blocked. | MTM |  |  |
| Sunday 22_N | TBE2: PT10 Min energy :LF DAQ armed but not the correct leds are glowing Probable origin : no exit for all applications in particular LF web. | MTM |  | Advise by Giorgio : reset and re-launch. |
| Monday 23_M | TBE2: Unable to put 2 A for warm up. Jack to be replaced in cable connected to the current leads. | MTM |  | Advises : -Be gentle with the cables: remove it using the connectors. - If something broken : pass the message in order to replace it asap. |