Vc/29-Mar-04

### Week 14  starting 29 March 04

### SM18 Operation Meeting (Tuesdays)

Bench status  and tests for this week :

| Bench Name Status Monday 1730 Temp - - TBB1 MBAR1051 Commissioning 300 K TBB2 MBBL3122 " 300 K TBC1 MBAR2016 Prep 1 Lyre Test 300 K TBC2 SSS0013 PT 7.1 detrain check 1.9 K TBD1 MBAL1058 CR 1 Leak Test 300 K Try Keep free After TBD2 MBBL3033 Prep 5 Pumping Cool Down Cooling TBE1 MBBL2063 PT 6.1 Training 1 1.9 K TBE2 MBBR3018 ICS 3 after warm test 300 K TBF1 1061 NEW Last night 300 K TBF2 1055 NEW Last night |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |

Test Coordinator this week:  Pierre Pugnat; MTM support on call  Federico:  (16-3602).

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Other General Issues

  - Keep D1 free for SSS Mech Preparation ( 6 hrs ? )   J.Halik

  - PLC For Water & Federico/Alessandro This week at night ?

| Date | Problem | Type | Application/soft | Action/advise |
| --- | --- | --- | --- | --- |
| Tuesday 23_M | TBF1: Difference in T magnet (2011). TT821 reading : 5.1 K. T cryo: 1.9 K HV test failed. | MTM |  | TT 821 sensor is shorted to the ground. |
| Tuesday 23_M | TBE: TEMA locked (Florence?). | AB-CO | TEMA | Giorgio called and unlocked TEMA |
| Tuesday 23_A | TBF1: Min Energy@11850 A analysis: Peak voltage limit more than 100 V for D2-L. D2-U is flat. | MTM |  | One HF acquisition cable replaced by Giorgio (flat curve). Mail reported to P.Pugnat as specialist (V too high). |
| Tuesday 23_A | TBD2: “Error at CPU address during the encoder pulse reception” during the machine cycle of the 2005. | AB-CO |  | On acknowledgment, the error disappeared. |
| Wednesday 24_N | TBE2: Temperature (TE148) sensor problem, which pilots the phases. | ACR |  | Conditioner changed |
| Wednesday 24_N | TBD2: After the quench for warm up, the vacuum degraded (reached 3 mbar). Origin not known. | ACR |  | Cryo team installed a mobile pump but temperature was falling. Vacuum recovered after 3 hours. Giorgio called and advised not to remove shaft as they were already at 200 K. They were removed after the warm up and re-calibrated. |
| Thursday 25_A | TBD1: HV test at cold: problem with SunMTA36. | AB/CO |  | Giorgio and Richard called Pedro. Test performed on TBE sun. RS232 transducer changed. |
| Friday 26_A | TBD2: HV test @ warm done by ICS for the 3033 magnet. Test OK but file not created or not transferred? | MTM/AB-CO |  | Richard is investigating. |
| Friday 26_A | TBD2: Launching IAP@warm, HWREC error message:” configuration files contain Oracle errors “ Test done with previous LF config file: Problem with T sensor (263 K). | MTM | HWREC | Giorgio informed. He came and created the proper config file (1 Qloc crate missing from the bench because of TB-B commissioning). Test OK. |
| Friday 26_N | TBF2: PC fault (water) before launching the SPA. Cannot reset the alarm. | AB/PO |  | Giorgio investigated and discovered a water flow too low internal to the PC (RMF). OK. |
| Sunday 28_A | TBD2: Configuration with anti-cryostat, flow of N2.But magnet without anti cryosytat. | MTM/ACR?? |  | To be clarified. |
| Sunday 28_A | Calibration bench: Water leak. Valves closed | MTM |  | Intervention of J.Halik Monday |
| Monday 29_M | TBF1: IAP @warm final .3 faulty files created (magnet 1078) | MTM |  | Files to be erased. Stop of acquisition forgotten after IAP? |
| Monday 29_M | TBE1: Training quench stopped, General problem of pumping. Warmer H1 down few minutes without alarms. Then it restarted but pressure was rising in the lines. | ACR |  | Now OK but investigation on the run to understand the origin of the stop. |