Vc/13-May-04

### Week20 starting 10 May 04

### SM18 Operation Meeting (Tuesdays)

Bench status  and tests for now & next few days :

| Bench | Name | STATUS 10/05 afternoon |  |
| --- | --- | --- | --- |
| TBA1 | _ | Not Activated |  |
| TBA2 | 3117 ? | Commissioning & RETEST | This week ? |
| TBB1 | MBAR2039 | Prep 1 Lyre Test | 300 K |
| TBB2 | MBAR3024 | PT 12 Warm Up | Warming |
| TBC1 | MBAL3148 | PT 12 Warm Up | Warming |
| TBC2 | SSS015 | PT 12 Warm Up | Warming |
| TBD1 | MBBL2061 | Prep 5 Pumping + Cool Down | Cooling |
| TBD2 | MBBL2078 | OUT OF USAGE | 300 K |
| TBE1 | MBBR3016 | PT 1 IAP @ cold | 1.9 K |
| TBE2 | MBBL1114 | PT 12 Warm Up | Warming |
| TBF1 | MBBL3135 | CR 1 Leak Test | 300 K |
| TBF2 | MBBL1097 | ICS 5 Traveller Last |  |

Test Coordinator this week: Marco ; MTM support on call  Giorgio:  (16-3602).

10 Dipoles  Disconnected after tests last week    ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

    - Repair of D2 Needs All benches above 80’ K

    - 36 hrs required ?    Plan warm up evening so work starts morning ?

A2  commissioning & re-test 3117 ?

- General Problems last week :  Cable rupture on F1  for Anticryostats

- SSS work :   GREY FILE IN SMCR  ….   To do list etc Nair + Sridhar ,

| Date | Bench | Problem | Type | Application/soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday 04_N | TBD1 | No possibility to close or open the electro-valves on the benches | MTM |  | Guy came and realised that the PLC valves control was not working. He opened the valves manually. Origine: Electrical overcharge which burned the fuse. Now the system is working (Federico) |
| Tuesday 04_N | TBD1 | 2049 magnet: oscillation present during the discharge of HF heaters at 1500 A. | MTM |  | 3 kA min energy for diode test performed: diagnostic confirmed. Test was stopped. Magnet was warmed up Wednesday afternoon because the QH YT211 is not working. |
| Wednesday 05_M | TBC1 | IAP@warm: HWREC problem which was solved. TEMA did not load the proper configuration | MTM | HWREC+TEMA | Many attempt. Still not OK. Reset of the racks then OK. |
| Wednesday 05_M | TBC1 | HV test stops. Launch 3 times run-recovery without success. | MTM |  | Rack changed. |
| Wednesday 05_A | ALL | SMTMS temperatures were not refreshed. | ACR |  | J.Axensalva working on the problem |
| Wednesday 05_A | TBD1 | 2049 HV test @cold: Lost of the cryo OK . | ACR |  |  |
| Wednesday 05_N | TBF1 | HF config. File not found. No data was found during the second training quench | AB-CO | HF |  |
| Tuesday 06 _M | TBE1 | 3016: Mobic did not accept the scanned data. | MTM? | HWREC | Wrong data in MTF for the anti-cryostats. |
| Tuesday 06 _M | ? | IAP@warm not possible: no trigger. | AB-CO | TEMA | TEMA re-started . Then OK |
| Tuesday 06 _A | TBF1 | Magnetic measurements: Wrong power supply selected (TBB). | MTM | Power | Power supply changed manually. |
| Tuesday 06 _A | TBE1 | HV test at warm on 3016. Problem with the first sequence. | MTM | HV | Richard informed. He said that the capacitance values reach the limits. Test is OK |
| Tuesday 06 _N | TBF1 | Full load line stopped because power supply went off. | AB-PO?? |  | Cycle re-started. |
| Tuesday 06 _N | TBC2 | Power supply could not be switched on. | MTM? |  | Guy came and found the digital communication cable was disconnected on PC. He also reboots the sun MTA19. |
| Tuesday 06 _N | TBC1 | HRWEC problem: INTERCOM missing. | ? | HWREC | Guy created the files manually. HWREC off for this test. |

| 8Sl. No | Date /Day | Shift | Bench/Cluster | Magnet | Problems | Remarks |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 8/5/04- Sat | N | D1 | MBBL 2061 | TBD1: leak problem 1*10-3 on the manchette. ICS will try to repar tomorrow. |  |
| 2 | 8/5/04- Sat | N | F1 | MBBL 3135 | TBF1 : The long anti cryostat cable was broken, ICS missing to disconnect this cable, when rocla came the wire was broken. |  |
| 3 | 8/5/04- Sat | N | Rocla Problm | --- | Rocla problem (bolding) a responsable (Pascal) is comming, the 1st rocla have a bolding problem, the 2nd will be repared for its electrical problem, monday morning. |  |
| 4 | 8/5/04- Sat | N | E2 | MBBL 1114 | TBE2 : Always a temperature sensor problem, a 0-200mA power supply has been taken in work shop. But not possible to have cryo OK. |  |
| 5. | 8/5/04- Sat | M | F2 | MBBL 1097 | TBF2: Anticryostat alarm after breaking the vacuum - temperature dropped to 0C. Temp. started rising after a while. | PLC reset was done. |
| 6 | 8/5/04- Sat | M | E2 | MBBL 1114 | Cryo temperature sensor (TT 130B) on one of the current leads not working. So cryo O.K. signal lost. | To be solved by cryo equipment support |
| 7 | 7-8/5/04 Fri-Sat | N/M | F2 | MBBL 1097 | Safety Matrix Rack on which Min Energy card is mounted is receiving permanent triggers. So 3KA diode test could not be carried out. The Min NRG card was disabled by morning shift. The diode test can be carried out in the afternoon shift | No triggers are seen when bench F1 is selected. |
| 8 | 7/05/04 Fri | N | E2 | MBBL 1114 | Problem of temperature sensor on the CFB, the current lead are freeze. |  |
| 9 | 7/05/04 Fri | N | F2 | MBBL 1097 | TBF2: Equipment Support had been called for a trigger always present on matrix Min. NRJ card, couldn't be solved, we continue the training because we can do this Diode Test tomorow. |  |
| 10 | 7/05/04 Fri | N | E2 | MBBL 1114 | TBE2: The problem will be resolved tomorow, because it is not possible to pass on an auxiliary sensor temperature, to give the cryo ok 1.9K. |  |
| 11 | 7/05/04 Fri | N | F2 | MBBL 1097 | TBF2: first training quench 2 files genereted (.a01 only for HF), full data in .a02 |  |
| 12 | 6/05/04 Thu | N | F1 | MBBL 1043 | TBF1: Full load line stopped and power supply off so we had to restart again |  |
| 13 | 6/05/04 Thu | N | C2 | SSS | TBC2:problem for switching on the power supply so we called Guy. |  |
| 14 | 6/05/04 Thu | N | C2 | SSS | TBC2: digital communication cable was disconnected on PC. Reconnected then OK |  |
| 15 | 6/05/04 Thu | N | C2 | SSS | TBC2: guy reboot sunmta19 |  |