Vc/28-Jun-04

### Week27  starting 28 June  04

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday 28 june 23.00 hrs

| TBA1 | MBAR2042 | CR 1 Leak Test | 302.1 |
| --- | --- | --- | --- |
| TBA2 | MBBL1123 | ICS 2 Connect Magnet (ICS) | 301.9 |
| TBB1 | MBBL2077 | ICS 4 Final connection | 301.3 |
| TBB2 | MBAR2523 | ICS 5 Traveller Last | 301.9 |
| TBC1 | MBAL1089 | Prep 2.4 Res. Meas. CDW1 | 300.8 |
| TBC2 | SSS016 | Prep 5 Pumping + Cool Down | 3 HV |
| TBD1 | SSS020 | New Magnet arrived | 301.4 |
| TBD2 | SSS037 | MM 4 Stretched Wire | 1.9 |
| TBE1 | MBAR2030 | PT 13.2 Resist. Meas. | 300.3 |
| TBE2 | MBBL1121 | ICS 6 Disconnect Magnet WP07 | 301.9 |
| TBF1 | MBBL3132 | PT 6.1 Training 1 | 1.92 |
| TBF2 | MBBL1126 | ICS 4 Final connection | 301.6 |

6/28/2004 10:52:19 PM    Temperature data are taken from TBSM18 site

Test Coordinator this week: marco  ; MTM support on call :  (16-3602).

8 Magnets( 7 dipoles+ 1 SSS)   Disconnected after tests last week    ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

  - SSS work :

Bench status  and tests for now( Monday 21/6 at 23.00)  & next few days :  see Video Display

Active Magnets page

Magnet TB Arrival Out Shafts WP04 HV CDW1 Res. CDW1 HV CDBP HV CDBP Res. CDPT CDAP HV CDW2 Res. WP07 HV MBAL1089 TBC1 27/06/2004 Tests Sequence MBAR2030 TBE1 23/06/2004 Tests Sequence MBAR2042 TBA1 28/06/2004 Tests Sequence MBAR2523 TBB2 21/06/2004 Tests Sequence MBBL1121 TBE2 18/06/2004 Tests Sequence MBBL1123 TBA2 28/06/2004 Tests Sequence MBBL1126 TBF2 25/06/2004 Tests Sequence MBBL2077 TBB1 26/06/2004 Tests Sequence MBBL3132 TBF1 23/06/2004 Tests Sequence SSS016 TBC2 24/06/2004 Tests Sequence SSS020 TBD1 25/06/2004 Tests Sequence SSS037 TBD2 15/06/2004 Tests Sequence

# Results by Week of week 26

| Magnets under Test from 20-06-2004 to 26-06-2004 |
| --- |
| - | Bench | Magnet | Shafts | Start Date . | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total |  |  |
| 1 | TBA2 | MBAR2015 |  | 15-06-2004 | 33 h. | 67 h. | 46 h. | 15 h. | 5 h. | 22-06-2004 | 168 hours |  |  |
| 2 | TBB1 | MBAR2031 |  | 17-06-2004 | 27 h. | 58 h. | 81 h. | 14 h. | 39 h. | 26-06-2004 | 216 hours |  |  |
| 3 | TBB2 | MBBL1120 |  | 14-06-2004 | 26 h. | 38 h. | 89 h. | 13 h. | 2 h. | 21-06-2004 | 168 hours |  |  |
| 4 | TBC1 | MBAL3140 |  | 14-06-2004 | 33 h. | 47 h. | 46 h. | 21 h. | 16 h. | 21-06-2004 | 168 hours |  |  |
| 5 | TBC2 | SSS31 |  | 10-06-2004 | 64 h. | 22 h. | 95 h. | 15 h. | 75 h. | 21-06-2004 | 264 hours |  |  |
| 6 | TBE1 | MBBL1102 |  | 16-06-2004 | 28 h. | 47 h. | 58 h. | 25 h. | 4 h. | 23-06-2004 | 168 hours |  |  |
| 7 | TBF1 | MBBL1097 |  | 14-06-2004 | 40 h. | 51 h. | 115 h. | 4 h. | 2 h. | 23-06-2004 | 216 hours |  |  |
| 8 | TBF2 | MBAL3142 |  | 15-06-2004 | 41 h. | 78 h. | 78 h. | 13 h. | 18 h. | 25-06-2004 | 240 hours |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAR2015 | 10896 A. (D1-U-S01) | 11570 A. (D1-L-S01) | 11935 A. (D2_L_S01) | 12693 A. (D1_L_S01) | 12849 A. | - | - | - | - | - | - | - | - |
| 2 | MBAR2031 | 9766 A. (D1_L) | 10957 A. (D2-U) | 11678 A. (D1-U) | 12080 A. (D2-L) | 12346 A. (D2-U) | 12094 A. (D2-L) | 12657 A. (D1_L) | 12602 A. (D1_L) | 12850 A. | - | - | - | - |
| 3 | MBBL1120 | 10287 A. (D2_U) | 11454 A. (D1-U) | 12096 A. (D2-L) | 12288.3 A. (D1 L) | 12356 A. (D2_L) | 12443 A. (D1_L) | 12581 A. (D1-L) | 12631 A. (D2-L) | 12731 A. (D1_U) |  | - | - | - |
| 4 | MBAL3140 | 11464 A. (D1 U) | 11100 A. (D2-U) | 12152 A. (D2-U) | 12280 A. (D2-U) | 12471 A. (D2_U) | 12656 A. (D2_U) | 12815 A. (D1-U) | 12850 A. | - | - | - | - | - |
| 5 | SSS31 | 12310 A. (E_12) | 12536 A. (I-12) | 12565 A. (I-12) | 12714 A. (I-34) | 12833 A. (I-34) | 0 A. | - | - | - | - | - | - | - |
| 6 | MBBL1102 | 11614 A. (D2-L) | 12173 A. (D1_U) |  | - | - | - | - | - | - | - | - | - | - |
| 7 | MBBL1097 | 11393 A. (D1-U-S13) | 11743 A. (D2-L-S13) | 11995 A. (D1-U-S12) | 12404 A. (D2-U-S13) | 12667 A. (D1-U-S13) |  | - | - | - | - | - | - | - |
| 8 | MBAL3142 | 10436 A. (D2-U) | 11853 A. (D1_U) | 12578 A. (D2-L) | 12748 A. (D2-U) | 12849 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - |

Back to Homepage

- General Problems last week :

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _22_A | TBD2 | HV test on orbit corrector on SSS37. Test failed at 1500 A even after condensation removal. | MTM |  | Richard launched the special sequence “dipole orbit investigation”. Problem on the dipole orbit external. Test repeated Thursday morning with an air blower on since tuesday then OK. Humidity on the current lead. |
| Tuesday _22_A | TBB2 | 2053 magnet: IAP at warm (LF) with the name of the old magnet. Wrong name also in TEMA. Scan of the magnet forgotten. | MTM |  | Test repeated |
| Wednesday_23_M | TBF1 | 1097 magnet : ICS opened the manchette with vacuum not broken. OK was not given by cryogenic team through the e-traveller, Fortunately the magnet was warm! | ICS | Task Tracking application | J.P. Lamboy informed by Stephane. He is investigating. |
| Wednesday_23_M | TBC | Water leak detected on RMC | AB/PO |  | Problem solved by AB/PO |
| Thursday _24_M | TBB1 | Fast de-excitation test on 2031 magnet: quench in Vbus-ext at 141 Amps. | MTM, ACR? |  | Origin investigated by Richard: potaim card “Vbusext” changed and then OK. Put the old potaim card back. It worked. Cryo problem? |
| Thursday _24_M | TBC1 | IAP at cold: for pot-aim card compensation no data in the HF file. | AB/CO | hf | Application quit and reset of crates. Test repeated and then OK. |
| Thursday _24_N | TBB2 | The data of the anticryostats were not all present and not correct for the long ones | MTM |  | P.Legrand requested to redo a scan. |
| Thursday _24_N | TBC1 | 2nd training ramp on the 2074 magnet. SPA at 11463 A. PC fault. The PC could not be reset. Origin : MCB fault at power converter. | AB/CO |  | Richard called the AB/PO first line. As no spare available, the MCB of the RMB was used because no cold test foreseen on B cluster. |
| Friday _25_N | TBB2 | IPA @ cold : difficult to adjust the offset of potaim cards D1_L-U, D2_U-L and D1-D2. | MTM |  |  |
| Saturday _26_M | TBA1 | 9th ramp to ultimate : quench at ultimate after the file extension was changed in ‘u’. | AB-CO | hf | Two files “u” were created and some data in hf files missing. |
| Saturday _26_M | TBE2 | De-ramp of the current at 5527 A during SSL at 4.4 K . interlocks are : SPA and SWOPRQ | ? |  |  |
| Saturday _26_A | TBE1 | Provoked quench at 1.5 kA (all HF). Quench in ext.bus bar. | MTM |  | Richard informed. He asked to repeat the quench triggering cards 1,2 and 3, 4. Symmetric quench? |
| Sunday _27_M | TBB2 | Capacitor discharge indicator is not glowing even when the capacitor is fully charged. | MTM |  | We do not have to rely on this information. Values measured at the multi-meter are the reference. |
| Sunday _27_A | TBB2 | MBAR2523: HV CDAP not OK, but the magnet was warming up. | MTM/ACR | HV | Richard advised to cool down to 4.4K and repeat the test. Second test indicated OK but file not containing all the data. Problem in the HV application ? The min. energy and warm-up have to be performed without other instruction from the test coordinator only if the HV-CDAP is OK. |
| Monday_28_M Monday_28_A | TBD2 | Lost of Cryo-OK Cryo problem : Cold compressor not working | ACR |  | J.P.Lamboy informed. |