Vc/19-Jul-04

### Week30 starting 19 July  04

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday 12 july  16.13hrs

| TBA1 | MBAR1136 | PT 6.3 Training 3 | Shafts+MM |
| --- | --- | --- | --- |
| TBA2 | MBAR1137 | Prep 5 Pumping + Cool Down |  |
| TBB1 | MBAL3176 | Prep 5 Pumping + Cool Down |  |
| TBB2 | MBAR2036 | Prep 5 Pumping + Cool Down | Shafts+MM |
| TBC1 | MBAR3026 | Finished Ready to go Final scan |  |
| TBC2 | SSS024 | MM 4 Stretched Wire |  |
| TBD1 | SSS023 | Prep 0 Scan Bench |  |
| TBD2 | SSS017 | PT 6.5 Training 5 |  |
| TBE1 | MBAL3179 | PT 6.1 Training 1 |  |
| TBE2 | MBAL3149 | Prep 5 Pumping + Cool Down |  |
| TBF1 | MBBL1129 | Prep 1 Lyre Test | Q-Antennae? |
| TBF2 | MBBL3128 | PT 4 Slow power abort check |  |

Coordinator this week: Nikolay  ; MTM support on call :  (16-3602).

11 Magnets( 11 dipoles+ 0 SSS)   Disconnected after tests last week    ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

Priorities & Board stickup

Tempr reading for SMTMS do not work since many days

  - SSS work :

  - SSW …faults, diagnostics, HV & orbit Correctors 55A supply….(      Shyam + ..)

Work by Alessandro/Federico/Eric for PLCCU + other (many) updates  [ cluster C on Tues aft ?]

Bench status  and tests for now( Monday 12/7at 16.00 hrs)  & next few days :  see Video Display

CHECK SHAFTS USAGE ISSUES

Magnet TB Arrival Out Shafts WP04 HV CDW1 Res. CDW1 HV CDBP HV CDBP Res. CDPT CDAP HV CDW2 Res. WP07 HV MBAR1136 TBA1 14/07/2004 Tests Sequence MBAR1137 TBA2 16/07/2004 Tests Sequence MBAL3176 TBB1 18/07/2004 Tests Sequence MBAR2036 TBB2 17/07/2004 Tests Sequence MBAR3026 TBC1 12/07/2004 Tests Sequence SSS024 TBC2 07/07/2004 Tests Sequence SSS023 TBD1 19/07/2004 Tests Sequence SSS017 TBD2 02/07/2004 Tests Sequence MBAL3179 TBE1 15/07/2004 Tests Sequence MBAL3149 TBE2 17/07/2004 Tests Sequence MBBL1129 TBF1 19/07/2004 Tests Sequence MBBL3128 TBF2 13/07/2004 Tests Sequence Actual Time 19 July 2004 - 16:28 Copyright CERN

# Results by Week for PREVIOUS week

## Tests

| Magnets under Test from 11-07-2004 to 17-07-2004 |
| --- |
| - | Bench | Magnet | Shafts | Start Date . | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total |  |  |
| 1 | TBA1 | MBBR2048 |  | 03-07-2004 | 35 h. | 103 h. | 49 h. | 70 h. | 8 h. | 14-07-2004 | 264 hours |  |  |
| 2 | TBA2 | MBAR1135 |  | 11-07-2004 | 30 h. | 35 h. | 17 h. | 18 h. | 8 h. | 15-07-2004 | 96 hours |  |  |
| 3 | TBA2 | MBBL1123 |  | 09-07-2004 | 7 h. | 62 h. | 47 h. | 11 h. | 29 h. | 11-07-2004 | 48 hours |  |  |
| 4 | TBB1 | MBAL3183 |  | 07-07-2004 | 44 h. | 45 h. | 17 h. | 17 h. | 5 h. | 13-07-2004 | 144 hours |  |  |
| 5 | TBB2 | MBAL3154 |  | 10-07-2004 | 45 h. | 29 h. | 51 h. | 31 h. | 16 h. | 17-07-2004 | 168 hours |  |  |
| 6 | TBC1 | MBBL1118 |  | 07-07-2004 | 58 h. | 10 h. | 30 h. | 16 h. | 5 h. | 12-07-2004 | 120 hours |  |  |
| 7 | TBE1 | MBBL3136 |  | 08-07-2004 | 186 h. | 71 h. | 55 h. | 13 h. | 6 h. | 15-07-2004 | 168 hours |  |  |
| 8 | TBE2 | MBBL1131 |  | 01-07-2004 | 46 h. | 92 h. | 92 h. | 17 h. | 19 h. | 12-07-2004 | 264 hours |  |  |
| 9 | TBE2 | MBBL3133 |  | 12-07-2004 | 24 h. | 51 h. | 20 h. | 16 h. | 11 h. | 17-07-2004 | 120 hours |  |  |
| 10 | TBF1 | MBAL3184 |  | 08-07-2004 | 19 h. | 69 h. | 18 h. | 16 h. | 9 h. | 13-07-2004 | 120 hours |  |  |
| 11 | TBF2 | MBBL1115 |  | 05-07-2004 | 26 h. | 66 h. | 75 h. | 16 h. | 6 h. | 13-07-2004 | 192 hours |  |  |

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBR2048 | 12789 A. (D2-L-S06) |  | - | - | - | - | - | - | - | - | - | - | - |
| 2 | MBAR1135 | 12098 A. (D2-L MRB) | 12251 A. (D2-L MRB) |  | - | - | - | - | - | - | - | - | - | - |
| 3 | MBBL1123 | 12189 A. (D2-L) | 12543 A. (D1-U) | 12513 A. (D1-L) |  | - | - | - | - | - | - | - | - | - |
| 4 | MBAL3183 | 11245 A. (d1-u) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 5 | MBAL3154 | 11238 A. (D2-U-S01) | 12630 A. (D2_U_S01) |  | - | - | - | - | - | - | - | - | - | - |
| 6 | MBBL1118 | 11295 A. (D2-L) | 12727 A. (D1-U) |  | - | - | - | - | - | - | - | - | - | - |
| 7 | MBBL3136 | 10961 A. (D2-U) | 11114 A. (D1-L) | 11343 A. (D2-U) | 11631 A. (D2-U) | 11888 A. (No Data) | 12203 A. (D2-U) | 12411 A. (D2_U) | 12707 A. (D2-U) | 12848 A. (D2_L) |  | - | - | - |
| 8 | MBBL1131 | 11951 A. (D2 L) | 11967 A. (D1-L) | 12282 A. (D1_L) | 12401 A. (D1-L) | 12376 A. (D1-L) | 12559 A. (D1-U) | 12709 A. (D1-L) | 12677 A. (D2-L) | 12806 A. (D1-L) |  | - | - | - |
| 9 | MBBL3133 | 11997 A. (D2-U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 10 | MBAL3184 | 12222 A. (D1-U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 11 | MBBL1115 | 11470 A. (D2-U) | 11732 A. (D1-L) | 11962 A. (D1-U) | 12192 A. (D1-L) | 12250 A. (D2-L) | 12411 A. (D2-L) | 12459 A. (D1-U) | 12577.2 A. (D1-U) | 12586 A. (D2_ U) |  | - | - | - |

General Problems last week :

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_13_N | TBD | Water flow interlock: 4 clusters in operation, too much!.. |  |  | Switch off the water for another cluster, reset the PC, then switch on the water again. |
| Tuesday_13_N | TBC2 | Roll angle measurements did not give good results on SSS24. Wire not centred in stage B. | MTM |  | Re-centring performed:then results OK. |
| Tuesday_13_N | TBD1 | SSS20: Axis measurements at cold : No signals. Repeated twice : no results | MTM |  | Return cable broken: repaired by Fabienne. |
| Tuesday_13_N | TBC2 | SSS24: SSW computer hung while transferring result files from data directory to magnet directory. | MTM |  | Worked again at the end of the shift: origin of the slow down? |
| Tuesday_13_N | TBB2 | Diode test on 3154 magnet current falls down below 3 kA : diode conducting during ramp at 50 A/s: happened twice. | MTM |  | Lower diode threshold than 5.5 V (spec between 5V-7V). Continue the test with ramp at 20 A/s. |
| Tuesday_13_N | TBA2 | IAO at cold, AC comp of potaim cards generator did not provide sine signals injection lead is found damaged Errors in D1L-U, D2 U-L, D1-D2. | MTM |  | Generator was removed by Federico and replaced by anew one |
| Wednesday_14_A | TBC2 | Motor is not working during sextupole measurement at cold. | MTM, |  | Switch ON/OFF the motor controller . Then it worked but the initial position has been lost. Axis measurements was launched but a confi.g file was missing. |
| Wednesday_14_N | TBB2 | Slow PA during LHC cycle. No cryo fault. Cryo lost during the second LHC cycle. | MTM ACR |  | Quench at 6.5 kA to prepare a new LHC cycle J.Milles advised to continue the second time: Comments by Stéphane : without a quench, results from the LHC cycle is not correct: |
| Fri16 July -N | TBD1 | SSS020: HV problem : corrector magnet selector switch problem | MTM | HV Insul | Solved by Eq Support |
| Fri 16 July _N | TBC2 | SSS024 Power supply CD “temp” interlock |  |  | Solved by Equip support |
|  | TBB2 | CDW2 resistances for Magnet 3154 NOK for MCS, MCD, MCS+MCO |  |  |  |
| Sat 17 July _N | TBA1 | TEMA blocked with “error recv info from UDP socket” | AB-CO | TEMA | Eq support needed CO support |
| Sun 18 July _M | D2 | SSS017: Doing PT4, LF-HF got armed but current NOT ramping up; OP found solution by putting PC on Standby manually & test relaunched | MTM or AB-CO |  |  |
| Sun 18 July _M | D2 | Alerts on Quench heaters ? John informed Doing PT6 , power ramped to 9845 A but dropped (SPA) to zero due to Cryo spike |  |  |  |
|  | A1 | Tcoil bad reading / Coeff recalibrated ? |  |  |  |