Vc/9-Aug-04

### Week33 starting 9Aug  04

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday 09 aug04  ~17.45 hrs

| TBA1 | MBAR1139 | PT 10 Min. NRJ 11850 A | 1.9 |
| --- | --- | --- | --- |
| TBA2 | MBBL2089 | Prep 5 Pumping + Cool Down | 1.94 |
| TBB1 | SSS041 | PT 6.2 Training 2 | 3.3 HV |
| TBB2 | MBBL1121 | Prep 5 Pumping + Cool Down | 302.5 |
| TBC1 | MBBL3194 | Prep 5 Pumping + Cool Down | 89 |
| TBC2 | MBBL2108 | PT 12 Warm Up | 300.8 |
| TBD1 | SSS007 | Stretched Wire Warm SW13 | 301.4 |
| TBD2 | SSS036 | ICS 5 Traveller Last | 301.9 |
| TBE1 | MBBL3193 | ICS 5 Traveller Last | 302.5 |
| TBE2 | MBAL3113 | ICS 3 Traveller after warm test | 298 |
| TBF1 | MBAL2047 | PT 3 QH Measure | 1.9 |
| TBF2 | MBBL1120 | PT 6.4 Training 4 | 1.9 |

8/9/2004 5:40:25 PM    Temperature data are taken from TBSM18 site

Coordinator this week: Gijs; MTM support on call :  (16-3602).

12 Magnets( 11 dipoles+ 1 SSS)   Disconnected after tests last week    ( Sat Midnight to Sat midnight)  ;

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.(Jerome ?)

  - SSS work :

  - RESISTANCE M’ments at warm at the beginning & end

1. Orbital corrector dipoles

2. Cryo heater

3. Temp Sensor TT 821

Item 2 & 3 can be included without any trouble

Howver for item 1 the availability of current source 10 mA, 10 VDC is Required ( Tcoil supply ?)

Bench status  and tests for now( Monday 2 Aug/7at 17.40 hrs)  & next few days :  see Video Display

| Magnets under Test on 09-08 5:46:25 PM |
| --- |
| - | Bench | Magnet | Shafts | Start | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | Last Test Desc | Time | I | Temp. | Total |
| 1 | TBA1 | MBAR1139 |  | 02-08 | 50 h. | 106 h. | 25 h. | - | - | PT 10 Min. NRJ 11850 A | 09-08 (16 : 56) | 11850 A. | 1.9 K | 181 hours |
| 2 | TBA2 | MBBL2089 |  | 06-08 | 27 h. | 44 h. | - | - | - | Prep 5 Pumping + Cool Down | 07-08 (21 : 30) | A. | Cooling | 71 hours |
| 3 | TBB1 | SSS041 |  | 02-08 | 132 h. | 32 h. | 11 h. | - | - | PT 6.2 Training 2 | 09-08 (16 : 56) | 12610 A. | 1.9 K | 175 hours |
| 4 | TBB2 | MBBL1121 |  | 05-08 | 48 h. | 44 h. | - | - | - | Prep 5 Pumping + Cool Down | 07-08 (21 : 31) | A. | Cooling | 92 hours |
| 5 | TBC1 | MBBL3194 |  | 07-08 | 28 h. | 15 h. | - | - | - | Prep 5 Pumping + Cool Down | 09-08 (02 : 42) | 2 A. | Cooling | 43 hours |
| 6 | TBC2 | MBBL2108 |  | 04-08 | 44 h. | 46 h. | 20 h. | 15 h. | - | PT 12 Warm Up | 09-08 (02 : 30) | 2 A. | Warming | 125 hours |
| 7 | TBD1 | SSS007 |  | 05-08 | 116 h. | - | - | - | - | Stretched Wire Warm SW13 | 08-08 (01 : 52) | 0 A. | 300 K | 116 hours |
| 8 | TBD2 | SSS036 |  | 28-07 | 156 h. | 4 h. | 63 h. | 22 h. | 2 h. | ICS 5 Traveller Last | 08-08 (01 : 38) | 0 A. | 300 K | 287 hours |
| 9 | TBE1 | MBBL3193 |  | 04-08 | 58 h. | 16 h. | 73 h. | - | - | ICS 5 Traveller Last | 09-08 (10 : 30) | 0 A. | 300 K | 147 hours |
| 10 | TBE2 | MBAL3113 |  | 07-08 | 46 h. | - | - | - | - | ICS 3 Traveller after warm test | 08-08 (01 : 30) | 0 A. | 300 K | 46 hours |
| 11 | TBF1 | MBAL2047 |  | 04-08 | 40 h. | 92 h. | 10 h. | - | - | PT 3 QH Measure | 09-08 (09 : 07) | 0 A. | 1.9 K | 142 hours |
| 12 | TBF2 | MBBL1120 |  | 05-08 | 49 h. | 7 h. | 37 h. | - | - | PT 6.4 Training 4 | 09-08 (14 : 22) | 12437 A. | 1.9 K | 93 hours |

Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAR1139 | 11634 Amp. (D1-L) | 12507 Amp. (D2-L) |  | - | - | - | - | - | - | - | - | - | - |
| 2 | MBBL2089 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 3 | SSS041 | 12281 Amp. (QI-12) | 12610 Amp. (No Data) |  | - | - | - | - | - | - | - | - | - | - |
| 4 | MBBL1121 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 5 | MBBL3194 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 6 | MBBL2108 | 10420 Amp. (D1-L) | 12404 Amp. (D2-L) |  | - | - | - | - | - | - | - | - | - | - |
| 7 | SSS007 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 8 | SSS036 | 12482 Amp. (Q_I_12) | 12491 Amp. (Q_I_12) | 12850 Amp. (Q-E-34) |  | - | - | - | - | - | - | - | - | - |
| 9 | MBBL3193 | 9536 Amp. (D2-U) | 10352 Amp. (D1_U) | 11515 Amp. (D2_U) | 11490 Amp. (D2_U) | 11809 Amp. (D2-U) | 12038 Amp. (D2_L) | 12143 Amp. (D2-U) | 12401 Amp. (D2-U) | 12659 Amp. (D2_U) |  | - | - | - |
| 10 | MBAL3113 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 11 | MBAL2047 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 12 | MBBL1120 | 12012 Amp. (D2-U-S13) | 11901 Amp. (D1_L_S13) | 12222 Amp. (D2-L-S13) | 12437 Amp. (D2-U-S13) |  | - | - | - | - | - | - | - | - |

# Results by Week for PREVIOUS week

| Magnets under Test from 01-08-2004 to 07-08-2004 |
| --- |
| - | Bench | Magnet | Shafts | Start Date . | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total |  |  |
| 1 | TBA1 | MBAL3177 |  | 25-07-2004 | 37 h. | 50 h. | 71 h. | 15 h. | 15 h. | 02-08-2004 | 192 hours |  |  |
| 2 | TBA2 | MBBL3191 |  | 31-07-2004 | 34 h. | 55 h. | 37 h. | 17 h. | 8 h. | 06-08-2004 | 144 hours |  |  |
| 3 | TBB1 | MBBL3190 |  | 26-07-2004 | 33 h. | 48 h. | 36 h. | 17 h. | 6 h. | 02-08-2004 | 168 hours |  |  |
| 4 | TBB2 | MBBL2090 |  | 29-07-2004 | 30 h. | 52 h. | 66 h. | 17 h. | 8 h. | 05-08-2004 | 168 hours |  |  |
| 5 | TBC1 | MBAL3110 |  | 03-08-2004 | 30 h. | 41 h. | 26 h. | 10 h. | 10 h. | 07-08-2004 | 96 hours |  |  |
| 6 | TBC1 | MBAL3172 |  | 26-07-2004 | 48 h. | 71 h. | 34 h. | 20 h. | 12 h. | 03-08-2004 | 192 hours |  |  |
| 7 | TBD1 | SSS023 |  | 19-07-2004 | 98 h. | 95 h. | 119 h. | 17 h. | 1 h. | 03-08-2004 | 360 hours |  |  |
| 8 | TBE1 | MBBL3137 |  | 28-07-2004 | 58 h. | 29 h. | 41 h. | 16 h. | 12 h. | 03-08-2004 | 144 hours |  |  |
| 9 | TBE2 | MBAL3170 |  | 02-08-2004 | 9 h. | 51 h. | 23 h. | 18 h. | 2 h. | 07-08-2004 | 120 hours |  |  |
| 10 | TBE2 | MBBL2035 |  | 23-07-2004 | 37 h. | 114 h. | 48 h. | 23 h. | 7 h. | 02-08-2004 | 240 hours |  |  |
| 11 | TBF1 | MBAR1140 |  | 27-07-2004 | 24 h. | 47 h. | 78 h. | 13 h. | 6 h. | 03-08-2004 | 168 hours |  |  |
| 12 | TBF2 | MBBL1119 |  | 29-07-2004 | 33 h. | 67 h. | 44 h. | 13 h. | 8 h. | 05-08-2004 | 168 hours |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAL3177 | 11627 A. (D2-U) | 12226 A. (D2-U) |  | - | - | - | - | - | - | - | - | - | - |
| 2 | MBBL3191 | 11167 A. (D1-U) | 12329 A. (D2-U) |  | - | - | - | - | - | - | - | - | - | - |
| 3 | MBBL3190 | 10479 A. (D2-L) | 12519 A. (D2_L) |  | - | - | - | - | - | - | - | - | - | - |
| 4 | MBBL2090 | 10787 A. (D1-U-S01) | 11153 A. (D2 L S01) | 12412 A. (D1_U S03) | 12816 A. (D2_L_S12) | 12850 A. | - | - | - | - | - | - | - | - |
| 5 | MBAL3110 | 12109 A. (D2_U) | 12406 A. (D1_U) |  | - | - | - | - | - | - | - | - | - | - |
| 6 | MBAL3172 | 11656 A. (D2_U) | 12851 A. | - | - | - | - | - | - | - | - | - | - | - |
| 7 | SSS023 | 12287 A. (QI-34) | 12616 A. (QI-12) | 12686 A. (QE_12) | 12845 A. (QE_12) | 12855 A. (QE_34) |  | - | - | - | - | - | - | - |
| 8 | MBBL3137 | 11089 A. (D2-L) | 12296 A. (D2_U) |  | - | - | - | - | - | - | - | - | - | - |
| 9 | MBAL3170 | 12497 A. (D2-U) | 12642 A. (D2_U) |  | - | - | - | - | - | - | - | - | - | - |
| 10 | MBBL2035 | 11580 A. (D2-U) | 11859 A. (D1_U) | 12616 A. (D1-L) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 11 | MBAR1140 | 11202 A. (D2-L) | 11645 A. (D2-L) | 12142 A. (D2-L) | 12383 A. (D1-L) | 12513 A. (D1-U) | 12597 A. (D2 L) | 12662 A. (D1-U) | 12700 A. (D2-U) | 12332 A. (D2_U) |  | - | - | - |
| 12 | MBBL1119 | 11940 A. (D2-U-S13) | 12119 A. (D1-L-S13) | 12231 A. (D2-U-S13) | 12498 A. (D2_U-S01) | 12691 A. (D2-U-S13) |  | - | - | - | - | - | - | - |

General Problems last week :

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _03_A | TBB2 | VME error during LHC cycle. Measurements stuck in one current plateau. Same problem on the load line. | AB-CO | MMP | LHC cycle aborted: Load line launched. Giorgio informed. Origin? Run recovery performed. |
| Wednesday_04_A Friday_06_N | TBD2 | HV Test failed on EF circuit (did not pass 800 V) for SSS36 NC remained after the power test. | MTM |  | TC informed. NC opened. OK to continue but power test on corrector EF will be performed after the MQ training after another HV test on EF if insulation is not degraded. |
| Wednesday _04_A | TBA2 | Training quench one on 3091: 2 DAQ field created and red alarm on QHeaters. “Heaters should fire after the DAQ trigger”. | MTM | AQA | Check by Giorgio and Stephane using LF DAQ. Acquisition problem. Reset the crate and OK to continue |
| Wednesday_04_N | TBE2 | Rotary switch did not work. Then TEMA blocked on E1 (communication problem between PLCSE and application. | MTM |  | Giorgio informed. He solved the problem |
| Thursday_05_M | TBE2 | HV test failed : TT821 vs ground. Connector inside the IFS box damaged during preparation | ICS |  | Giorgio and Kevin repaired the pin which was bent. |
| Thursday _05_M | TBD2 | SPA happened at 5000 A .Giorgio was there: he checked the PLC before resetting: no apparent fault | AB/CO? |  | Origin not known: Communication problem between PLCSE and PC |
| Friday_06_N | TBD2 | TCLaux error on the PLC. T problem on the f auxiliary current leads | ACR |  | Giorgio called but it is a cryo problem. Solved by using a potentiometer to keep the valves opened. |
| Friday_06_N | TBD1 | IAP @warm: Vtap signals not visible in IAP but visible in lfoff except IAP_QE_P12. | MTM | IAP | Giorgio informed: He told Nair that the reason is that the SSS is of B type and the corresponding name is IAP_QE_P34. |
| Saturday _07_A | TBD1 | Axis measurements on SSS07 at warm: physical offsets in stage A and B were above the allowable limits. Repeat again: the wire snapped. Repeat third time: Results not displayed in the screen? | MTM |  | Origin not known: Communication problem between PLCSE and PC |
| Saturday _07_N | TBE1 | Training quench Nº9: three files a09, a10, a11. The first 2 are erroneous files. a11 is the real quench. Data could not be found. | AB/CO |  | Giorgio comments: the data are present but on the local machine only. Error when copying to storage disk. |
| Sunday _08_A | TBD2 | HV test at warm: Connector C11 broken. | MTM |  | Giorgio advised to continue the other steps. Then he came and unbent the pin: then OK. |
| Sunday _08_A | TBF2 | 1120 magnet: Provoked quench at 1.5 kA. No trigger generated because no quench. Quench heaters 3&4 in protection did not discharge. Test repeated with 1&2 in protection. Quench appeared. | MTM |  | Giorgio informed: he compared with the results of the first run: the same! As protection is done by HF heaters, the test continued. Stephane comments: a NC has to be opened and the magnet in stand by in any case. |
| Monday _09_M | TBF2 | SPA occurred during LHC cycle. | ? |  | Origin unknown. |