Vc/1-Feb-05

### Week 5 starting 31 Jan 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday 31 Jan 05 @   1/31/2005 3:30:40 PM

|  | A1 | MBAL1173 | Prep 5 Pumping + Cool Down |  |
| --- | --- | --- | --- | --- |
|  | A2 | MBAR3278 | Prep 4 TRU Connection |  |
|  | B1 | SSS056 | PT13.3: WP17 HV Test |  |
|  | B2 | MBBR3271 | PT 12 Warm Up |  |
|  | C1 | MBAL1174 | Current Lead upgrade |  |
|  | C2 | MBAR3253 | Current Lead Upgrade |  |
|  | D1 | SSS036 | PT13.3: WP17 HV Test |  |
|  | D2 | SSS058 | Prep 5 Pumping + Cool Down |  |
|  | E1 | MBAR3280 | Prep 5 Pumping + Cool Down |  |
|  | E2 | MBAL3108 | Prep 5 Pumping + Cool Down |  |
|  | F1 | MBAR3283 | Prep 0 Scan Bench |  |
|  | F2 | MBAL2123 | ICS 5 Traveller Last |  |

1/31/2005 3:30:40 PM    Temperature data are taken from TBSM18 site

Coordinator this week: Vittorio; MTM support on call :   16-3602 .

ONLY 4  Magnets Tested  !

4  dipoles( 0 with MM & long MM)+ 0  SSS(0  with MM) +  0  SSSS     Disconnected after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Situation on Repair :

      - Silver Shoes ( problems to finish  in shutdown?)

      - F Cluster & Quench in busbars or … Instrumentation ?

General Points

Problems commutation  Dipole to SSS on B cluster (salem)

SSW Measurements

Cryo Capacity & Followup  + priorities

AOB

## Tests

| Magnets under Test from 23-01-2005 to 29-01-2005 |
| --- |
| - | Bench | Magnet | Shafts | Start Date . | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total |  |  |
| 1 | TBA1 | MBAL2127 |  | 19-01-2005 | 139 h. | 58 h. | 18 h. | 16 h. | 7 h. | 29-01-2005 | 240 hours |  |  |
| 2 | TBA2 | MBBL1108 |  | 19-01-2005 | 123 h. | 45 h. | 20 h. | 20 h. | 5 h. | 28-01-2005 | 216 hours |  |  |
| 3 | TBB2 | MBAL1175 |  | 19-01-2005 | 139 h. | 33 h. | 16 h. | 19 h. | 7 h. | 28-01-2005 | 216 hours |  |  |
| 4 | TBE2 | MBAL1074 |  | 19-01-2005 | 139 h. | 25 h. | 39 h. | 20 h. | 8 h. | 29-01-2005 | 240 hours |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAL2127 | 11374.5 A. (D2_U) | 12074.1 A. (D2_L) |  | - | - | - | - | - | - | - | - | - | - |
| 2 | MBBL1108 | 11940 A. (D2_U) | 12471 A. (D1_U) |  | - | - | - | - | - | - | - | - | - | - |
| 3 | MBAL1175 | 12490 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 4 | MBAL1074 | 11571 A. (D2-U) | 12313 A. (D2-U) |  | - | - | - | - | - | - | - | - | - | - |
| From 23-01-2005 to 29-01-2005 |
| - | Bench | Magnet Name | shafts | Start Date . | Connect Time | Wait +Cooling | Cold Tests | Wait +Warming | Disconnect Time | End Date . | Total | Q |
| 1 | TBE2 | MBAL1074 | False | 19-01-2005 | 139 h. | 25 h. | 39 h. | 20 h. | 8 h. | 29-01-2005 | 10 days | 2 |
| 2 | TBB2 | MBAL1175 | False | 19-01-2005 | 139 h. | 33 h. | 16 h. | 19 h. | 7 h. | 28-01-2005 | 9 days | 1 |
| 3 | TBA2 | MBBL1108 | False | 19-01-2005 | 123 h. | 45 h. | 20 h. | 20 h. | 5 h. | 28-01-2005 | 9 days | 2 |
| 4 | TBA1 | MBAL2127 | False | 19-01-2005 | 139 h. | 58 h. | 18 h. | 16 h. | 7 h. | 29-01-2005 | 10 days | 2 |

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Monday _24_N | TBD1 | C11 connector damaged (pin broken) | MTM |  | Used as is carefully: Richard informed A.Siemko for an action with the MEL |
| Tuesday _25_M | TBF2 | Trigger of one potaim card Problem in one lf crate | AB/CO | Hf, lf | Federico informed. DAQ hf and lf blocked because one hidden application was running even if it closed.The process was killed. |
| Friday 28_N | TBF | hf DAQ problem | AB/CO | Hf | Federico called Adrian Several problems were found: pins in the MXI connectors bended the CPU card did not boot normally. Card changed. Adrian managed to make the hf DAQ functioning. |
| Friday 28_N | TBF | “ Arm error 2” when arming DAQ hf | MTM |  | Federico changed the ADC card in slot 2 |
| Tuesday _25_A | TBD1 | Anti-cryostat T problem | MTM |  | Solved after re charging the good coefficients from MTF |
| Tuesday _25_N |  | No pumping from 4.2 K to 1.9 K possible: oil leak in one group. Problem in one water controller in the second group | ACR |  | Solved Wednesday mid day . At 1.9 K Wednesday 15h. 16h lost. |
| Tuesday _25_N | TBA2 | Offset too high in potaim cards | MTM |  | Giorgio adjusted the offsets Wednesday morning |
| Wednesday_26_A | TBB2 | LF quench heater provoked quench. No signal seen by hf DAQ | MTM | lf | One cable used for the QH tests was disconnected because a SSSS was tested before. Solved by Giorgio. |
| Wednesday_26_A | TBB2 | LHF quench heater provoked quench. Vmin too high for one heater(12 V): red alarm | MTM |  | Stephane decided to repeat the quench. Same results. Check of the QH discharge and MIITS: OK OK to continue. |
| Wednesday_26_N |  | Leak in compressors: problem in cooling down the SSS in TBD1 | ACR |  |  |
| Wednesday_26_N | TBE | Power supply problem when starting the SPA at 1 kA: RME could not be turned on. | ? |  | Federico called. He reset the communication card of the RME then OK. |
| Wednesday_26_N | TBE2 | 1.5 kA QH test : red alarm DVT too high. Non consistency of the signals in the D1 pole | MTM |  | Stephane and Federico informed. Problem in one gain of the Potaim card. Test repeated: OK. |
| Thursday_27_M | TBE2 | Diode test: MIITS too high (26 instead of 23): red alarm | MTM |  | Stephane informed: check the quench parameters with Pierre (heater delay, QH discharge). OK to continue. |
| Thursday_27_A | TBD1 | SSS36: HV test MQT vs Ground failed at cold at 600 V. | MTM |  | Richard checked the rack, made investigations. NC opened. |
| Thursday_27_A | TBB1 | SSS056 : Resistance measurements slightly out of range | MTM |  | Stephane informed. He advised to open a NC and to proceed in the connection. |
| Thursday_27_A | TBB1 | SSS056: scanning problem Origin : Different naming in MTF OP team did not have the rights | TS/CE | HWREC | Maryline asks T.Ladzinski: to correct the name give the right access. |
| Thursday_27_N | TBF2 | 2123: Quench at 6.2 kA in the bus bar. | MTM |  | Stephane informed. Checks of electronic cards. No results. Quench repeated 2 times: same results. Warm up of the magnet and opening with MTM experts. |
| Friday 28_A | TBF1 | IAP at cold on 2128 magnet: DAQ is not working | MTM | IAP | Federico called. He restarted IAP application: solved. |
| Friday 28_A | TBA2 | ICS found the circuits opened for the short anti-cryostats. | MTM |  | Federico called. He found the good values of R. He explained to ICS team how to use the test box. |
| Friday 28_A | TBE1 | Warm up quench : DVT too high (156V) Red alarm. | MTM |  | TC informed. Quench cleaned. Warm up confirmed. |
| Friday 28_A | TBD1 | During extra HV test Richard found that one current lead was too hot (~100 C) | ACR |  | Cryo team informed: One sensor was out of used. |
| Saturday_29_A | TBE2 | IAP at warm : Tau jump error `6.1 % | MTM |  | TC informed. Test repeated. Then OK. |
| Saturday_29_N | TBB1 | TCL aux (current lead auxiliary) are not OK | ACR |  | Cryo team informed |
| Sunday_30_M | TBB1 | SSS056 : One training quench performed (12600 A). Then faulty quenches (DAQ triggered). No data in the a02 file. | ACR |  | What happened? |
| Sunday_30_M | TBB1 | MXI error while launching PT8 on SSS056 | ACR |  | Federico advised to reset the DAQ crates :solved. |
| Sunday_30_M | TBB1 | HWREC error (equipment missing) when launching the powering of corrector CD, EF. Ramp of current impossible : SPA | MTM AB/PO |  | Federico called. He decided to switch off HWREC Federico found that L is not what it is expected (18 mH). He tried some values and found that L=3 mH is matching. Ramp done manually. |
| Sunday_30_A | TBA2 | During cool down, ACR tema informed that Tanticryosta less than 0 C | MTM |  | Stephane and Federico informed. Federico found that a cable comimg from the PLC ART was disconnected The heating restart. Stephane advised to put the shafts when Tanticryostat stable and above 20 C.. |
| Sunday_30_N | TBB1 | Tcryo~30 K but magnet was still at 80 K (liquefactor problem , BP too high). Traction value too high 18 Kg. | ACR |  | TC informed. Shaft insertion stopped by Stephane. Shaft inserted Monday by MTM mechanics |
| Monday_31_M | TBE | Ripple observed during the current discharge | MTM |  | Salem informed. |

- Cluster A: MXI-VME module base address switch position modified

- Cluster B: VD50 modules swapped

- Cluster C: MXI cable badly connected (pins touch)

- Cluster D: MXI-VME module base address switch position modified

- Cluster E: Ok (until now)

- Cluster F: MXI cable badly connected (pins touch)