Vc/21-Feb-05

### Week 08 starting 21 Feb 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday 21 feb 05 @ 2/21/2005 3:21:51 PM

| 8 | A1 | MBBR3291 | ICS 4 Final connection | 291.87 |
| --- | --- | --- | --- | --- |
| 4 | A2 | MBBR3269 | Prep 5 Pumping + Cool Down | 1.93 HV |
| 2 | B1 | SSS110 | PTE 3 Correctors Powering | 1.90 |
| 5 | B2 | MBAR3288 | PT 13.1 IAP @ Warm | 299.11 |
| 10 | C1 |  |  | 291.87 |
| 1 | C2 | MBAL1080 | Prep 5 Pumping + Cool Down | 81.71 |
| 6 | D1 | SSS073 | Prep 2.3 Res. Meas. CDW1 | 91.97 |
| 9 | D2 |  | Current Lead repair | 292.94 |
| 3 | E1 | MBAL3171 | PT 12 Warm Up | 283.48 |
| 11 | E2 |  |  | 295.18 |
| 12 | F1 |  |  | 293.57 |
| 7 | F2 | MBBL3153 | ICS 4 Final connection | 291.78 |

2/21/2005 3:22:54 PM    Temperature data are taken from TBSM18 site

Coordinator this week: Nikolay; MTM support on call :   16-3602 .

14  Magnets Tested   ( not counting sss108)

10  dipoles( 2 with MM)+ 4  SSS(0  with MM) +  0  SSSS

Disconnected after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Situation on Repair :

      - Silver Shoes bench F etc

General Points

e-Traveller & Paper Traveller

Power Converters  Slow PA & filter by AB-PO

AOB

| Magnets under Test from 13-02-2005 to 19-02-2005 |
| --- |
| - | Bench | Magnet | Shafts | Start Date . | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total |  |  |
| 1 | TBA1 | MBAL2123 |  | 04-02-2005 | 50 h. | 30 h. | 142 h. | 14 h. | 22 h. | 15-02-2005 | 264 hours |  |  |
| 2 | TBA2 | MBAR3278 |  | 13-02-2005 | 26 h. | 35 h. | 47 h. | 16 h. | 9 h. | 18-02-2005 | 120 hours |  |  |
| 3 | TBB1 | SSS034 |  | 10-02-2005 | 36 h. | 19 h. | 30 h. | 7 h. | 29 h. | 16-02-2005 | 144 hours |  |  |
| 4 | TBB1 | SSS108 |  | 17-02-2005 | 23 h. | - | - | - | - | 18-02-2005 | 24 hours |  |  |
| 5 | TBB2 | MBAL3165 |  | 11-02-2005 | 39 h. | 39 h. | 30 h. | 16 h. | 13 h. | 17-02-2005 | 144 hours |  |  |
| 6 | TBC1 | MBAR3286 |  | 08-02-2005 | 36 h. | 29 h. | 56 h. | 13 h. | 13 h. | 14-02-2005 | 144 hours |  |  |
| 7 | TBC2 | MBAL3168 |  | 14-02-2005 | 32 h. | 21 h. | 31 h. | 18 h. | 22 h. | 19-02-2005 | 120 hours |  |  |
| 8 | TBC2 | MBAR3253 |  | 08-02-2005 | 23 h. | 36 h. | 49 h. | 13 h. | 19 h. | 14-02-2005 | 144 hours |  |  |
| 9 | TBD1 | SSS003 |  | 14-02-2005 | 46 h. | 13 h. | 24 h. | 14 h. | 33 h. | 19-02-2005 | 120 hours |  |  |
| 10 | TBD1 | SSS032 |  | 09-02-2005 | 43 h. | 14 h. | 25 h. | 9 h. | 20 h. | 14-02-2005 | 120 hours |  |  |
| 11 | TBD2 | SSS107 |  | 08-02-2005 | 87 h. | 22 h. | 16 h. | 8 h. | 27 h. | 15-02-2005 | 168 hours |  |  |
| 12 | TBE1 | MBAL3141 |  | 12-02-2005 | 25 h. | 38 h. | 33 h. | 13 h. | 13 h. | 17-02-2005 | 120 hours |  |  |
| 13 | TBE2 | MBBL2146 |  | 10-02-2005 | 35 h. | 34 h. | 25 h. | 19 h. | 10 h. | 16-02-2005 | 144 hours |  |  |
| 14 | TBF1 | MBAR3285 |  | 10-02-2005 | 88 h. | 21 h. | 25 h. | 16 h. | 10 h. | 17-02-2005 | 168 hours |  |  |
| 15 | TBF2 | MBBR3270 |  | 10-02-2005 | 32 h. | 27 h. | 32 h. | 21 h. | 6 h. | 15-02-2005 | 120 hours |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAL2123 | 11320.6 A. (D2-U-S13) | 12523.5 A. (D1-L-S10) | 12675.9 A. (D1-U-S01) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 2 | MBAR3278 | 12009 A. (D2-U) | 12255 A. (D2-U-S01) | 12419 A. (D2-U-S09) | 12590.7 A. (D2-U-S05) | 12689 A. (D1 - U) | 12850 A. | 12850 A. | 12850 A. | - | - | - | - | - | - | - |
| 3 | SSS034 | 12281.4 A. (QE_12) | 12256.3 A. (QE-12) | 12664 A. (QE-12) | 12687 A. (QE-34) | 12850 A. | - | - | - | - | - | - | - | - |  |  |
| 4 | SSS108 |  | - | - | - | - | - | - | - | - | - | - | - | - |  |  |
| 5 | MBAL3165 | 10106 A. (D2_L) | 12303.3 A. (D2 - L) |  | - | - | - | - | - | - | - | - | - | - |  |  |
| 6 | MBAR3286 | 11801.2 A. (D2-U) | 12128.9 A. (D2-U) |  | - | - | - | - | - | - | - | - | - | - |  |  |
| 7 | MBAL3168 | 9734 A. (D2_L) | 11471 A. (D2 - U) | 12612 A. (D2-U) | 12574 A. (D2-U) | 12755 A. (D2-U) | 12850 A. | - | - | - | - | - | - | - |  |  |
| 8 | MBAR3253 | 4858 A. (Vcla int) | 6690 A. (D2-L) | 11379.6 A. (D2-L) | 12125 A. (D2_L) | 12950.7 A. (D1-L) |  | - | - | - | - | - | - | - |  |  |
| 9 | SSS003 | 11604.3 A. (Q_I_12) | 12111.5 A. (No Data) |  | - | - | - | - | - | - | - | - | - | - |  |  |
| 10 | SSS032 | 12818.8 A. (QI-34) | 12683 A. (QI-34) | 12850 A. | - | - | - | - | - | - | - | - | - | - |  |  |
| 11 | SSS107 | 12249.3 A. (QI_34) | 12848 A. (QI_34) |  | - | - | - | - | - | - | - | - | - | - |  |  |
| 12 | MBAL3141 | 11514.6 A. (D2-U) | 12745 A. (D2_U) | 12850 A. | 12856 A. | - | - | - | - | - | - | - | - | - | - |  |
| 13 | MBBL2146 | 10351.4 A. (D2_U) | 12230.1 A. (D1-L) | 12727 A. (D1_L) |  | - | - | - | - | - | - | - | - | - |  |  |
| 14 | MBAR3285 | 8391 A. (D2-L) | 12782 A. (D2-L) |  | - | - | - | - | - | - | - | - | - | - |  |  |
| 15 | MBBR3270 | 12526 A. (D1-U-S02) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |  |  |
| From 13-02-2005 to 19-02-2005 |
| - | Bench | Magnet Name | shafts | Start Date . | Connect Time | Wait +Cooling | Cold Tests | Wait +Warming | Disconnect Time | End Date . | Total | Q |
| 1 | TBF2 | MBBR3270 | True | 10-02-2005 | 32 h. | 27 h. | 32 h. | 21 h. | 6 h. | 15-02-2005 | 5 days | 1 |
| 2 | TBF1 | MBAR3285 | False | 10-02-2005 | 88 h. | 21 h. | 25 h. | 16 h. | 10 h. | 17-02-2005 | 7 days | 2 |
| 3 | TBE2 | MBBL2146 | False | 10-02-2005 | 35 h. | 34 h. | 25 h. | 19 h. | 10 h. | 16-02-2005 | 6 days | 3 |
| 4 | TBE1 | MBAL3141 | False | 12-02-2005 | 25 h. | 38 h. | 33 h. | 13 h. | 13 h. | 17-02-2005 | 5 days | 2 |
| 5 | TBD2 | SSS107 | False | 08-02-2005 | 87 h. | 22 h. | 16 h. | 8 h. | 27 h. | 15-02-2005 | 7 days | 2 |
| 6 | TBD1 | SSS003 | False | 14-02-2005 | 46 h. | 13 h. | 24 h. | 14 h. | 33 h. | 19-02-2005 | 5 days | 2 |
| 7 | TBD1 | SSS032 | False | 09-02-2005 | 43 h. | 14 h. | 25 h. | 9 h. | 20 h. | 14-02-2005 | 5 days | 2 |
| 8 | TBC2 | MBAL3168 | False | 14-02-2005 | 32 h. | 21 h. | 31 h. | 18 h. | 22 h. | 19-02-2005 | 5 days | 5 |
| 9 | TBC2 | MBAR3253 | False | 08-02-2005 | 23 h. | 36 h. | 49 h. | 13 h. | 19 h. | 14-02-2005 | 6 days | 5 |
| 10 | TBC1 | MBAR3286 | False | 08-02-2005 | 36 h. | 29 h. | 56 h. | 13 h. | 13 h. | 14-02-2005 | 6 days | 2 |
| 11 | TBB2 | MBAL3165 | False | 11-02-2005 | 39 h. | 39 h. | 30 h. | 16 h. | 13 h. | 17-02-2005 | 6 days | 2 |
| 12 | TBB1 | SSS034 | False | 10-02-2005 | 36 h. | 19 h. | 30 h. | 7 h. | 29 h. | 16-02-2005 | 6 days | 4 |
| 13 | TBB1 | SSS108 HV orbit Cor | False | 17-02-2005 | 23 h. | 88 h. | -87 h. | 1 h. | 8 h. | 18-02-2005 | 1 days | 0 |
| 14 | TBA2 | MBAR3278 | True | 13-02-2005 | 26 h. | 35 h. | 47 h. | 16 h. | 9 h. | 18-02-2005 | 5 days | 5 |
| 15 | TBA1 | MBAL2123 | True | 04-02-2005 | 50 h. | 30 h. | 142 h. | 14 h. | 22 h. | 15-02-2005 | 11 days | 3 |

### Week 14 Feb05 to 20Feb 05

OP statistics for SM18 OP during the period from 14/02/2005 06:00:00 to 21/02/2005 21:00:00

| Lines | in Super Cycle | in Faults | Availabilities (%) |
| --- | --- | --- | --- |
| TBA1 | 126 [h] 44 [min] 43[s] | 13 [h] 02 [min] 16[s] | 89.71 |
| TBA2 | 126 [h] 44 [min] 43[s] | 11 [h] 17 [min] 25[s] | 91.09 |
| TBB1 | 141 [h] 34 [min] 37[s] | 17 [h] 45 [min] 22[s] | 87.46 |
| TBB2 | 126 [h] 44 [min] 43[s] | 38 [min] 01[s] | 99.5 |
| TBC1 | 126 [h] 44 [min] 43[s] | 38 [min] 01[s] | 99.5 |
| TBC2 | 126 [h] 44 [min] 43[s] | 17 [h] 40 [min] 55[s] | 86.05 |
| TBD1 | 126 [h] 44 [min] 43[s] | 27 [h] 45 [min] 09[s] | 78.1 |
| TBD2 | 126 [h] 44 [min] 43[s] | 1 [h] 06 [min] 08[s] | 99.13 |
| TBE1 | 142 [h] 44 [min] 43[s] | 16 [h] 38 [min] 01[s] | 88.35 |
| TBE2 | 126 [h] 44 [min] 43[s] | 5 [h] 08 [min] 01[s] | 95.95 |
| TBF1 | 126 [h] 44 [min] 43[s] | 3 [h] 34 [min] 32[s] | 97.18 |
| TBF2 | 142 [h] 44 [min] 43[s] | 16 [h] 38 [min] 01[s] | 88.35 |
| TOTAL: | 1567 [h] 46 [min] 30[s] | 131 [h] 51 [min] 52[s] | 91.59 |

| Faults Groups | Faults Names | Duration |
| --- | --- | --- |
| Cryogenics | Miscellaneous | 14 [h] 06 [min] 08[s] |
| Electronics | ----- | 12 [h] 46 [min] 49[s] |
| HV insulation | ----- | 44 [h] 24 [min] 15[s] |
| Hardware | Power converters | 4 [h] 51 [min] 28[s] |
| ICS | ----- | 3 [h] 37 [min] 06[s] |
| Miscellaneous | ----- | 21 [h] 40 [min] 53[s] |
| Network | ----- | 3 [h] 13 [min] 02[s] |
| Software | MTF | 21 [min] 15[s] |
|  | Test master | 1 [h] 32 [min] 53[s] |

Top of Form

Bottom of Form

[SM18 OP] 17/02/2005 05:47:56

Still fighting with RMD. AB/PO second line was called. Richard.

[SM18 OP] 17/02/2005 04:00:00

RMD is returning a "PC fault" signal, that we are unable to clear. AB/PO first line was called. Richard.

[SM18 OP] 16/02/2005 23:10:00

Local PLC for water control was found to be in TEMA mode, due to which the valves could not be opened. Contacted Richard and he advised to put in LOCAL mode. Problem solved.

[SM18 OP] 15/02/2005 23:50:36

TBA2: PT1 - IAP @ cold. Dc offset errors in QA compensation (eg. CF_S07_D). Don't know how to adjust. Talked to Richard... he will look into it but told us to proceed futher

[SM18 OP] 15/02/2005 19:52:38

TBD1 &2: Problem with generation of test reports of magnet SSS074 and SSS003 was solved by Marco. It was due to a template file corrupted. Richard.

[SM18 OP] 15/02/2005 19:50:53

TBC1: HV problem on this bench is now understood. The defect comes from the water cooled power cables. Richard.

[SM18 OP] 15/02/2005 18:02:48

ALL: Test reports of HV tests performed during the night were generated and new password implemented in all HV applications. Richard.

[SM18 OP] 15/02/2005 11:27:07

Could not launch test from Tema. PLCCU was set to manual instead of TEMA??? Richard.

[SM18 OP] 15/02/2005 11:23:58

Potaim card replaced by Richard and problem solved.

[SM18 OP] 15/02/2005 10:03:19

Trigger on Potaim card of B2 being attended by Richard.

[SM18 OP] 15/02/2005 00:15:51

HV report cannot be generated due to change of sm18 password-Richard will take necesseray action tuesday morning

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _15_A |  | Liquefactor stopped. | ACR |  | 4h hours of stop |
| Tuesday _15_A | TBD | Problem in generation of HV test reports on SSS003 and SSS74 | AB/CO | HV | Template file was corrupted. Solved by Marco M. |
| Tuesday _15_N | TBA2 | IAP @ cold. Dc offset errors in LQA compensation. Cannot be adjusted. | MTM | IAP | Richard advised to proceed. No compensation needed for the LQA the moment. |
| Tuesday _15_N | TBE2 | MBBL2146 taken away by ROCLA before signing for stripping in paper or electronically. | ICS |  |  |
| Wednesday_16_M | TBA1 | CDW1--HV test report is not loaded into HV_MTF folder. But test was successful. | MTM | HV | File generated but wrong folder. Magnet is in the first run not in the third one. Re-direction of HV and Quench files performed by Richard Friday. |
| Wednesday_16_N |  | Opening/closing of water valves could not be managed from PLCCU in TEMA mode. impossible to launch the power application from TEMA, | AB/CO? |  | This mode does not work. In this bench. Opening water valve with the PLC in local mode. Test launched manually. |
| Wednesday_16_N | TBD1 | PC fault in RMD when launching the SPA test at 1000 A. | MTM |  | Richard called the AB/PO first line. Second line. Key of the polarity switch not in right position. |
| Thursday_17_M | TBD1 | Not able to power CD correct in SSS003 | MTM |  | Giorgio suspected a problem in the clamp there is NC opened for Res.Measurements (CQW0) also. The V Clamp D signal was by-passed to solve the problem temporarily. |
| Friday_18_M | TBB1 | SSS108: NC detected during HV test (Orbit dipole/ground) Resistance values for Orbit dipole outside the specs. |  |  | Investigation by A.Siemko and Richard : caused by the coating of voltage taps wires being damaged. Connection for correctors not OK: Richard advised to connect the P16 connector. |
| Wednesday_18_M | TBA1 | MM1:Over-range in sector 22-23 for cmp. | MTM |  | Discussed with Maryline and Stephane. Go on with LHC cycle. If Over-range persists, Richard will reduce the gain in amplifiers |
| Wednesday_18_N | TBA1 | PTE (diode test): voltage signals have spikes (warning, D1,D2~11V) and MIITS just over the limit (24.2 MA2s). | MTM |  | Richard called. He advised to continue. |
| Wednesday_18_N | TBC,TBE TBF1 | No sufficient He liquid capacity to cool down all the magnets: During PT8 De-Excitation, SPA occurred probably due to low level of LHe (As per Cryo Operator). | ACR |  | Put the 2 magnets in TBC and TBE in stand by. Test re-launched later. |
| Saturday_19_A | TBF1 | Potaim Card Trigger from D2 U-L and Vtotal . | MTM |  | Richard called. He adjusted offset of card 1 and changed card 7 |
| Saturday_19_A | TBF1 | The two MOBIC scanners are not working . One scanner was re-configured. | MTM |  | Information entered manually for SSS073 and MB1080. |
| Sunday_20_M | TBF1 | Discharged of quench heater during switch on power supply. | MTM |  | A connector was not properly connected. |
| Sunday_20_M | TBF1 | During warm up after the break of the vacuum from cryo, the MRB is without contact of the "interlock support MRB". | ACR |  | Richard has been informed. |
| Sunday_20_M Sunday _20_N | TBE2 TBA | The flashing box high current is not working. The HV rack XMT.MY95 has a connector for passive busbar broken HV on 1186: C11 connector broken! No spare connectors! | MTM |  |  |
| Sunday_20_N | TBF1 | Two quench heaters are not getting discharged while conducting PT5.1 on SSS110. AQA says quench not dectected.. | MTM | AQA | Test repeated 3 times, the third time at 2000 A. Quench not detected? NC had to be open. |