Vc/8-Mar-05

### Week 10 starting 07 March 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday 07 March 05 @ 3/7/2005 11:30:55 AM

| 4 | A1 | MBAL3144 | Prep 5 Pumping + Cool Down | 100.17 |
| --- | --- | --- | --- | --- |
| 5 | A2 | MBAL2129 | PT 12 Warm Up | 164.13 |
| 2 | B1 | SSS601 | Survey @ Cold SC24 | 1.90 |
| 7 | B2 | MBAL2156 | PT 12 Warm Up | 85.51 |
| 1 | C1 | MBAL2149 | PT 6.3 Training 3 | 5.86 |
| 9 | C2 | MBAL3150 | ICS 4 Final connection | 290.71 |
| 3 | D1 | SSS106 | PT 5.2 1.5 KA Quench HF | 1.91 |
| 12 | D2 | SSS129 | ICS 2 Connect Magnet (ICS) | 291.78 |
| 6 | E1 | MBAR3287 | PT 12 Warm Up | 87.69 |
| 8 | E2 | MBBL2137 | PT 12 Warm Up | 238.02 |
| 10 | F1 | MBBR2525 | ICS 4 Final connection | 290.53 |
| 11 | F2 | MBAL2154 | PT 13.2 Resist. Meas. | 292.49 |

Temperature data are taken from TBSM18 site

Coordinator this week: Marta; MTM support on call :   16-3602 .

12 Magnets Tested

9  dipoles( 2 with MM)+ 3  SSS(1  with MM) +  0  SSSS

Disconnected after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

    - Problem on C2 lastFriday. Problem A2 weekend

- Situation on Repair :

      - Silver Shoes bench F etc

General Points

SSSS

Power Converters  Slow PA etc

Monthly  & overall Report

AOB

| Magnets under Test from 27-02-2005 to 05-03-2005 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total |  |  |
| 1 | TBA1 | MBBL3235 |  |  | 25-02-2005 | 30 h. | 83 h. | 39 h. | 12 h. | 10 h. | 04-03-2005 | 168 hours |  |  |
| 2 | TBA2 | MBAL2129 |  |  | 24-02-2005 | 38 h. | 30 h. | 69 h. | 17 h. | 0 h. | 02-03-2005 | 144 hours |  |  |
| 3 | TBB2 | MBBL2121 |  |  | 23-02-2005 | 58 h. | 22 h. | 46 h. | 16 h. | 8 h. | 02-03-2005 | 168 hours |  |  |
| 4 | TBC1 | MBBL1196 |  |  | 27-02-2005 | 22 h. | 31 h. | 46 h. | 16 h. | 10 h. | 04-03-2005 | 120 hours |  |  |
| 5 | TBC2 | MBAL2151 |  |  | 25-02-2005 | 24 h. | 30 h. | 18 h. | 14 h. | 16 h. | 01-03-2005 | 96 hours |  |  |
| 6 | TBD1 | SSS073 |  |  | 23-02-2005 | 15 h. | 13 h. | 59 h. | 8 h. | 13 h. | 28-02-2005 | 120 hours |  |  |
| 7 | TBD1 | SSS100 |  |  | 28-02-2005 | 41 h. | 17 h. | 31 h. | 7 h. | 15 h. | 05-03-2005 | 120 hours |  |  |
| 8 | TBD2 | SSS080 |  |  | 25-02-2005 | 41 h. | 29 h. | 17 h. | 9 h. | 18 h. | 02-03-2005 | 120 hours |  |  |
| 9 | TBE1 | MBBL2147 |  |  | 27-02-2005 | 37 h. | 26 h. | 35 h. | 17 h. | 8 h. | 04-03-2005 | 120 hours |  |  |
| 10 | TBE2 | MBAL2148 |  |  | 26-02-2005 | 27 h. | 39 h. | 21 h. | 14 h. | 9 h. | 03-03-2005 | 120 hours |  |  |
| 11 | TBF1 | MBBL1199 |  |  | 26-02-2005 | 33 h. | 40 h. | 83 h. | 13 h. | 14 h. | 05-03-2005 | 168 hours |  |  |
| 12 | TBF2 | MBAL2131 |  |  | 25-02-2005 | 29 h. | 37 h. | 46 h. | 13 h. | 11 h. | 03-03-2005 | 144 hours |  |  |

### Quench Information Week 27Feb05

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBL3235 | 10927 A. (D2-U-S13) | 12189 A. (D2_U_S13) |  | - | - | - | - | - | - | - | - | - | - |
| 2 | MBAL2129 | 10659 A. (D2-U-S13) | 11823 A. (D2-U-S13) | 12128 A. (D1-L-S05) | 12342 A. (D1_L_S07) | 12566 A. (D1_U_S01) | 12717 A. (D1-U-S12) | 12708 A. (D2-U-S13) | 12702 A. (D2-U-S13) |  | - | - | - | - |
| 3 | MBBL2121 | 10725.7 A. (D1-L) | 11486.7 A. (D2_L) | 11930.8 A. (D2-L) | 12331 A. (D2-L) | 12333 A. (D1-L) | 12515 A. (D2-U) | 12663 A. (D2-L) | 12829 A. (D2_U) | 12846 A. | - | - | - | - |
| 4 | MBBL1196 | 11270.9 A. (D1_L) | 11670 A. (D1 _ L) | 11807 A. (D1_L) | 11830 A. (D1 _ U) | 11964 A. (D1_L) | 12036 A. (D1-U) | 12191.4 A. (D1_L) | 12028 A. (D1-L) | 12301.5 A. (D1-L) |  | - | - | - |
| 5 | MBAL2151 | 11886.8 A. (D2-U) | 12615 A. (D2-L) |  | - | - | - | - | - | - | - | - | - | - |
| 6 | SSS073 | 11456 A. (QI_12) | 12521 A. (QI_12) | 12650 A. (QE_12) | 12800 A. | - | - | - | - | - | - | - | - | - |
| 7 | SSS100 | 12031 A. (QE_34) | 12161 A. (QE_34) |  | - | - | - | - | - | - | - | - | - | - |
| 8 | SSS080 | 12567.5 A. (QE_12) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 9 | MBBL2147 | 11386.4 A. (D1-L) | 11763.2 A. (D1-L) | 12336.3 A. (D2-U) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 10 | MBAL2148 | 11496 A. (D1 - L) | 0 A. (No Data) | 12413 A. (D1-L) |  | - | - | - | - | - | - | - | - | - |
| 11 | MBBL1199 | 10553.3 A. (D1_L) | 11694 A. (D1-L) | 12030 A. (D1-L) | 12327.6 A. (D2-L) | 12461 A. (D2-U) | 12620 A. (D2_U) | 12728 A. (D2-U) | 12820 A. (D2-L) | 12789 A. (D2-U) |  | - | - | - |
| 12 | MBAL2131 | 10676 A. (D1-L) | 11583.3 A. (D1_U) | 12548 A. (D1_L) | 12823 A. (D1_U) | 12800 A. (D1_U) | 12852 A. (D1_U) | 12850 A. | - | - | - | - | - | - |

### Week 27 Feb05 to 5 March 05

| From 27-02-2005 to 05-03-2005 |
| --- |
| - | Bench | Magnet Name | shafts | Start Date . | Connect Time | Wait +Cooling | Cold Tests | Wait +Warming | Disconnect Time | End Date . | Total | Q |
| 1 | TBF2 | MBAL2131 | False | 25-02-2005 | 29 h. | 37 h. | 46 h. | 13 h. | 11 h. | 03-03-2005 | 6 days | 6 |
| 2 | TBF1 | MBBL1199 | False | 26-02-2005 | 33 h. | 40 h. | 83 h. | 13 h. | 14 h. | 05-03-2005 | 7 days | 9 |
| 3 | TBE2 | MBAL2148 | False | 26-02-2005 | 27 h. | 39 h. | 21 h. | 14 h. | 9 h. | 03-03-2005 | 5 days | 3 |
| 4 | TBE1 | MBBL2147 | False | 27-02-2005 | 37 h. | 26 h. | 35 h. | 17 h. | 8 h. | 04-03-2005 | 5 days | 3 |
| 5 | TBD2 | SSS080 | False | 25-02-2005 | 41 h. | 29 h. | 17 h. | 9 h. | 18 h. | 02-03-2005 | 5 days | 1 |
| 6 | TBD1 | SSS073 | True | 23-02-2005 | 15 h. | 13 h. | 59 h. | 8 h. | 13 h. | 28-02-2005 | 5 days | 3 |
| 7 | TBD1 | SSS100 | False | 28-02-2005 | 41 h. | 17 h. | 31 h. | 7 h. | 15 h. | 05-03-2005 | 5 days | 2 |
| 8 | TBC2 | MBAL2151 | False | 25-02-2005 | 24 h. | 30 h. | 18 h. | 14 h. | 16 h. | 01-03-2005 | 4 days | 2 |
| 9 | TBC1 | MBBL1196 | False | 27-02-2005 | 22 h. | 31 h. | 46 h. | 16 h. | 10 h. | 04-03-2005 | 5 days | 9 |
| 10 | TBB2 | MBBL2121 | False | 23-02-2005 | 58 h. | 22 h. | 46 h. | 16 h. | 8 h. | 02-03-2005 | 7 days | 8 |
| 11 | TBA2 | MBAL2129 | True | 24-02-2005 | 38 h. | 30 h. | 69 h. | 17 h. | 0 h. | 02-03-2005 | 6 days | 8 |
| 12 | TBA1 | MBBL3235 | True | 25-02-2005 | 30 h. | 83 h. | 39 h. | 12 h. | 10 h. | 04-03-2005 | 7 days | 2 |

| Period | Dipoles |
| --- | --- |
|  | Fresh | Repeat | Total | MM | % Repeat | % MM |
| Year 2002-2003 | 95 | 0 | 95 | 0 | 0.00 | 0.00 |
| Year 2004 | 358 | 38 | 396 | 65 | 10.61 | 18.16 |
| Year 2005 | 45 | 4 | 49 | 4 | 8.89 | 8.89 |
| Year 2006 | 0 | 0 | 0 | 0 | 0.00 | 0.00 |
| Year 2007 | 0 | 0 | 0 | 0 | 0.00 | 0.00 |
| Total | 498 | 42 | 540 | 69 | 8.43 | 13.86 |
|  |  |  | 1232 |  |  |  |
| LHC Statistics |  |  |  |  |  |
| Total Fresh Magnets Tested | 566 |  | out of Expected 1894 ( 1706 + 11% ) |
| Total Repeat Magnets Tested | 47 |  |  |  |  |  |
| Total Magnets Tests | 613 |  |  |  |  |  |
| Total % Retested | 8.30 |  |  |  |  |  |
| ARC-SSS |  | IR-SSS |
| Fresh | Repeat | Total | MM | % Repeat | % MM |  | Fresh | Repeat | Total | % Repeat |
| 2 | 0 | 2 | 0 | 0.00 | 0.00 |  | 0 | 0 | 0 | 0.00 |
| 49 | 3 | 52 | 9 | 6.12 | 18.37 |  | 3 | 0 | 3 | 0.00 |
| 13 | 1 | 14 | 1 | 7.69 | 7.69 |  | 1 | 1 | 2 | 100.00 |
| 0 | 0 | 0 | 0 | 0.00 | 0.00 |  | 0 | 0 | 0 | 0.00 |
| 0 | 0 | 0 | 0 | 0.00 | 0.00 |  | 0 | 0 | 0 | 0.00 |
| 64 | 4 | 68 | 10 | 6.25 | 15.63 |  | 4 | 1 | 5 | 25.00 |
|  |  | 360 |  |  |  |  |  |  | 114 |  |

OP statistics for SM18 OP during the period from 28/02/2005 06:00:00 to 07/03/2005 13:00:00

| Lines | in Super Cycle | in Faults | Availabilities (%) |
| --- | --- | --- | --- |
| TBA1 | 161 [h] | 9 [h] 04 [min] 14[s] | 94.37 |
| TBA2 | 161 [h] | 20 [h] 58 [min] 52[s] | 86.97 |
| TBB1 | 161 [h] | 3 [h] 42 [min] 57[s] | 97.69 |
| TBB2 | 161 [h] | 3 [h] 29 [min] 48[s] | 97.83 |
| TBC1 | 161 [h] | 2 [h] 47 [min] 44[s] | 98.26 |
| TBC2 | 161 [h] | 2 [h] 47 [min] 44[s] | 98.26 |
| TBD1 | 161 [h] | 6 [h] 27 [min] 44[s] | 95.99 |
| TBD2 | 161 [h] | 4 [h] 03 [min] 11[s] | 97.48 |
| TBE1 | 161 [h] | 2 [h] 47 [min] 44[s] | 98.26 |
| TBE2 | 161 [h] | 2 [h] 47 [min] 44[s] | 98.26 |
| TBF1 | 161 [h] | 9 [h] 22 [min] 40[s] | 94.18 |
| TBF2 | 161 [h] | 2 [h] 47 [min] 44[s] | 98.26 |
| TOTAL: | 1932 [h] | 71 [h] 08 [min] 06[s] | 96.32 |
| Faults Groups | Faults Names | Duration |
| Cryogenics | Cryo NOT OK | 45 [min] |
|  | Miscellaneous | 19 [h] 33 [min] 01[s] |
| Electronics | ----- | 42 [min] 04[s] |
| HW recognition | ----- | 4 [h] 16 [min] 02[s] |
| Hardware | TRU and shafts | 6 [h] 40 [min] 51[s] |
| Miscellaneous | ----- | 7 [h] 51 [min] 08[s] |
| Rocla | ----- | 34 [min] 56[s] |
| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| Tuesday _01_A | TBC1 | No quench at 1.5 kA with lf heaters in MB1196 | MTM |  | Guy advised to proceed with a trial at 2000 A |
| Tuesday _01_A | TBD2 | HWREC error on IAP test | MTM | HWREC | Guy switched off HWREC. |
| Wednesday_02_A | TBD1 | Magnet name problem with SSS100 . name refused by MTF, NC was opened Monday by Maryline. | ICS |  | Magnet name entered manually for electrical warm tests. HWREC switched off by Guy. Configuration files generated manually Test without HWREC. |
| Wednesday _02_A | TBD1 | HF heater quench at 15 kA: no quench detected | MTM |  | Test repeated, same result. Test repeated at 2000 A. |
| Wednesday _02_N | TBC1 | SSL @ 4.4 K. quench at 9262 A.....Normal value is >10100A at PT102 Pressure of 1350 mbar. | MTM |  | Test repeated |
| Thursday _03_M | all | Liquefactor problem: cryo OK will be lost soon. Stephane decided to stop the load line in TBA1 | MTM |  | 3 hours of stop |
| Thursday _03_A | TBA1 | SPA at 5000 A during the ramp down on the loadline performed on the MB 3235 Stephane found that the wrong setting was used for the full load line (that of TBA2), Mistake in the preparation of the setting. | MTM |  | Stephane informed. He checked the current level of the stop and decided not to repeat the test. Guy informed. With Maryline he will correct the magnet name, shaft name. |
| Friday _04_A | TBA, TBC | The option “ repeat magnet” in TEMA is not always working properly :store sometimes results in wrong directories. | AB/CO | TEMA | Re direction of the results of the MB2129 in the folder MB2129_2. |
| Saturday_05_M | TBA2 | Four quenches outside the coils. Cryo problem? During LHC cycle, quench at 7900 A out of the main coil, while ramping up, due to cryo electrovalve | ACR |  | J:P:Lamboy called. Current leads B Temaperature not stable. Then Quench OK. |
| Saturday_05_M | TBB1 | Survey team not available for the survey of the SSS601. | MTM |  | Survey performed by Guy. |
| Saturday_05_M | TBD1 | final cold HV failed during the first step: quadrupole vs ground (other steps are OK). | MTM |  | Guy called and checked. He asked to open a NC. |
| Saturday_05_M | TBB2 | SSS601 : SSW measurement not correct | MTM |  | Guy called and checked the results. He found that the file name was not correct. |
| Saturday_05_A | TBF1 | Magnet 1199: WP07 not signed after completion by ICS. The mention “stand by “ could not be put by OP team. It was removed from the bench without the traveller signed “ stripped or stand by “. New magnet 3225 coming: It came two day after the cold test and it is signed “ stripped”. | ICS |  |  |

Top of Form

Bottom of Form