Vc/29-Mar-05

### Week 13  starting 28  March 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  28 March 05 @ 3/28/2005 6:43:58 PM

| 1 | A1 | MBBL1196 | PT 13.2 Resist. Meas. | 299.65 |
| --- | --- | --- | --- | --- |
| 3 | A2 | MBBL3130 | PT 6.3 Training 3 | 3.95 HV |
| 7 | B1 | SSS125 | PT 12 Warm Up | 202.58 |
| 12 | B2 | MBBL2157 | PT 1 IAP @ cold | 1.90 |
| 10 | C1 | MBAL1200 | ICS 1 WP04 HV Test Warm | 294.1 |
| 8 | C2 | MBAL2125 | PT 6.1 Training 1 | 1.90 |
| 6 | D1 | SSS028 | PT 14 Magnet MAPS | 294.91 |
| 5 | D2 | SSS025 | PT 6.7 Training 7 | 1.91 |
| 11 | E1 | MBBR3294 | Finished Ready to go Final scan | 295.89 |
| 4 | E2 | MBBR3312 | Finished Ready to go Final scan | 295.8 |
| 2 | F1 | MBAR3299 | PT 13.2 Resist. Meas. | 300.82 |
| 9 | F2 | MBBL2162 | Prep 5 Pumping + Cool Down | 5.49 HV |

Temperature data are taken from TBSM18 site

All in RED are Finished Magnets , some since  Many Hours ( even yesterday, Sunday) & there are no new magnets coming in ( ICS, MRB, Rocla etc problems over Easter: Ask Vittorio !)

Coordinator this week: Stephane; MTM support on call :   16-3602 .

15 Magnets Tested

11  dipoles( 1 with MM)+ 3  SSS(0  with MM) +  1  SSSS (with MM)

Disconnected after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Situation on Repair :

General Points

IAP  problems     R init 20.7    etc etc   …

AQA & Alarms & New Version Distribution    -- when ??

SSSS  Reporting for Strip/standby

AOB

## Tests

| Magnets under Test from 20-03-2005 to 26-03-2005 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total |  |  |
| 1 | TBA1 | MBBL1184 |  |  | 14-03-2005 | 32 h. | 43 h. | 42 h. | 15 h. | 10 h. | 20-03-2005 | 142 h. |  |  |
| 2 | TBA2 | MBBR3275 |  |  | 18-03-2005 | 56 h. | 30 h. | 50 h. | 16 h. | 12 h. | 24-03-2005 | 164 h. |  |  |
| 3 | TBB1 | SSS602 |  |  | 14-03-2005 | 77 h. | 15 h. | 119 h. | 8 h. | 26 h. | 25-03-2005 | 245 h. |  |  |
| 4 | TBB2 | MBAL1202 |  |  | 20-03-2005 | 22 h. | 28 h. | 32 h. | 19 h. | 24 h. | 25-03-2005 | 125 h. |  |  |
| 5 | TBB2 | MBAR3251 |  |  | 14-03-2005 | 30 h. | 27 h. | 38 h. | 28 h. | 16 h. | 20-03-2005 | 139 h. |  |  |
| 6 | TBC1 | MBAL1209 |  |  | 17-03-2005 | 29 h. | 30 h. | 17 h. | 14 h. | 10 h. | 21-03-2005 | 100 h. |  |  |
| 7 | TBC2 | MBAL1204 |  |  | 17-03-2005 | 37 h. | 22 h. | 14 h. | 14 h. | 9 h. | 21-03-2005 | 96 h. |  |  |
| 8 | TBC2 | MBBL1165 |  |  | 21-03-2005 | 26 h. | 25 h. | 16 h. | 18 h. | 11 h. | 25-03-2005 | 96 h. |  |  |
| 9 | TBD1 | SSS103 |  |  | 18-03-2005 | 37 h. | 15 h. | 34 h. | 9 h. | 17 h. | 22-03-2005 | 112 h. |  |  |
| 10 | TBD1 | SSS103 |  |  | 22-03-2005 | 2 h. | - | - | - | - | 23-03-2005 | 14 h. |  |  |
| 11 | TBD2 | SSS019 |  |  | 20-03-2005 | 41 h. | 18 h. | 19 h. | 9 h. | 13 h. | 24-03-2005 | 100 h. |  |  |
| 12 | TBD2 | SSS128 |  |  | 15-03-2005 | 30 h. | 21 h. | 35 h. | 9 h. | 25 h. | 20-03-2005 | 120 h. |  |  |
| 13 | TBE2 | MBBR3293 |  |  | 17-03-2005 | 31 h. | 26 h. | 37 h. | 13 h. | 10 h. | 22-03-2005 | 117 h. |  |  |
| 14 | TBF1 | MBAL1203 |  |  | 19-03-2005 | 31 h. | 25 h. | 29 h. | 14 h. | 10 h. | 24-03-2005 | 109 h. |  |  |
| 15 | TBF2 | MBAL2101 |  |  | 20-03-2005 | 19 h. | 30 h. | 42 h. | 13 h. | 16 h. | 25-03-2005 | 120 h. |  |  |
| 16 | TBF2 | MBAR3303 |  |  | 11-03-2005 | 22 h. | 78 h. | 93 h. | 13 h. | 7 h. | 20-03-2005 | 213 h. |  |  |

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBL1184 | 12844.1 A. (D2-U-S01) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 2 | MBBR3275 | 9193 A. (D1-L-S01) | 11685 A. (D2-L-S01) | 12322 A. (D2-L-S01) | 12738 A. (D2-L-S01) | 12763 A. (D2 - U) | 12851 A. | - | - | - | - | - | - | - |
| 3 | SSS602 | 5430 A. (No Data) | 5850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 4 | MBAL1202 | 12133.5 A. (D1-L) | 12284 A. (D2_U) |  | - | - | - | - | - | - | - | - | - | - |
| 5 | MBAR3251 | 10417 A. (D2-L) | 11356 A. (D2-U) | 11980 A. (D1-U) | 12678 A. (D2-L) | 12618 A. (D1-L) | 12604 A. (D1-L) | 12750 A. (D1_L) | 12850 A. | - | - | - | - | - |
| 6 | MBAL1209 | 12129 A. (D2_L) | 12513 A. (D2_U) |  | - | - | - | - | - | - | - | - | - | - |
| 7 | MBAL1204 | 12196 A. (D1_L) | 12386 A. (D1-L) |  | - | - | - | - | - | - | - | - | - | - |
| 8 | MBBL1165 | 11750 A. (D1_U) | 12393 A. (D2-L) |  | - | - | - | - | - | - | - | - | - | - |
| 9 | SSS103 | 12000 A. (QE-34) | 11566.3 A. (QE-34) | 12056.3 A. (QE_12) | 12569 A. (QE-12) | 12686 A. (QI_12) | 12771.2 A. (QI-34) | 12812 A. (QE - 12) | 12638 A. (QE - 12) | 12850 A. | - | - | - | - |
| 10 | SSS103 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 11 | SSS019 | 12588 A. (QE-34) | 12462 A. (QE_12) |  | - | - | - | - | - | - | - | - | - | - |
| 12 | SSS128 | 11852 A. (QE - 34) | 11796 A. (QE_34) | 12349 A. (QI_34) | 12551 A. (QI-12) | 12681 A. (QE-12) | 12803 A. (QI - 12) | 12850 A. | - | - | - | - | - | - |
| 13 | MBBR3293 | 11568.8 A. (D2-U) | 11770.9 A. (D1-U) | 11910 A. (D2_L) | 12109 A. (D2-L) | 12249 A. (D2-U) | 12142 A. (D2-L) | 12191.6 A. (D2-L) | 12448.9 A. (D1_U) | 12550 A. (D2-U) |  | - | - | - |
| 14 | MBAL1203 | 12221.8 A. (D1_U) | 12413 A. (D2-L) |  | - | - | - | - | - | - | - | - | - | - |
| 15 | MBAL2101 | 10372.1 A. (D2_L) | 12186.6 A. (D1_L) |  | - | - | - | - | - | - | - | - | - | - |
| 16 | MBAR3303 | 11417 A. (D1-L) | 11678 A. (D1 - U) | 12130 A. (D2-U) | 12096 A. (D2 - U) | 12331 A. (D2-U-S01) | 12420 A. (D2-L) | 12287.8 A. (D2_L) | 12610 A. (D2_U) | 12646 A. (D2_U) |  | - | - | - |

### Week 13 march 20 to 26 March 05

### [3275 waiting for 1184 or was it special MM or MM problems?

### 3303 long cooldown problem  or cohabitation ?  ]

| From 20-03-2005 to 26-03-2005 |
| --- |
| - | Bench | Magnet Name | shafts | Start Date . | Connect Time | Wait +Cooling | Cold Tests | Wait +Warming | Disconnect Time | End Date . | Total | Q |
| 1 | TBF2 | MBAL2101 | False | 20-03-2005 | 19 h. | 30 h. | 42 h. | 13 h. | 16 h. | 25-03-2005 | 120 h. | 2 |
| 2 | TBF2 | MBAR3303 | False | 11-03-2005 | 22 h. | 78 h. | 93 h. | 13 h. | 7 h. | 20-03-2005 | 213 h. | 9 |
| 3 | TBF1 | MBAL1203 | False | 19-03-2005 | 31 h. | 25 h. | 29 h. | 14 h. | 10 h. | 24-03-2005 | 109 h. | 2 |
| 4 | TBE2 | MBBR3293 | False | 17-03-2005 | 31 h. | 26 h. | 37 h. | 13 h. | 10 h. | 22-03-2005 | 117 h. | 9 |
| 5 | TBD2 | SSS019 | False | 20-03-2005 | 41 h. | 18 h. | 19 h. | 9 h. | 13 h. | 24-03-2005 | 100 h. | 2 |
| 6 | TBD2 | SSS128 | False | 15-03-2005 | 30 h. | 21 h. | 35 h. | 9 h. | 25 h. | 20-03-2005 | 120 h. | 6 |
| 7 | TBD1 | SSS103 | False | 18-03-2005 | 37 h. | 15 h. | 34 h. | 9 h. | 17 h. | 22-03-2005 | 112 h. | 8 |
| 8 | TBD1 | SSS103 | False | 22-03-2005 | 2 h. | 137 h. | -136 h. | 4 h. | 7 h. | 23-03-2005 | 14 h. | 0 |
| 9 | TBC2 | MBAL1204 | False | 17-03-2005 | 37 h. | 22 h. | 14 h. | 14 h. | 9 h. | 21-03-2005 | 96 h. | 2 |
| 10 | TBC2 | MBBL1165 | False | 21-03-2005 | 26 h. | 25 h. | 16 h. | 18 h. | 11 h. | 25-03-2005 | 96 h. | 2 |
| 11 | TBC1 | MBAL1209 | False | 17-03-2005 | 29 h. | 30 h. | 17 h. | 14 h. | 10 h. | 21-03-2005 | 100 h. | 2 |
| 12 | TBB2 | MBAL1202 | False | 20-03-2005 | 22 h. | 28 h. | 32 h. | 19 h. | 24 h. | 25-03-2005 | 125 h. | 2 |
| 13 | TBB2 | MBAR3251 | False | 14-03-2005 | 30 h. | 27 h. | 38 h. | 28 h. | 16 h. | 20-03-2005 | 139 h. | 7 |
| 14 | TBB1 | SSS602 | True | 14-03-2005 | 77 h. | 15 h. | 119 h. | 8 h. | 26 h. | 25-03-2005 | 245 h. | 1 |
| 15 | TBA2 | MBBR3275 | True | 18-03-2005 | 56 h. | 30 h. | 50 h. | 16 h. | 12 h. | 24-03-2005 | 164 h. | 5 |
| 16 | TBA1 | MBBL1184 | False | 14-03-2005 | 32 h. | 43 h. | 42 h. | 15 h. | 10 h. | 20-03-2005 | 142 h. | 1 |

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _22_A | TBA2 | Over-ranges in 3275 magnet during MM1 | MTM |  | Connector not screwed properly. |
| Tuesday _22_A | TBA2 | QH discharge. Temperature of Lead B is not correct. | ACR |  |  |
| Tuesday _22_A | TBB2 | HV test: automatic loading of the magnet name not possible. Error : PLC-SE does not allow to start the test. | MTM | HV | Richard loaded the name manually. Investigation by Richard. |
| Tuesday _22_N | TBE | Connection between PLC and HV rack MY93 not possible. | ACR |  |  |
| Wednesday _23_M | TBF2 | Cryo OK lost. | ACE |  | Decrease of He liquid in the CFB |
| Wednesday _23_M | TBD2 | Corrector circuit cannot to be powered. Trigger generated every time. | ICS |  | Connection not correct. TC decided to end up the cold test. |
| Wednesday _23_M | TBD1 | P11 connector has been broken during magnet movement. | MTM |  |  |
| Wednesday _23_A | TBA2 | SPA occurred on the training number 6. | ACR? |  | Cryo problem? |
| Wednesday _23_A | TBA1 | MITTS alert on 3 kA diode test. | MTM | AQA | Pierre contacted. MIITS limit increased with the new AQA version 27 instead of 23 so the alarm became a warning. |
| Thursday_24_M | TBC2 | Training 1: Test could not be started 2 times | MTM | TEMA | TEMA switch off and re-started. Then OK. |
| Thursday_24_A | TBC1 | SPA at 1 kA: unable to launch it due to interlocks “leadTempA_X1” and “LeadTempB_x1”which are not re-settable. | ACR/AB-CO |  | Maryline called. CTH is sending the signal which creates the interlocks. C.Dehavay has to decrease the thresholds of these signals. Software will be modified by A.Tovar Tuesday. |
| Thursday_24_A Friday:25_M | TBB2 TBE | IAP at warm after test not launched. DAQ could not be armed. Same problem | AB/CO | Hf,lf | Local sun server switched off/on and TEMA re-launched. |
| Friday_25_M | TBD2 | Rinit to low for QH1&4 (20 ohm) | MTM | hfoff | Maryline informed. She advised to use the hfoff application of the sun MTA 24: values OK now. |
| Friday_25_A | TBD2 | SSS25 : HV test at warm: YT311 vs ground has a wrong value of capacitance. | MTM |  | Maryline informed. Investigation through the IFS box. |
| Friday_25_A | TBD1 | Cryo OK lost during the powering of the EF circuit. | ACR |  | Cryo lost. TC advised to re-launch the test. |
| Friday_25_N | TBB1 | SSS125: QH could not be charged. | ICS |  | Maryline informed.Vtaps in the internal lead clamp not connected. ICS did the connection again. |
| Friday_25_N | TBA1 | 1196 magnet : lf QH test : Iearth and QH signals not found. | AB/CO |  | Maryline advised to reset the crates before proceeding with the next quench. |
| Saturday_26_N | TBA1 | Training quench number 5: -lf file not present in the control room sun. - QH timing error with AQA found in hf file | AB/CO |  | Analysis performed on lf file in local sun: signals OK. |
| Sunday_27_M | TBA1 | DAQ could not be loaded correctly. | AB/CO |  | Reset the crates but hf DAQ cannot be armed. Reset again but then MXI communication error. Maryline changed the first card of the hf crate. |
| Sunday_27_M | TBD2 | SSS025:Trigger which prevents to power the EF correctors | MTM |  | Training of the MQ proceeded. Maryline is informed. She shorted the signals in the security matrix. |
| Sunday_27_N | TBB1 | HF quench at 1.5 kA on SSS125: QH3 and QH4 did not discharge. Test repeated: same result. | MTM |  | Maryline suggested to reset all the DAQ system (crates, lf, hf applications,): same result. Provoked quench at 2000A. Hf trigger is received and all Heaters discharged. AQA alarm tFire<-2000 |
| Sunday_27_N | TBB1 | Powering EF circuit not possible: | MTM |  | Maryline changed the inductance value~3 mH |
| Monday_28_M | TBD2 | Slow PA during training ramp | ? |  |  |