Vc/17-May-05

### Week 18  starting 2 May 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  02 May 05 @ 5/2/2005 3:04:22 PM

| 3 | A1 | MBAL2126 | PT 12 Warm Up | 4.25 |
| --- | --- | --- | --- | --- |
| 12 | A2 | MBAL2135 | PT 13.2 Resist. Meas. | 298.67 |
| 2 | B1 | SSS090 | PT 5.2 1.5 KA Quench HF | 1.90 |
| 5 | B2 | SSS171 | PT 12 Warm Up | 292.85 |
| 10 | C1 | MBAR3325 | Prep 5 Cool Down | 297.59 |
| 7 | C2 | MBAR3324 | Prep 5 Cool Down | 212.28 |
| 11 | D1 | SSS126 | PT 14 Magnet MAPS | 297.41 |
| 1 | D2 | SSS615 | PT 7 Magnet Trained | 1.90 |
| 8 | E1 | MBAL2172 | PT 12 Warm Up | 267.38 |
| 9 | E2 | MBAL2173 | Prep 5 Cool Down | 220.62 |
| 4 | F1 | MBAL1179 | PTE Min. NRJ 3KA Diode Test | 1.91 |
| 6 | F2 | MBBL2164 | Prep 5 Cool Down | 135.88 |

Coordinator this week: Marta; MTM support on call :   16-3602

18  complete Magnet Tests cycles   ( 2171 ROBTC)

12  dipoles( 1 with MM)+ 6  SSS( 0  with MM) +  0  SSSS (with MM)

Disconnected or ROBTC after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Situation on Repair :   J Halik

- General Points

AOB

## Tests

| Magnets under Test from 24-04-2005 to 30-04-2005 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBBR3331 |  |  | 23-04-2005 | 0 h. | 21 h. | 23 h. | 16 h. | 19 h. | 26-04-2005 | 79 h. | Stripping |  |  |
| 2 | TBA2 | MBBR3293 |  |  | 20-04-2005 | 34 h. | 25 h. | 43 h. | 16 h. | 12 h. | 25-04-2005 | 130 h. | Stand By |  |  |
| 3 | TBB1 | SSS096 |  |  | 20-04-2005 | 26 h. | 26 h. | 18 h. | 9 h. | 28 h. | 25-04-2005 | 107 h. | Stripping |  |  |
| 4 | TBB1 | SSS176 |  |  | 26-04-2005 | 35 h. | 11 h. | 26 h. | 9 h. | 16 h. | 30-04-2005 | 97 h. | Stripping |  |  |
| 5 | TBB2 | SSS079 |  |  | 24-04-2005 | 27 h. | 13 h. | 26 h. | 8 h. | 16 h. | 27-04-2005 | 90 h. | Stripping |  |  |
| 6 | TBC1 | MBBL1127 |  |  | 22-04-2005 | 23 h. | 27 h. | 38 h. | 14 h. | 14 h. | 26-04-2005 | 116 h. | Stripping |  |  |
| 7 | TBC2 | MBBL1105 |  |  | 21-04-2005 | 21 h. | 23 h. | 35 h. | 14 h. | 24 h. | 26-04-2005 | 117 h. | Stripping |  |  |
| 8 | TBC2 | MBBL1220 |  |  | 26-04-2005 | 43 h. | 24 h. | 13 h. | 14 h. | 8 h. | 30-04-2005 | 102 h. | Stripping |  |  |
| 9 | TBD1 | SSS105 |  |  | 21-04-2005 | 22 h. | 25 h. | 16 h. | 7 h. | 18 h. | 25-04-2005 | 88 h. | Stripping |  |  |
| 10 | TBD1 | SSS175 |  |  | 25-04-2005 | 25 h. | 11 h. | 16 h. | 7 h. | 12 h. | 28-04-2005 | 71 h. | Stand By |  |  |
| 11 | TBD2 | SSS070 |  |  | 20-04-2005 | 31 h. | 12 h. | 26 h. | 7 h. | 10 h. | 25-04-2005 | 86 h. | Stripping |  |  |
| 12 | TBE1 | MBBL2164 |  |  | 22-04-2005 | 0 h. | 24 h. | 8 h. | 16 h. | 14 h. | 24-04-2005 | 62 h. | Stand By |  |  |
| 13 | TBE1 | MBBR3317 |  |  | 24-04-2005 | 22 h. | 20 h. | 14 h. | 12 h. | 17 h. | 28-04-2005 | 85 h. | Stripping |  |  |
| 14 | TBE2 | MBAR3326 |  |  | 21-04-2005 | 25 h. | 39 h. | 18 h. | 13 h. | 12 h. | 25-04-2005 | 107 h. | Stripping |  |  |
| 15 | TBE2 | MBBL1219 |  |  | 25-04-2005 | 46 h. | 25 h. | 17 h. | 16 h. | 15 h. | 30-04-2005 | 119 h. | Stripping |  |  |
| 16 | TBF1 | MBBR3310 |  |  | 22-04-2005 | 22 h. | 22 h. | 22 h. | 14 h. | 20 h. | 27-04-2005 | 100 h. | Stripping |  |  |
| 17 | TBF2 | MBAL2171 |  |  | 22-04-2005 | 23 h. | 29 h. | 45 h. | - | - | 27-04-2005 | 97 h. | Not Found or Thermal cycle |  |  |
| 18 | TBF2 | MBAL2171 |  |  | 27-04-2005 | 0 h. | 15 h. | 11 h. | 15 h. | 6 h. | 29-04-2005 | 47 h. | Stripping |  |  |
| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| 1 | MBBR3331 | 12336.4 A. (D1-U) | 12650 A. (D2-U) |  | - | - | - | - | - | - | - | - | - | - |
| 2 | MBBR3293 | 10616 A. (D1 - L) | 12512 A. (D2-U) | 12827.5 A. (D2-U-S01) |  | - | - | - | - | - | - | - | - | - |
| 3 | SSS096 | 11333 A. (QI-12) | 12319 A. (QI-34) | 12789 A. (QI-12) |  | - | - | - | - | - | - | - | - | - |
| 4 | SSS176 | 11204 A. (QI-12) | 11406 A. (Q_E_12) | 11702 A. (Q_E_34) | 12250 A. (Q_I_12) | 12659 A. (Q_E_12) | 12839.1 A. (Qe-12) | 12850 A. | - | - | - | - | - | - |
| 5 | SSS079 | 11413 A. (QE-34) | 11852 A. (QI-34) | 12383 A. (QI-34) | 12515.7 A. (QI-34) | 12628 A. (QI-34) |  | - | - | - | - | - | - | - |
| 6 | MBBL1127 | 11665.7 A. (D1-L) | 12269 A. (D2 - U) | 12578 A. (D1_U) |  | - | - | - | - | - | - | - | - | - |
| 7 | MBBL1105 | 12189.2 A. (D2-L) | 12404.3 A. (D1-L) | 12572.9 A. (D2-U) |  | - | - | - | - | - | - | - | - | - |
| 8 | MBBL1220 | 11795 A. (D2 - L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 9 | SSS105 | 12140 A. (QI-34) | 12452 A. (QE-34) | 12628 A. (QI_12) |  | - | - | - | - | - | - | - | - | - |
| 10 | SSS175 | 12344 A. (QE-34) | 12620 A. (QI-34) | 12688 A. (QE-12) |  | - | - | - | - | - | - | - | - | - |
| 11 | SSS070 | 11234.8 A. (QI_34) | 11228 A. (QE_12) | 11923 A. (QE-34) | 12360 A. (QI-34) | 12599 A. (QI-12) | 12850 A. | - | - | - | - | - | - | - |
| 12 | MBBL2164 | 11949 A. (D1 - L) | 11537 A. (D1 - L) |  | - | - | - | - | - | - | - | - | - | - |
| 13 | MBBR3317 | 10577.1 A. (D2-U) | 12303 A. (D2-U) | 12333.1 A. (D1-U) |  | - | - | - | - | - | - | - | - | - |
| 14 | MBAR3326 | 11746 A. (D1-U) | 11791 A. (D1-U) | 12608 A. (D2-L) | 12832 A. (D2-U) |  | - | - | - | - | - | - | - | - |
| 15 | MBBL1219 | 12518 A. (D2_U) | 12630 A. (D2_L) | 12794 A. (D2_L) |  | - | - | - | - | - | - | - | - | - |
| 16 | MBBR3310 | 9130 A. (D2-U) | 12073 A. (D2-U) | 12007 A. (D2-U) | 12391 A. (D1-L) | 12525 A. (D2-U) |  | - | - | - | - | - | - | - |
| 17 | MBAL2171 | 10169 A. (D2-U) | 11326 A. (D2_L) | 12034 A. (D1-L) | 12254.6 A. (D2-U) | 11749 A. (D2-U) | 11839 A. (D2-U) | 11849.2 A. (D2-U) | 12388 A. (D2_U) | 12630 A. (D1_U) |  | - | - | - |
| 18 | MBAL2171 | 12162 A. (D1-L) | 12390 A. (D2_U) |  | - | - | - | - | - | - | - | - | - | - |

### Week 17  :  24 April to 30 April 05  (18 Magnet Tests)

### [   2171 On bench rapid thermal cycle   ]

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Wednesday _27_M | TBD1 | Quench at 520 A during the powering of the EF corrector circuit of SSS175. | ICS? |  | Investigation by Giorgio. He suggested to open a non conformity because a bad clamping. |
| Wednesday _27_M | All | Oracle problem: MTF stopped | IT | MTF |  |
| Thursday _28_N | TBA | DAQ blocked. Resetting the crates is not sufficient. | AB/CO | Hf, lf | A.Rijllart came and changed the MXI card and the cards in the crate. |
| Thursday _28_N | TBB1 | Connection of voltage taps in the clamp for SSS176 is not good, detected by the corrector test at warm | ICS |  | Sleeve re-opened and reconnection done. |
| Thursday _28_N | TBD1 | One of the voltage tap wires for the main Quad is broken too far into the CFB to be re-soldered. | ICS |  | Maryline called Nikolay who asked to continue the connection without the VT. |
| Thursday _28_N | TBD2 | One of the voltage tap wires for the main Quad is broken too far into the CFB to be re-soldered. | ICS |  | Maryline called Nikolay who asked to continue the connection without the VT. |
| Friday _29_M | TBD1 | Permanent trigger because of configuration of bench corresponding to S4 | MTM |  |  |
| Friday _29_M | TBE1 | Not MTF: cannot upload WP04 | IT | MTF |  |
| Friday _29_A | TBA2 | IAP errors: Q_loc offset | MTM |  | Maryline solved the problem by replacing the cable from the TRU to the rack. |
| Friday _29_N | TBA2 | Faulty quench on MB2135 trigger by potaim card 3 (Vbus EXT-Vbus INT) | MTM |  |  |
| Saturday_30_M | TBF2 | Voltage during lyre test is 12.06V. | MTM |  | Maryline gave the OK to continue Voltage at warm acceptable as long as it stays stable during the test. |
| Saturday_30_M | TBC1 | MB3158: Data File ( L0504301147.ka01) not transferred to storage. | AB/CO MTM |  | Maryline will transfer it Monday. |
| Saturday_30_A | TBF2 | Continuous message coming from the PLCSE. | MTM? |  | Maryline advised to restart Tema, Server, PLC in this sequence. |
| Saturday_30_A | TBF2 | HVrack MY92 -- connector D20 broken Lack of rack for HV tests. | MTM |  |  |
| Sunday_1_M | TBB2 | Training could not be launched because of TCLaux_X2 and Cryo1_X2 signal not ok for PLCSE. | ACR |  | Problem rectified by the cryo-team. |
| Sunday_1_M | TBB2 | 4 channels missing in hf DAQ | AB/CO |  | Maryline investigated the problem: one HF card faulty (to be confirmed) |
| Sunday_1_A | TBF1 | HV tests: Profibus is not working. | ]MTM |  | Maryline disabled the interlocks signals coming from the PLCSE to the PLC of the HV rack until the repair Monday Morning. |