Vc/13-Jun-05

### Week 23  starting 6 June 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  06 June 05 @ 6/6/2005 1:46:37 PM

| 5 | A1 | MBAL3101 | PT 5.2 1.5 KA Quench HF | 1.90 |
| --- | --- | --- | --- | --- |
| 3 | A2 | MBAR1141 | PT 12 Warm Up | 54.25 |
| 9 | B1 | SSS135 | PT 14 Magnet MAPS | 296.34 |
| 4 | B2 | SSS277 | PT 6.4 Training 4 | 4.16 |
| 8 | C1 | MBBL3124 | Prep 1 Lyre Test | 295.18 |
| 1 | C2 | MBAR3301 | PT 12 Warm Up | 287.94 |
| 10 | D1 |  |  | 296.34 |
| 11 | D2 | SSS613 | PT 14 Magnet MAPS | 296.88 |
| 12 | E1 | MBBR3356 | ICS 6 Disconnect Magnet WP07 | 299.56 |
| 7 | E2 | MBBL3096 | ICS 3 Traveller after warm test | 297.68 |
| 6 | F1 | MBAL1229 | Prep 5 Cool Down | 97 |
| 2 | F2 | MBBL3134 | PT 12 Warm Up | 176.91 |

Coordinator this week: Nikolay; MTM support on call :   16-3602

17  complete Magnet Tests cycles   ( 3147  ROBTC )

12  dipoles( 2 with MM )+ 5  SSS( 1  with MM) +  0 SSSS (with MM)

Disconnected or ROBTC after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- Situation on Repair :   3 silver shoes done , 9 to go ?

- General Points

AOB

## Tests

| Magnets under Test from 29-05-2005 to 04-06-2005 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBAR3298 |  |  | 30-05-2005 | 31 h. | 21 h. | 22 h. | 15 h. | 13 h. | 04-06-2005 | 102 h. | Stripping |  |  |
| 2 | TBA1 | MBAR3319 |  |  | 26-05-2005 | 26 h. | 24 h. | 39 h. | 15 h. | 9 h. | 30-05-2005 | 113 h. | Stripping |  |  |
| 3 | TBA2 | MBAR3344 |  |  | 27-05-2005 | 26 h. | 36 h. | 75 h. | 15 h. | 10 h. | 03-06-2005 | 162 h. | Stripping |  |  |
| 4 | TBB1 | SSS178 |  |  | 27-05-2005 | 39 h. | 13 h. | 14 h. | 9 h. | 8 h. | 30-05-2005 | 83 h. | Stripping |  |  |
| 5 | TBB1 | SSS194 |  |  | 31-05-2005 | 31 h. | 10 h. | 18 h. | 7 h. | 6 h. | 02-06-2005 | 72 h. | Stripping |  |  |
| 6 | TBB2 | SSS172 |  |  | 31-05-2005 | 26 h. | 14 h. | 20 h. | 9 h. | 12 h. | 04-06-2005 | 81 h. | Stripping |  |  |
| 7 | TBB2 | SSS197 |  |  | 27-05-2005 | 24 h. | 23 h. | 16 h. | 11 h. | 13 h. | 31-05-2005 | 87 h. | Stripping |  |  |
| 8 | TBC1 | MBBL2179 |  |  | 28-05-2005 | 17 h. | 32 h. | 20 h. | 14 h. | 7 h. | 01-06-2005 | 90 h. | Stand By |  |  |
| 9 | TBC2 | MBBR3337 |  |  | 26-05-2005 | 25 h. | 21 h. | 21 h. | 13 h. | 9 h. | 30-05-2005 | 89 h. | Stripping |  |  |
| 10 | TBC2 | MBBR3338 |  |  | 30-05-2005 | 21 h. | 25 h. | 16 h. | 14 h. | 13 h. | 03-06-2005 | 89 h. | Stripping |  |  |
| 11 | TBD1 | SSS156 |  |  | 20-05-2005 | 42 h. | 11 h. | 160 h. | 39 h. | 8 h. | 31-05-2005 | 260 h. | Stripping |  |  |
| 12 | TBE1 | MBAR3346 |  |  | 29-05-2005 | 23 h. | 25 h. | 15 h. | 14 h. | 7 h. | 02-06-2005 | 84 h. | Stripping |  |  |
| 13 | TBE1 | MBBL2185 |  |  | 23-05-2005 | 29 h. | 35 h. | 50 h. | 16 h. | 10 h. | 29-05-2005 | 140 h. | Stripping |  |  |
| 14 | TBE2 | MBAR3327 |  |  | 28-05-2005 | 23 h. | 25 h. | 22 h. | 14 h. | 11 h. | 01-06-2005 | 95 h. | Stripping |  |  |
| 15 | TBF1 | MBAL3147 |  |  | 28-05-2005 | 0 h. | 25 h. | 11 h. | 14 h. | 13 h. | 31-05-2005 | 63 h. | Stripping |  |  |
| 16 | TBF1 | MBAR3351 |  |  | 31-05-2005 | 31 h. | 27 h. | 18 h. | 13 h. | 11 h. | 04-06-2005 | 100 h. | Stripping |  |  |
| 17 | TBF2 | MBAR3175 |  |  | 24-05-2005 | 36 h. | 43 h. | 27 h. | 16 h. | 9 h. | 29-05-2005 | 131 h. | Stripping |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAR3298 | 11568.3 A. (D2_U) | 12017 A. (D1_U) | 12378 A. (D1-L) |  | - | - | - | - | - | - | - | - | - |
| 2 | MBAR3319 | 9338 A. (D2-L-S01) | 11245 A. (D1-U-S11) | 0 A. (No Data) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 3 | MBAR3344 | 12850 A. | - | - | - | - | - | - | - | - | - | - | - | - |
| 4 | SSS178 | 12613 A. (QI _34) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 5 | SSS194 | 12239 A. (QE-12) | 12340 A. (QE-34) | 12462 A. (QE-34) |  | - | - | - | - | - | - | - | - | - |
| 6 | SSS172 | 10511 A. (QE_12) | 11777 A. (QE_12) | 12280 A. (QE_12) | 12776 A. (QI-34) |  | - | - | - | - | - | - | - | - |
| 7 | SSS197 | 11932 A. (QI-34) | 12109 A. (QI-34) | 12850 A. (QI-34) |  | - | - | - | - | - | - | - | - | - |
| 8 | MBBL2179 | 11421 A. (D1 - U) | 12470 A. (D2-U) | 12581 A. (D2-L) |  | - | - | - | - | - | - | - | - | - |
| 9 | MBBR3337 | 10455 A. (D1-U) | 12849 A. (No Data) | 11998.6 A. (D1-U) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 10 | MBBR3338 | 12180 A. (D2 - U) | 12711 A. (D1-L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 11 | SSS156 | 11916.5 A. (QI-12) | 12224.6 A. (QI-12) |  | - | - | - | - | - | - | - | - | - | - |
| 12 | MBAR3346 | 12210 A. (D2-L) | 12236 A. (D1-U) | 12818 A. (D2-U) |  | - | - | - | - | - | - | - | - | - |
| 13 | MBBL2185 | 9550 A. (D1-L) | 11487 A. (D1-L) | 12220 A. (D2-L) | 12429 A. (D1-L) | 12615 A. (D1-L) | 12645 A. (D1-L) | 12750 A. (D1-L) | 12840 A. (D1-U) | 12900 A. | - | - | - | - |
| 14 | MBAR3327 | 9935.7 A. (D2-L) | 12900 A. | - | - | - | - | - | - | - | - | - | - | - |
| 15 | MBAL3147 | 12616 A. (D2-U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 16 | MBAR3351 | 12461.3 A. (D2_L) | 12686 A. (D1-L) | 12850 A. (D2-U) |  | - | - | - | - | - | - | - | - | - |
| 17 | MBAR3175 | 12457 A. (D2-U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |

Back to Homepage | For print

### Week 23  :  29 May to 04 June 05  (17 Magnet Tests)

### [3147 had ROBTC ]

### END May05 STATISTICS & PERFORMANCE CURVES

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _31_A | TBD2 | SPA on RMD when going to stand by mode | AB/PO |  | Memory fault cleared by Federico. Reset of the PC14 kA before launching the test. |
| Tuesday _31_N | TBE1 | Offset on Potaim card Vsupra_ext not stable + VClamp CD triggering | MTM |  | Federico solved many problems: 1 potaim card is out of order, several Vtaps missing at the input of potaim crate nr.2 i.e. clamp VT1 to VT4 |
| Thursday_02_A | TBA2 | T_sensor in the MB3344 is not showing the good slope although there is a perfect straight line. | MTM |  | Federico called. He explained that the calibration coefficients for the interpolation were changed for the TT821 and TT148 both for in AQA and IAP |
| Thursday_02_A | TBD1 | AC compensation on SSS202 could not be carried out.. Trigger is on Potaim cards Vbus_Ext & D1-D2_lg. | MTM |  | Solved using the adaptation. |
| Friday_03_N | TBE2 | High current indicator on the CFB is not flashing | MTM |  |  |
| Saturday_04_M | TBD1 | During the generation of the “ub file” for PT8, DAQ received a trigger “ read out starts”, system switched off. Same message received when the system is switched off. | AB/CO? |  | E-mail sent to Giorgio. |
| Saturday_04_M | TBD2 | Polarity check on SSS613. Voltage applied of. 10 V, current indicated less than 1 A. | MTM |  | Giorgio called. He advised to apply the max Voltage of 40 V and to see if the current increased proportionally. He suggested to do the test as for the arc SSS. In any case there is no diode for this SSS. |
| Saturday_04_M | TBC1 | HV test after cold test using HV rack XMT MY92 on MB3295.Wrong voltage applied in the first step: 700V instead of 600 V. Test was stopped with a comment "Applied voltage exceeded by more than 10% then specified" Run recovery didn't worked. Also it was found that the C11 connector was damaged for this | MTM |  | Rack was changed and test proceed smoothly? Test failed because cryo-heater connector was not working. Giorgio advise to go on with the test, and to open NC to have the connector repaired later on. |
| Saturday_04_N | TBB1 | PLCSE blocked. | MTM |  | Switched off/on the PLCSE. |
| Saturday_04_N | TBB1 | De-excitation test PT8.2 for SSS. The PC did not follow the 350A/s ramp rate specified by the test. | MTM |  | The de-ramp happens at ~120A/s. |
| Sunday_05_M | All | MTF not working since mid night? | IT? |  | Restart to work at 8..30 Am |
| Sunday_05_M | TBF2 | MM1 on MB3134. Test cannot be launched. Magnets setting file had no Temperature data for flat measurements | MTM | MMP | Ramp rate was put at 0A/s. Changed with 10A/s and test can be launched.. Giorgio corrected it, then TT821 is OK |
| Sunday_05_A | All | Quench during MM2 (LHC cycle) for MB3134. | ACR? |  | Giorgio contacted. Possible explanation : too long time at 11850 A: Tleads B too high and SPA occurred quench during the de-ramp. |
| Monday_06_M | TBE2 | Error in Prep 2.2 (QH discharge HF); tminV for H4 is 1624. | IT? |  | Maryline checked and advised to continue. |