Vc/13-Jun-05

### Week 24  starting 13 June 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  13 June 05 @: 13/06/2005 14:33:29

| 3 | A1 | MBAR3305 | PT 13.2 Resist. Meas. | 300.82 |
| --- | --- | --- | --- | --- |
| 10 | A2 |  |  | 298.04 |
| 2 | B1 | SSS285 | ICS 6 Disconnect Magnet WP07 | 298.58 |
| 12 | B2 |  |  | 298.04 |
| 11 | C1 | MBAL1224 | Prep 2.2 IAP @ warm HF | 296.88 |
| 4 | C2 | MBBL3100 | Prep 5 Cool Down | 91.24 |
| 6 | D1 | SSS255 | Prep 5 Cool Down | 2.35 |
| 5 | D2 | SSS639 | PTE 3 Correctors Powering | 1.90 |
| 7 | E1 | MBBL1231 | Temporary test and preparation | 135.96 |
| 8 | E2 | MBBL3125 | CR 1 Leak Test | 298.04 |
| 9 | F1 | MBAR3347 | Prep 5 Cool Down | 2.35 |
| 1 | F2 | MBAL1225 | PT 10 Min. NRJ 11850 A | 1.92 |

Temperature data are taken from TBSM18 site

Coordinator this week: V.Chohan; MTM support on call :   16-3602

19  complete Magnet Tests cycles   ( S277  ROBTC )

12  dipoles( 1 with MM )+ 6  SSS( 0  with MM) +  1 SSSS (with MM)

[3093 & 3096 put in standby because there was no diode]

Disconnected or ROBTC after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

- General Points

  - F1 & F2

  - SSS639 (  Salem)

AOB

| Magnets under Test from 05-06-2005 to 11-06-2005 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBAL3101 |  |  | 04-06-2005 | 28 h. | 21 h. | 23 h. | 14 h. | 13 h. | 08-06-2005 | 99 h. | Stripping |  |  |
| 2 | TBA2 | MBAR1141 |  |  | 03-06-2005 | 34 h. | 23 h. | 16 h. | 19 h. | 8 h. | 08-06-2005 | 100 h. | Stripping |  |  |
| 3 | TBB1 | SSS135 |  |  | 03-06-2005 | 28 h. | 14 h. | 15 h. | 6 h. | 23 h. | 06-06-2005 | 86 h. | Stripping |  |  |
| 4 | TBB1 | SSS201 |  |  | 06-06-2005 | 25 h. | 15 h. | 24 h. | 9 h. | 14 h. | 10-06-2005 | 87 h. | Stripping |  |  |
| 5 | TBB2 | SSS277 |  |  | 04-06-2005 | 32 h. | 12 h. | 26 h. | 6 h. | 0 h. | 07-06-2005 | 76 h. | Stand By |  |  |
| 6 | TBB2 | SSS277 |  |  | 07-06-2005 | 0 h. | 12 h. | 7 h. | 10 h. | 22 h. | 09-06-2005 | 51 h. | Stand By |  |  |
| 7 | TBC1 | MBAR3295 |  |  | 01-06-2005 | 22 h. | 29 h. | 19 h. | 13 h. | 7 h. | 06-06-2005 | 90 h. | Stripping |  |  |
| 8 | TBC1 | MBBL3124 |  |  | 06-06-2005 | 30 h. | 23 h. | 13 h. | 13 h. | 4 h. | 09-06-2005 | 83 h. | Stripping |  |  |
| 9 | TBC2 | MBAR3301 |  |  | 03-06-2005 | 20 h. | 20 h. | 18 h. | 14 h. | 16 h. | 07-06-2005 | 88 h. | Stripping |  |  |
| 10 | TBC2 | MBBL3093 |  |  | 07-06-2005 | 24 h. | 21 h. | 19 h. | 13 h. | 13 h. | 11-06-2005 | 90 h. | Stand By |  |  |
| 11 | TBD1 | SSS202 |  |  | 31-05-2005 | 32 h. | 23 h. | 29 h. | 7 h. | 9 h. | 06-06-2005 | 100 h. | Stripping |  |  |
| 12 | TBD1 | SSS210 |  |  | 07-06-2005 | 29 h. | 11 h. | 26 h. | 8 h. | 17 h. | 11-06-2005 | 91 h. | Stripping |  |  |
| 13 | TBD2 | SSS613 |  |  | 27-05-2005 | 28 h. | 36 h. | 109 h. | 12 h. | 51 h. | 06-06-2005 | 236 h. | Stand By |  |  |
| 14 | TBE1 | MBAL2158 |  |  | 06-06-2005 | 32 h. | 21 h. | 30 h. | 15 h. | 14 h. | 11-06-2005 | 112 h. | Stripping |  |  |
| 15 | TBE1 | MBBR3356 |  |  | 02-06-2005 | 24 h. | 28 h. | 20 h. | 16 h. | 10 h. | 06-06-2005 | 98 h. | Stripping |  |  |
| 16 | TBE2 | MBAR3345 |  |  | 01-06-2005 | 23 h. | 26 h. | 30 h. | 16 h. | 9 h. | 05-06-2005 | 104 h. | Stripping |  |  |
| 17 | TBE2 | MBBL3096 |  |  | 05-06-2005 | 26 h. | 34 h. | 18 h. | 14 h. | 15 h. | 10-06-2005 | 107 h. | Stand By |  |  |
| 18 | TBF1 | MBAL1229 |  |  | 04-06-2005 | 26 h. | 22 h. | 19 h. | 14 h. | 10 h. | 08-06-2005 | 91 h. | Stripping |  |  |
| 19 | TBF2 | MBBL3134 |  |  | 02-06-2005 | 29 h. | 27 h. | 29 h. | 14 h. | 15 h. | 07-06-2005 | 114 h. | Stripping |  |  |

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAL3101 | 9028.9 A. (D1_L) | 11615 A. (D1_U) | 12500 A. (D1-U) | 12693.8 A. (D2-L) |  | - | - | - | - | - | - | - | - |
| 2 | MBAR1141 | 11955 A. (D1_L) | 12308 A. (D2_L) | 12679.1 A. (D1-L) |  | - | - | - | - | - | - | - | - | - |
| 3 | SSS135 | 12632 A. (QE_34) | 12652.6 A. (QE_12) | 12850 A. (QI-12) |  | - | - | - | - | - | - | - | - | - |
| 4 | SSS201 | 11986 A. (QE-12) | 12028 A. (QE-12) | 12219 A. (QI-34) |  | - | - | - | - | - | - | - | - | - |
| 5 | SSS277 | 10859 A. (QI-34) | 11011 A. (QI-34) | 11046 A. (QI-34) | 11206 A. (QI-34) | 11349 A. (QI-34) | 11076 A. (QI-34) | 11215 A. (QI_34) | 11177 A. (QI-34) | 11159 A. (QI-34) |  | - | - | - |
| 6 | SSS277 | 10769 A. (QI-34) | 10650 A. (QI-34) |  | - | - | - | - | - | - | - | - | - | - |
| 7 | MBAR3295 | 12166.5 A. (D2-U) | 12713 A. (D1 - L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 8 | MBBL3124 | 12174 A. (D2_U) | 12480 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 9 | MBAR3301 | 11706 A. (D2-U) | 12712 A. (D2-L) | 12591 A. (D2-L) |  | - | - | - | - | - | - | - | - | - |
| 10 | MBBL3093 | 9380.4 A. (D2-L) | 10758 A. (D1 - U) | 12328 A. (D1 - L) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 11 | SSS202 | 12143 A. (QE_12) | 12469 A. (QE_34) | 12667 A. (QE-12) |  | - | - | - | - | - | - | - | - | - |
| 12 | SSS210 | 11871 A. (QE-12) | 12193 A. (QE-12) | 12538 A. (QI-34) |  | - | - | - | - | - | - | - | - | - |
| 13 | SSS613 | 5820 A. | - | - | - | - | - | - | - | - | - | - | - | - |
| 14 | MBAL2158 | 10229.1 A. (D1-U) | 11274.5 A. (D2-U) | 12446 A. (D1 - L) | 12619 A. (D2 - L) |  | - | - | - | - | - | - | - | - |
| 15 | MBBR3356 | 9232 A. (D1_U) | 11864 A. (D1 - U) | 12850 A. (D1 - L) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 16 | MBAR3345 | 9633 A. (D2_U) | 11213.9 A. (D2-U) | 11450.7 A. (D2-U) | 12001.7 A. (D2-U) | 12307 A. (D2 - U) | 12652 A. (D2 - U) | 12759 A. (D2-U) | 12900 A. | - | - | - | - | - |
| 17 | MBBL3096 | 9611.2 A. (D2_U) | 12641.2 A. (D1 - L) | 12857 A. | - | - | - | - | - | - | - | - | - | - |
| 18 | MBAL1229 | 12482 A. (D1-L) | 12761.1 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 19 | MBBL3134 | 11213.6 A. (D1-L) | 12032.8 A. (D2-U-S01) | 12218.2 A. (D2-U/S01) |  | - | - | - | - | - | - | - | - | - |

Back to Homepage | For print

### Week 24  :  5 June to 11 June 05  (19 Magnet Tests)

### [ S277 had  ROBTC ]

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _31_A | TBD2 | SPA on RMD when going to stand by mode | AB/PO |  | Memory fault cleared by Federico. Reset of the PC14 kA before launching the test. |
| Tuesday _31_N | TBE1 | Offset on Potaim card Vsupra_ext not stable + VClamp CD triggering | MTM |  | Federico solved many problems: 1 potaim card is out of order, several Vtaps missing at the input of potaim crate nr.2 i.e. clamp VT1 to VT4 |
| Thursday_02_A | TBA2 | T_sensor in the MB3344 is not showing the good slope although there is a perfect straight line. | MTM |  | Federico called. He explained that the calibration coefficients for the interpolation were changed for the TT821 and TT148 both for in AQA and IAP |
| Thursday_02_A | TBD1 | AC compensation on SSS202 could not be carried out.. Trigger is on Potaim cards Vbus_Ext & D1-D2_lg. | MTM |  | Solved using the adaptation. |
| Friday_03_N | TBE2 | High current indicator on the CFB is not flashing | MTM |  |  |
| Saturday_04_M | TBD1 | During the generation of the “ub file” for PT8, DAQ received a trigger “ read out starts”, system switched off. Same message received when the system is switched off. | AB/CO? |  | E-mail sent to Giorgio. |
| Saturday_04_M | TBD2 | Polarity check on SSS613. Voltage applied of. 10 V, current indicated less than 1 A. | MTM |  | Giorgio called. He advised to apply the max Voltage of 40 V and to see if the current increased proportionally. He suggested to do the test as for the arc SSS. In any case there is no diode for this SSS. |
| Saturday_04_M | TBC1 | HV test after cold test using HV rack XMT MY92 on MB3295.Wrong voltage applied in the first step: 700V instead of 600 V. Test was stopped with a comment "Applied voltage exceeded by more than 10% then specified" Run recovery didn't worked. Also it was found that the C11 connector was damaged for this | MTM |  | Rack was changed and test proceed smoothly? Test failed because cryo-heater connector was not working. Giorgio advise to go on with the test, and to open NC to have the connector repaired later on. |
| Saturday_04_N | TBB1 | PLCSE blocked. | MTM |  | Switched off/on the PLCSE. |
| Saturday_04_N | TBB1 | De-excitation test PT8.2 for SSS. The PC did not follow the 350A/s ramp rate specified by the test. | MTM |  | The de-ramp happens at ~120A/s. |
| Sunday_05_M | All | MTF not working since mid night? | IT? |  | Restart to work at 8..30 Am |
| Sunday_05_M | TBF2 | MM1 on MB3134. Test cannot be launched. Magnets setting file had no Temperature data for flat measurements | MTM | MMP | Ramp rate was put at 0A/s. Changed with 10A/s and test can be launched.. Giorgio corrected it, then TT821 is OK |
| Sunday_05_A | All | Quench during MM2 (LHC cycle) for MB3134. | ACR? |  | Giorgio contacted. Possible explanation : too long time at 11850 A: Tleads B too high and SPA occurred quench during the de-ramp. |
| Monday_06_M | TBE2 | Error in Prep 2.2 (QH discharge HF); tminV for H4 is 1624. | IT? |  | Maryline checked and advised to continue. |