vc/27-Jun-05

### Week 26  starting 27 June 05

### SM18 Operation Meeting (Tuesdays)

Bench status  on Monday  27 June 05 @: 6/27/2005 11:18:06 AM

| 10 | A1 | MBBL1232 | PT 14 Magnet MAPS | 300.82 |
| --- | --- | --- | --- | --- |
| 7 | A2 | MBAL2193 | ICS 4 Final connection | 300.1 |
| 8 | B1 | SSS121 | Prep 2.2 IAP @ warm HF | 300.46 |
| 3 | B2 | SSS290 | PT 13.2 Resist. Meas. | 301.89 |
| 9 | C1 | MBBL1234 | PT 13.2 Resist. Meas. | 300.28 |
| 4 | C2 | MBBR3357 | Prep 5 Cool Down | 5.86 |
| 1 | D1 | SSS280 | Thermal Cycle | 5 HV |
| 12 | D2 | SSS616 | EX1 Extra Test | 299.83 |
| 11 | E1 | MBBL2165 | Finished Ready to go Final scan | 301.35 |
| 2 | E2 | MBBR3355 | PT 13.2 Resist. Meas. | 303.05 |
| 5 | F1 | MBAL1228 | PT 12 Warm Up | 147.42 |
| 6 | F2 | MBAL3330 | Prep 5 Cool Down | 288.03 |

Temperature data are taken from TBSM18 site

Coordinator this week: P.Pugnat;

MTM support on call :  16-3602

12 complete Magnet Tests cycles   ( 2191  ROBTC )

09  dipoles( 1 with MM )+ 3  SSS( 0  with MM) +  0 SSSS (with MM)

[2524 brought again on another cluster but cables too short – now finally rejected]

Disconnected or ROBTC after tests last week ( Sat Midnight to Sat midnight)

- Situation From Cryogenics:  J-P Lamboy or ACR Rep.

Lots of problems in Cryo last week?

- General Points

AOB

| Magnets under Test from 19-06-2005 to 25-06-2005 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBAR3343 |  |  | 14-06-2005 | 30 h. | 54 h. | 33 h. | 15 h. | 11 h. | 20-06-2005 | 143 h. | Stripping |  |  |
| 2 | TBA2 | MBAR2524 |  |  | 13-06-2005 | 39 h. | 24 h. | 39 h. | 20 h. | 13 h. | 19-06-2005 | 135 h. | Stripping |  |  |
| 3 | TBB1 | SSS223 |  |  | 17-06-2005 | 41 h. | 18 h. | 16 h. | 8 h. | 18 h. | 22-06-2005 | 101 h. | Stripping |  |  |
| 4 | TBB2 | SSS219 |  |  | 17-06-2005 | 41 h. | 36 h. | 19 h. | 9 h. | 16 h. | 22-06-2005 | 121 h. | Stripping |  |  |
| 5 | TBC1 | MBAR3252 |  |  | 17-06-2005 | 30 h. | 42 h. | 41 h. | 21 h. | 20 h. | 24-06-2005 | 154 h. | Stand By |  |  |
| 6 | TBC2 | MBAL2191 |  |  | 15-06-2005 | 35 h. | 27 h. | 70 h. | 15 h. | 0 h. | 21-06-2005 | 147 h. | Not Found or Thermal cycle |  |  |
| 7 | TBC2 | MBAL2191 |  |  | 22-06-2005 | 0 h. | 27 h. | 23 h. | 19 h. | 10 h. | 25-06-2005 | 79 h. | Stripping |  |  |
| 8 | TBD1 | SSS200 |  |  | 16-06-2005 | 30 h. | 18 h. | 20 h. | 6 h. | 15 h. | 20-06-2005 | 89 h. | Stripping |  |  |
| 9 | TBE1 | MBAL3103 |  |  | 16-06-2005 | 43 h. | 37 h. | 16 h. | 14 h. | 24 h. | 22-06-2005 | 134 h. | Stripping |  |  |
| 10 | TBE2 | MBBR3362 |  |  | 16-06-2005 | 45 h. | 50 h. | 15 h. | 20 h. | 23 h. | 22-06-2005 | 153 h. | Stripping |  |  |
| 11 | TBF1 | MBBR3361 |  |  | 16-06-2005 | 33 h. | 23 h. | 36 h. | 17 h. | 25 h. | 21-06-2005 | 134 h. | Stripping |  |  |
| 12 | TBF2 | MBAL2196 |  |  | 18-06-2005 | 38 h. | 34 h. | 64 h. | 17 h. | 12 h. | 25-06-2005 | 165 h. | Stripping |  |  |
| 13 | TBF2 | MBAR2524 |  |  | 25-06-2005 | 1 h. | 0 h. | 0 h. | 0 h. | 0 h. | 25-06-2005 | 1 h. | Not Found or Thermal cycle |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAR3343 | 10258.5 A. (D2 - L) | 11012.9 A. (D1_U) | 12628 A. (D2-L) | 12720 A. (D1-U) |  | - | - | - | - | - | - | - | - |
| 2 | MBAR2524 | 10833 A. (D2_L) | 11852 A. (D1-L) | 12158 A. (D1-L) | 12338 A. (D2-L) | 12411.4 A. (D2-U) | 12569 A. (D-L) | 12759 A. (D2-U) | 12827 A. (D2-U) | 12850 A. | - | - | - | - |
| 3 | SSS223 | 12367 A. (QE-34) | 12851 A. (QI_12) |  | - | - | - | - | - | - | - | - | - | - |
| 4 | SSS219 | 11592 A. (QE-34) | 12708 A. (QI-12) | 12818.1 A. (QI_34) |  | - | - | - | - | - | - | - | - | - |
| 5 | MBAR3252 | 11573.6 A. (D1-L) | 11775 A. (D1-L) | 12077 A. (D1_L) | 12332.6 A. (D1-L) | 12609.6 A. (D1-L) | 12773 A. (D1 - L) | 12850 A. | - | - | - | - | - | - |
| 6 | MBAL2191 | 11335 A. (D2-L) | 11720 A. (D1-U) | 12148.1 A. (D2-L) | 12398 A. (D2-L) | 12620 A. (D2-L) | 12720 A. (D1-U) | 12410 A. (D1-U) | 12740 A. (D1-U) | 12848.5 A. (D1-L) |  | - | - | - |
| 7 | MBAL2191 | 11982 A. (D2-U) | 12475 A. (D2-L) |  | - | - | - | - | - | - | - | - | - | - |
| 8 | SSS200 | 11731.6 A. (QI_12) | 12497.9 A. (QI_34) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 9 | MBAL3103 | 12469.2 A. (D1_U) | 12900 A. | - | - | - | - | - | - | - | - | - | - | - |
| 10 | MBBR3362 | 10149 A. (D1 - L) | 12352 A. (D1 - L) | 12650 A. (D1_L) |  | - | - | - | - | - | - | - | - | - |
| 11 | MBBR3361 | 12551.9 A. (D1_L) | 12449 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 12 | MBAL2196 | 11088 A. (D1_U) | 12302 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 13 | MBAR2524 |  | - | - | - | - | - | - | - | - | - | - | - | - |

According SMTMS and e-traveller the time lost has been :

A2 (New MM2 application: who, for what & who sanctioned it ?) :1235   23/06/05     from : 8:44    till : 20:51   equals : 12:07 hrs

D1 (Because of special SSS cohabitation & not able to switch Electronics & non-24 hr working for SSSS ):

SSS280  22/06/05   from : 8:23  till : 20:45   equals : 12:22 hrs

F2 (LEICA field measurements ( not only parasitic ) , New Mini NRJ application, New MM2 application  - why this work ? ..) :

2196  21/06/05  from : 8:55   till : 17:51   equals : 8:56 hrs   

 So Adding all above .  giving   Total at least  33h for special tests on different clusters but finally leading to de-phasing in cold /warm   requirements for optimal throughput.

E Cluster PROBLEMS

This explains  in part the poor week (13 magnets), but there were also hardware problems as well as Cryo problems :  see elogbook.

So ICS oversight on C is only a minor detail added to all above.
Top of Form

[SM18 OP] 25/06/2005 22:27:00

TBE1 : No data in HFDAQ since HF aquisition did not trigger. Richard was informed. He suggested to use lfoff data for finding quench location and proceed for next stage.

[SM18 OP] 22/06/2005 14:42:00

TBE1 : HF acqusition system had problem in arming , crates were reseted but the problem came again message was :Crate 1:slot1 MX in error. HF crate being changed.

Top of Form

[SM18 OP] 22/06/2005 04:32:00

TBE2: MBBR3362: We had to switch off HF and LF crates twice before starting again the process to launch PT13.1

[SM18 OP] 22/06/2005 03:31:00

TBE2:MBBR3362: HF status: AM Arm Error 2 Impossible to launch iap @ warm (PT13.1).

Top of Form

C Cluster PROBLEMS

Bottom of Form

[SM18 OP] 23/06/2005 12:16:43

TBC2: There was warning during first training ,TT821 is showing 1.548 ,Richa was informed. Richa checked and advised to open non-confirmity for temprature sensor

[SM18 OP] 23/06/2005 10:09:16

TBC2 : Bad connection problem solved by Maryiline

[SM18 OP] 23/06/2005 00:07:15

TBC2:we found a "quench origin out of the main coils" (D1-D2 = 0 volts) we checked the bench and around the voltage taps A and B there is some water. We informed cryo and they'll have a look.

[SM18 OP] 22/06/2005 18:55:00

TBC2: CFB now dryer, proceeding with HV tests

[SM18 OP] 22/06/2005 17:21:50

TBC2: Ice forming on CFB, possible cause of HV problem, PT2. Waiting for cryo team to clean up.

[SM18 OP] 21/06/2005 01:24:40

TBC2: During second SSL Quench@4 K, SPA occurred at around 9500A due to cryo loss. Test launched again after cryo OK.

[SM18 OP] 20/06/2005 20:51:50

TBC2: HV final@4.4K. Failed Dipoles vs Ground at 3050V due to condensation on the current leads. Richard checked and asked us to skip that step and continue.

D Cluster PROBLEMS

This explains  in part the poor week (13 magnets), but there were also hardware problems as well as Cryo problems :  see elogbook.

So ICS oversight on C is only a minor detail added to all above.
Top of Form

C Cluster PROBLEMS

C Cluster PROBLEMS C Cluster PROBLEMS Top of Form

Bottom of Form

[SM18 OP] 26/06/2005 11:36:00

TBD1 : "Contacting PLC_SE..." after HV @ cold PT9, the PLC SE was not contacting by server. After restart server, PLC no amelioration, after restart the sun communication become OK.

[SM18 OP] 24/06/2005 04:00:00

TBD1:Cryo OK now, but all cryo pumping is slow.

[SM18 OP] 24/06/2005 02:56:05

TBD1: Cryo not OK not able to proceed with powering.

[SM18 OP] 23/06/2005 23:50:00

TBD1: Slow PA test. The PC would not go to STANDBY. We found that the Power Supply panel showed Dipole instead of Arc SSS. We changed the inductance to SSS manually, and the PC worked. We then restarted TEMA and re-ran the test, and the Power application configured correctly. The problem reappeared and the test was launched from local server bypassing TEMA

[SM18 OP] 23/06/2005 22:00:00

TBD1: AIP@cold. Problems adjusting VSum. We found that the cables on channels 10 and 11 were inversed.

[SM18 OP] 20/06/2005 13:10:00

TBD1: SSS200 is a new type (MS) and the the values are OK

Top of Form

Bottom of Form

[SM18 OP] 23/06/2005 00:33:35

TBD2: The nominal strength m/m was repeated but the required files are not created. The test was repeated twice, but still the same problem

[SM18 OP] 23/06/2005 00:33:05

TBD2: SPA happened between the measurement due to cryo interlock. Again repeated the test

[SM18 OP] 23/06/2005 00:07:45

TBD2:FS12 nominal strength measurement problem. The measurement was contnuing in a loop with the message "checking the motion status.." . The measurement was stopped and repeated.

[SM18 OP] 22/06/2005 19:12:24

TBD2 : Slow PA because of spurious trigger on power converter

[SM18 OP] 22/06/2005 18:28:27

TBD2 : Slow PA because of cryo OK lost. the test has ok by chance.

[SM18 OP] 22/06/2005 03:43:42

TBD2: Return wire was not connected to the input of the Metrolab card. Previous measurement must be checked, by expert, tomorrow morning, as we might need to repeat the survey if it is not good. Richard.

[SM18 OP] 22/06/2005 01:51:29

TBD2: problem with FS22 -no data valid so we contact Richard who advise to check the wire connection and if the wire is touching the anticryostat .

F Cluster PROBLEMS This explains in part the poor week (13 magnets), but there were also hardware problems as well as Cryo problems : see elogbook. So ICS oversight on C is only a minor detail added to all above. Top of Form

Top of Form

[SM18 OP] 23/06/2005 10:09:01

TBF2 : PT10 was repeated after consulting with stephan and Pierre, Same alarm occured --- D1-2:Direct voltage peak exceeded 200V

[SM18 OP] 22/06/2005 07:10:00

TBF2: MM2 works now

[SM18 OP] 22/06/2005 06:50:00

TBF2: problems in launching MM2; relaunched MM2 after VME crate reset

[SM18 OP] 21/06/2005 07:03:09

TBF2: communications problem with PLC, reset

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday _21_A | TBB2 | After HV test at cold, error on the PLC: try to connect PLC | AB/CO |  | Reset of the PLC |
| Tuesday _21_N | TBD2 | SSW measurements on SSS116.Values out of ranges. Repeated : same results | MTM |  | Richard came and found that the return cable was not connected to the input of the Metrolab card. Test repeated. Stephane and Nikolay checked the results : OK |
| Tuesday _21_N | TBE2 | Arm Error 2 Impossible to launch iap @ warm. | AB/CO | hf | Switch off hf and lf crates twice before starting again IAP. Adrian replaced the hf crate Wednesday. |
| Wednesday _22_M | TBF2 | LHC cycle cannot be launched: VME communication error | AB/CO |  | Reset the VME crate and launch the test. |
| Wednesday _22_A | TBA2 | MM1: encoder signal not delivered properly during rotation in aperture 2 | MTM |  | Solved by F.Flamand |
| Wednesday _22_A | TBC2 | Ice appearing on the current leads on the CFB: HV test may be not performed in correct conditions. | MTM |  | Test performed after making the CFB current leads dry. |
| Wednesday _22_A | TBA2 | Problem to arm hf DAQ. Unable to perform High Field quench at 1.5 kA: MXI connector of the crate hf was not properly screwed. 2 erronous triggers on LF/HF acquisition, and same error message, unable to re-arm hf. | MTM | hf | Reset crate twice, then ok. Richard screwed the MXI connector. MXI card replaced by Richard. |
| Wednesday _22_N | TBC2 | Hf quench at 1.5 kA: "quench origin out of the main coils" (D1-D2 = 0 volts) | MTM |  | Richard and Stéphane informed. Stéphane asked a description of the signals observed and ask to stop the test till the morning. Lf signals are OK so not a card problem. Maryline and Richard found that a splitter for the hf crate was damaged. |
| Wednesday _22_N | TBC2 | Min energy at 11850 A on MB2196 . Alarm DTV too high (~280 V) | MTM |  | Stephane asked to stop the warm up. He checked the signals: consistency is there. He contacted Pierre who made a special investigation: Min Energy with different LF heater to provoke the quench. Special HV test. DVT lower after provoking the quench with a lower energy (380 V instead of 500 V) |
| Wednesday _22_N | TBD2 | Value of the strength at nominal not running correctly. The measurement was continuing in a loop with the message "checking the motion status." | MTM |  | Problem in the motor controller.Test stopped and repeated but the required files are not created. The test was repeated twice, but still the same problem. To repeate the test in good conditions the motor controler have to be switched off and on. |
| Wednesday _22_N | TBD2 | SPA happened between the measurements due to cryo interlock. | ACR |  |  |
| Thursday _23_M | TBC2 | Tsensor on MB2191 not showing the correct value: 1.55 K instead of 1.9 K. Warning during the training. | MTM |  | Richard advised to open a non conformity. |
| Thursday _23_A | TBF2 | HF arming error while launching training. | AB/CO | hf | Quit LF & HF applications, switched off crate power supply. |
| Thursday _23_N | TBD1 | AIP@cold. Problems adjusting VSum. | MTM |  | The cables on channels 10 and 11 were inverted. |
| Thursday _23_N | TBD1 | SPA test cannot be launched. RMD not going to stamd_by. Inductance displayed is 90 mH not that for an SSS. | MTM |  | Good value of inductance entered. Test ran form local sun. |
| Thursday _23_N | TBA2 | SPA during LHC cycle. | MTM |  | Good value of inductance entered. Test ran form local sun. |
| Thursday _23_N | TBA2 | Micro cut of power: One automate blocked. No alarm to ACR. Other pumps overloaded. Cryo pumping weakened. |  |  | No pumping. Cryo OK lost . Magnet going to 4 K. Loadline in MB1235 aborted. LHC cycle on MB1235 not done in correct conditions: temperature fluctuating. Test repeated. |
| Friday_24_A | TBA2 | No hf DAQ during training quench. | AB/CO | hf | Richard informed. He changed the CPU card and made some trials. |
| Saturday_25_M | TBB2 | Value of the temperature given by the TT821 sensor on the SSS290 is not correct. | MTM |  | Non conformity opened |
| Saturday_25_M | TBA1 | MBB1232: wrong current reading: 1423A instead of 1550 A. MIITS not correct. | MTM |  | ? |
| Saturday_25_A | TBF2 | MBAR2524 unable to be connected by ICS because of cables are to short. | MTM |  | Magnet changed with the MB330? |
| Saturday_25_A | TBF1 | MBB1228: Unable to reach the sufficient voltage to provoke the quench at 3 kA. | MTM |  | Richard informed. He asked to to provoke the quench with QH1&2 |
| Saturday_25_N | TBE1 | No data in hf DAQ since hf acquisition did not trigger. | AB/CO | hf | Richard advised to check the results using the lf DAQ |
| Saturday_25_N | TBA1 | Quench heaters are not charged but TEMA indicated that they are charged. | AB/CO | TEMA |  |
| Saturday_25_N | TBE2 | IAP @cold on MBBR3355 : no data | AB/CO |  | Crate reseted. |
| Sunday_26_M | TBE2 | MBBR3355 -Diode bump is around 500A which seems to be very less compare to usual bump. | MTM |  |  |
| Sunday_26_M | TBD1 | Bad communication between PLCSE and the server after HV test at cold. | AB/CO |  | Sun re-started then OK. |
| Sunday_26_A | TBA2 | Mobic not working (internal error) | ? |  |  |
| Sunday_26_N | TBA2 | LF Aquisition not working . | AB/CO | lf | Crate reseted. Then OK |