| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_30_M | TBC2 | On MB3388: CDBP failed MB-QH vs ground above 600 V. | MTM |  | Salem investigated. Two heaters have insulation weakness. Ask ICS to dismount the IFS box. Problem remains. Magnet warm up and changed. |
| Tuesday-Friday _02_N | TBD2 | Air leak on SSS606 | MTM |  | Leak on the manchette (joint). Leak on one foot (can live with this leak) Leak on the MRB side. |
| Tuesday_30_N | TBF | PLC/HV connector damaged. |  |  | PLC (carefully) disabled in the HV application. |
| Wednesday _31_M | TBA2 | HV test on MB2142: Water Condensation on Leads |  |  | HV test launched after removing the leads. |
| Wednesday _31_A | TBA2 | AQA analysis was not possible for warm up quench ("a03") from control room console | AB/CO | AQA | Quench analysed locally. Salem fixed the AQA problem. |
| Thursday_01_A | TBC2 | PLCSE blocked.: short circuit on Digital Input? | AB/CO? |  | Salem came but could not solve the problem. He decided to call C.Dehavay Friday morning.in the morning the PLCSE was working again. Investigation by C.Dehavay on –going. |
| Friday _02_N | TBB1 | Alarm positioning MRB received by cryo-team : OVC test stopped. | MTM |  | Salem informed Saturday morning |
| Saturday _03_M | TBC2 | Air leak during OVC test. Joint in the manchette not OK | ICS |  | Cryo team asked ICS for intervention. |
| Saturday _03_N | TBE1 | Prep2.1 and Prep2.2 test on E1 : there was an interlock PowerQH_OK. Disappear switching from E1 to E2. Back going to E1 bench. | MTM |  |  |
| Sunday _04_A | TBC2 | Polarity switch blocked. | MTM |  | Salem unblocked it. Maintenance needed: 3 days of stop. |
| Sunday _04_N | TBF2 | Qloc compensation not correct. | MTM |  | OK after shaft in Qloc position.. |