### Chohan /28--Nov-05

### Week 48 starting  28 Nov 05

### SM18 Operation Meeting (Tuesdays)

Bench status on Monday 28 Nov 05:

11/28/2005 2:21:44 PM

| 6 | A1 | MBAR1289 | PT 12 Warm Up | 176.4 |
| --- | --- | --- | --- | --- |
| 11 | A2 |  |  | 291.33 |
| 5 | B1 | SSS114 | PT 12 Warm Up | 300.28 |
| 1 | B2 | SSS281 | PT 8 Fast De-excitation | 1.90 |
| 7 | C1 | MBAL2133 | Prep 5 Cool Down | 96.91 |
| 9 | C2 | MBBR2239 | ICS 4 Final connection | 289.99 |
| 8 | D1 | SSS083 | Prep 3 CDW1 HV | 4.54 HV |
| 4 | D2 | SSS605 | PT 6.1 Training 1 | 4.62 HV |
| 10 | E1 | MBAL3413 | ICS 5 Traveller Last | 294.1 |
| 2 | E2 | MBBL2159 | PT 6.1 Training 1 | 19.8 |
| 3 | F1 | MBAR2247 | PT 11 4 K Quench SSL | 2.35 |
| 12 | F2 |  |  | 288.48 |

Tempr. from TBSM18 site

Coordinator this week: vittorio; MTM support: 16-3602

20 complete Magnet Tests cycles  13 Dipoles  + 7 SSS + 0 SSSS

Disconnected or ROBTC after tests last week ( Sat Midnight to Sat midnight)

[2201, 3388 2nd time ; 3139 3nd time ]

Situation From Cryogenics:  J-P Lamboy or ACR Rep.

Special Points:

- Why 4.4 K quenches on 2224( without shafts ) & Tempr problems

- Special SSS Q6 on D2 (nikolay, Marco)  SSW  issues @  weekend

AOB

| Magnets under Test from 20-11-2005 to 26-11-2005 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBAR2226 |  |  | 20-11-2005 | 33 h. | 23 h. | 22 h. | 19 h. | 9 h. | 24-11-2005 | 106 h. | Stripping |  |  |
| 2 | TBA1 | MBBL2201 |  |  | 14-11-2005 | 30 h. | 26 h. | 36 h. | 25 h. | 10 h. | 20-11-2005 | 127 h. | Stripping |  |  |
| 3 | TBA2 | MBAL3378 |  |  | 16-11-2005 | 54 h. | 25 h. | 36 h. | 15 h. | 18 h. | 23-11-2005 | 148 h. | Stripping |  |  |
| 4 | TBB1 | SSS149 |  |  | 20-11-2005 | 30 h. | 16 h. | 24 h. | 7 h. | 12 h. | 24-11-2005 | 89 h. | Stripping |  |  |
| 5 | TBB1 | SSS161 |  |  | 16-11-2005 | 36 h. | 16 h. | 20 h. | 14 h. | 26 h. | 20-11-2005 | 112 h. | Stripping |  |  |
| 6 | TBB2 | SSS147 |  |  | 17-11-2005 | 36 h. | 28 h. | 14 h. | 11 h. | 21 h. | 22-11-2005 | 110 h. | Stripping |  |  |
| 7 | TBB2 | SSS224 |  |  | 22-11-2005 | 27 h. | 15 h. | 13 h. | 10 h. | 13 h. | 25-11-2005 | 78 h. | Stripping |  |  |
| 8 | TBC1 | MBAR2249 |  |  | 17-11-2005 | 27 h. | 23 h. | 25 h. | 15 h. | 16 h. | 22-11-2005 | 106 h. | Stripping |  |  |
| 9 | TBC1 | MBBL3410 |  |  | 22-11-2005 | 22 h. | 23 h. | 15 h. | 19 h. | 23 h. | 26-11-2005 | 102 h. | Stripping |  |  |
| 10 | TBC2 | MBAR2246 |  |  | 19-11-2005 | 30 h. | 24 h. | 18 h. | 12 h. | 9 h. | 23-11-2005 | 93 h. | Stripping |  |  |
| 11 | TBD1 | SSS165 |  |  | 15-11-2005 | 44 h. | 13 h. | 26 h. | 17 h. | 15 h. | 20-11-2005 | 115 h. | Stripping |  |  |
| 12 | TBD1 | SSS208 |  |  | 20-11-2005 | 34 h. | 17 h. | 23 h. | 7 h. | 28 h. | 25-11-2005 | 109 h. | Stripping |  |  |
| 13 | TBD2 | SSS289 |  |  | 18-11-2005 | 25 h. | 16 h. | 18 h. | 8 h. | 12 h. | 21-11-2005 | 79 h. | Stripping |  |  |
| 14 | TBE1 | MBAL1088 |  |  | 19-11-2005 | 28 h. | 30 h. | 25 h. | 15 h. | 16 h. | 24-11-2005 | 114 h. | Stripping |  |  |
| 15 | TBE2 | MBBR3341 |  |  | 17-11-2005 | 24 h. | 21 h. | 13 h. | 15 h. | 23 h. | 21-11-2005 | 96 h. | Stripping |  |  |
| 16 | TBE2 | MBBR3388 |  |  | 21-11-2005 | 33 h. | 20 h. | 28 h. | 14 h. | 15 h. | 26-11-2005 | 110 h. | Stripping |  |  |
| 17 | TBF1 | MBAL3139 |  |  | 14-11-2005 | 33 h. | 28 h. | 37 h. | 19 h. | 20 h. | 20-11-2005 | 137 h. | Stripping |  |  |
| 18 | TBF1 | MBBL1192 |  |  | 20-11-2005 | 25 h. | 24 h. | 17 h. | 13 h. | 9 h. | 24-11-2005 | 88 h. | Stripping |  |  |
| 19 | TBF2 | MBAR1288 |  |  | 16-11-2005 | 27 h. | 21 h. | 13 h. | 25 h. | 20 h. | 21-11-2005 | 106 h. | Stripping |  |  |
| 20 | TBF2 | MBBR3339 |  |  | 21-11-2005 | 43 h. | 20 h. | 31 h. | 14 h. | 23 h. | 26-11-2005 | 131 h. | Stripping |  |  |

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAR2226 | 12209.9 A. (D2_u) | 12696.9 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 2 | MBBL2201 | 11958.2 A. (D2-L-S01) |  | - | - | - | - | - | - | - | - | - | - | - |
| 3 | MBAL3378 | 12336.3 A. (D2_L_S02) | 12851 A. | - | - | - | - | - | - | - | - | - | - | - |
| 4 | SSS149 | 11740.1 A. (QI_34) | 12479.6 A. (QI-12) | 12666 A. (QI-34) |  | - | - | - | - | - | - | - | - | - |
| 5 | SSS161 | 11617.6 A. (QE 12) | 12217.7 A. (QI 12) | 12388 A. (QI-12) |  | - | - | - | - | - | - | - | - | - |
| 6 | SSS147 | 12847.1 A. (No Data) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 7 | SSS224 | 11153.6 A. (QE-34) | 12059 A. (QE-34) | 12652.1 A. (QI-12) |  | - | - | - | - | - | - | - | - | - |
| 8 | MBAR2249 | 9414 A. (D2_L) | 12087 A. (D1_U) | 12804.1 A. (D1_L) |  | - | - | - | - | - | - | - | - | - |
| 9 | MBBL3410 | 11349 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 10 | MBAR2246 | 11523.6 A. (D2_U) | 11770.9 A. (D1_L) | 12390.7 A. (D1-U) | 12776 A. (D1_U) |  | - | - | - | - | - | - | - | - |
| 11 | SSS165 | 12501.5 A. (QE-12) | 12836.4 A. (QI-12) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 12 | SSS208 | 12177.9 A. (QE-34) | 12387.6 A. (QE-34) | 12433.8 A. (QE-34) |  | - | - | - | - | - | - | - | - | - |
| 13 | SSS289 | 11098 A. (QI_34) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 14 | MBAL1088 | 12296.7 A. (D1-U) | 12539.5 A. (D1-L) | 12777.7 A. (D1_L) |  | - | - | - | - | - | - | - | - | - |
| 15 | MBBR3341 | 9808.5 A. (D2_U) | 10970.5 A. (D2_L) | 12856 A. | - | - | - | - | - | - | - | - | - | - |
| 16 | MBBR3388 | 9728 A. (D1- L) | 9577.3 A. (D1- U) | 11152.4 A. (D2_U) | 11667 A. (D1-U) | 12656.7 A. (D1-U) | 12900 A. | - | - | - | - | - | - | - |
| 17 | MBAL3139 | 12112.4 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 18 | MBBL1192 | 12529.6 A. (D1-L) | 12648 A. (D2-L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 19 | MBAR1288 | 12641.6 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 20 | MBBR3339 | 12728.4 A. (D1-L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |

| 06Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_22_A | TBE1 | Diode test on MB1088: Rcathode exceeded the limit of 15  (also wrong at cold). | MTM |  | Equip support called Pierre who asked to open a NC |
| Tuesday_22_N | TBF1 | MBBL1192: AQA alarm . MIITS AQA exceeds the limit 37(MIITS AQA = 45.1)in PT10 Min Energy. Quench was out of the main coil. | MTM |  | Equip.support contacted during night. Why? TC contacted. He advised to do a SSL at 4.5 K. If the current is greater than 10200 A the magnet can be warmed up. |
| Wednesday_23_M | TBA1 | Loss of cryo OK when discharging the heaters after PT8. | ACR |  |  |
| Wednesday_23_A | TBD1 | SSS208 :PT8 could no be launched. LF DAQ not getting armed | AB/CO | lf | Equip Support called. CPU card error. Spare will be provided by Adrian. |
| Wednesday_23_A | TBB1 | SSS149 : Resistance value of the sextupole voltage taps does not match with the type of connection presented here but with the new connection type although this SSS is not belonging to the appropriate list. | MTM |  | Equip Support called. He will confirm the type of connection. |
| Wednesday_23_N | TBF | Polarity switch blocked (not possible to go in right position) | MTM |  | Equip support called. Polarity inverter forced. |
| Thursday_24_M | TBA1 | Lf DAQ not getting armed. | AB/CO | hf | Guy solved the problem: CPU card not working, Card taken from cluster A has been taken. |
| Thursday_24_N | TBC | Rotary switch blocked. Interlock present. | MTM |  | Equip support called. IDCCT off. Put on in the power supply. |
| Friday_25_M | TBD2 | Lyre test and HV tests done on SSS605. Electrical tests stopped because magnet scheme was missing and there is no traveller. | MTM |  | Traveller brought in the afternoon. |
| Friday_25_A | TBE | LF DAQ is not working. | AB/CO | lf | The CPU card was moved in another bench. Problem fixed by Adrian. |
| Friday_25_A | TBD2 | IAP on SSS605 could not be launched. | MTM | IAP | Guy found that lf config files are not correct. File were corrupted: need to be transferred again. |
| Friday_25_A | TBC2 | MB1283: Provoked quench with LF heaters: AQA warning - Quench location outside main coil. Test repeated. Same result. MB1283: Provoked quench with LF heaters: AQA did not detect the quench. MB1283: Provoked quench with HF heaters: AQA indicated a quench location outside main coil. | MTM | AQA | The AC compensation for D2-U-L was 30% more than what it is normally. The threshold was adjusted. Equip support advised to perform HF provoked quench. Equip support will check the results. |
| Friday_25_N | TBA2 | SPA because of cryo loss. | ACR |  |  |
| Friday_25_N | TBD1 | SSS083 could not be connected by ICS because 3 wires cut on CFB side | ACR |  | Equip support called. Problem fixed. |
| Friday_25_N | TBA2and C2 | Tests could not be launched because HWREC is not responding. | IT |  | Maryline called. Problem solved by Oracle expert (Oracle files corrupted in the DB). |
| Saturday_26_A | TBD2 | SSW stages are not correctly placed, i.e. stage A is on the CFB side and stage B is on the MRB side. | MTM |  | Nikolay asked Guy to installed properly the stages and give the correct sag_calibration_files for SSS605. Done. |
| Saturday_26_N | TBD2 | SSW at warm on SSS605. Movement in opposite direction for Y direction with respect to the command sent by the motor. | MTM |  | DB file was copied from system 1 and did not correspond to the environment of system 3. Guy will correct the DB file. |
| Saturday_26_N | TBC2 | IAP @ Warm final on MB1283 not possible to launch because of a wrong temperature phase "Magnet @ 300K”. Temperature was below 300K but still above 250 K and ACR team did not give the OK | ACR |  | OK when the Temperature was closed to 300 K. |
| Sunday_27_M | TBA1 | PT4 on MB1289: quench at 780 A. Test repeated and same result. AQA did not detect a quench in the two times. | MTM |  | Maryline found that the AC compensation D1-D2 is not correct (offset above 100 mV). However no fault detected in IAP. |
| Sunday_27_A | TBF2 | HV test before powering on MB3406 : HV test with rack MY93 could not start. Same with MY 95 rack. Fault ; wrong cryo phase HV test could be launched but during the first step of test "CDW1" (dipole quench heaters vs. ground) the voltage was out of tolerance | ACR MTM |  | Maryline called. She advised to re-start the PLCSE. In fact problem came fro cryo signal which cannot be delivered because magnet too cold (288K instead of 290 K). Maryline checked and asked ICS to disconnect the magnet and re-do the WP04. Test done after disconnection: problem remained. Decision is to remove the magnet from the bench and open the IFS box. |
| Sunday_27_N | All | 2 Roclas out of 3 are not working. |  |  |  |
| Sunday_27_N | TBF1 | MB2247: Quench below 12000 A (IQ=11160 A) during ramp up for PT8 (SPA test at 12000 A). | MTM |  | Magnet detrained. Re-test needed with quench antenna. |
| Sunday_27_N | TBF1 | MB2247: Noise on direct pole voltages. | MTM | AQA | No alarm: TC advises to go on. |