### Week 19 starting 08 May 06

### SM18 Operation Meeting (Tuesdays) 9 May 06

Bench status  on Monday  08-05-2006 at: 5/8/2006 9:47:38 AM

| 9 | A1 | MBBL1391 | ICS 4 Final connection | 299.56 |
| --- | --- | --- | --- | --- |
| 7 | A2 | MBAL1385 | Prep 5 Cool Down | 76.22 |
| 10 | B1 | SSS229 | Prep 3 CDW1 HV | 299.47 |
| 1 | B2 | SSS504 | PT 10.1: Prov. Quench 1/2 HF 11850 | 1.90 |
| 6 | C1 | MBAR2322 | ICS 6 Disconnect Magnet WP07 | 299.65 |
| 5 | C2 | MBBR1280 | PT 14 Magnet MAPS | 299.65 |
| 12 | D1 | SSS635 | Prep 0 Scan Bench | 298.94 |
| 11 | D2 | SSS609 | Mech not working w'ends/nights | 299.65 |
| 4 | E1 | MBBR2221 | PT 6.2 Training 2 | 19.92 |
| 8 | E2 | MBBR1350 | Prep 5 Cool Down | 91.4 |
| 3 | F1 | MBBR2254 | Extra Provoked Quench | 1.90 |
| 2 | F2 | MBBR2264 | PT 13.2 Resist. Meas. | 299.65 |

Supervisor: Arjan/        Eq support : 16-3602 :

15 Magnet Tests completed :  11 Dipoles (),    3 Arc-SSS  ( 182   REPEAT)

No  500-series ()

No  600-series ( 636 , had 342 hrs but no tests , 609 finished

but not disconnected & removed )

- { D1:635  connection ; D2 : 609 Disconnection & removal

BOTH under  Waiting during past whole weekend 6-7 May  }

ROCALA Repair in front of F cluster last  Thursday 4th May

ROCLA problems 8 May 06 morning

General Points :

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- News /Situation from (SSS, SSSS)  :  J Halik

- Situation SSSS  :OP issues :

- A.O.B.

## Tests

| Magnets under Test from 30-04-2006 to 06-05-2006 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA2 | MBAL1387 |  |  | 25-04-2006 | 26 h. | 48 h. | 32 h. | 19 h. | 10 h. | 01-05-2006 | 135 h. | Stripping |  |  |
| 2 | TBA2 | MBBR2238 |  |  | 01-05-2006 | 29 h. | 18 h. | 25 h. | 21 h. | 19 h. | 06-05-2006 | 112 h. | Stripping |  |  |
| 3 | TBB1 | SSS182 |  |  | 29-04-2006 | 0 h. | 2 h. | 42 h. | 7 h. | 23 h. | 02-05-2006 | 74 h. | Stand By |  |  |
| 4 | TBB2 | SSS232 |  |  | 01-05-2006 | 23 h. | 12 h. | 52 h. | 6 h. | 11 h. | 05-05-2006 | 104 h. | Stand By |  |  |
| 5 | TBB2 | SSS234 |  |  | 26-04-2006 | 35 h. | 20 h. | 18 h. | 13 h. | 9 h. | 30-04-2006 | 95 h. | Stripping |  |  |
| 6 | TBC1 | MBBL2180 |  |  | 26-04-2006 | 28 h. | 27 h. | 24 h. | 22 h. | 0 h. | 30-04-2006 | 101 h. | Stripping |  |  |
| 7 | TBC1 | MBBL2180 |  |  | 30-04-2006 | 18 h. | 21 h. | 26 h. | 15 h. | 18 h. | 04-05-2006 | 98 h. | Stripping |  |  |
| 8 | TBC2 | MBAR1323 |  |  | 28-04-2006 | 46 h. | 24 h. | 29 h. | 17 h. | 9 h. | 03-05-2006 | 125 h. | Stripping |  |  |
| 9 | TBD1 | SSS636 |  |  | 19-04-2006 | 145 h. | - | - | - | - | 03-05-2006 | 342 h. | Not Found or Thermal cycle |  |  |
| 10 | TBE1 | MBAR2317 |  |  | 25-04-2006 | 23 h. | 33 h. | 47 h. | 19 h. | 16 h. | 01-05-2006 | 138 h. | Stripping |  |  |
| 11 | TBE1 | MBBR1348 |  |  | 01-05-2006 | 30 h. | 25 h. | 21 h. | 15 h. | 13 h. | 05-05-2006 | 104 h. | Stripping |  |  |
| 12 | TBE2 | MBAR2236 |  |  | 29-04-2006 | 0 h. | 22 h. | 16 h. | 15 h. | 23 h. | 02-05-2006 | 76 h. | Stand By |  |  |
| 13 | TBE2 | MBBR1262 |  |  | 02-05-2006 | 25 h. | -25 h. | 64 h. | 14 h. | 13 h. | 06-05-2006 | 91 h. | Stripping |  |  |
| 14 | TBF1 | MBBR1344 |  |  | 28-04-2006 | 33 h. | 37 h. | 19 h. | 14 h. | 8 h. | 03-05-2006 | 111 h. | Stripping |  |  |
| 15 | TBF2 | MBAR1297 |  |  | 28-04-2006 | 28 h. | 20 h. | 52 h. | 18 h. | 22 h. | 04-05-2006 | 140 h. | Stripping |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAL1387 | 11212.5 A. (D1-L) | 12236.9 A. (D2_L) | 12501.1 A. (D1_L) |  | - | - | - | - | - | - | - | - | - |
| 2 | MBBR2238 | 11236.3 A. (D1_L) | 12242.4 A. (D1_U) | 11756.6 A. (D1_L) | 12689 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - |
| 3 | SSS182 | 12044.2 A. (QE_34) | 11846.5 A. (No Data) | 11998.9 A. (No Data) | 12426.5 A. (QI-12) | 12497 A. (QE-34) | 12578 A. (QE_12) | 12736.7 A. (QE-12) | 12797.5 A. (QE-34) | 12562.7 A. (QE-34) | 12824 A. (QE_34) |  | - | - |
| 4 | SSS232 | 12147.1 A. (QI_12) | 12367.7 A. (QE-34) |  | - | - | - | - | - | - | - | - | - | - |
| 5 | SSS234 | 12021.3 A. (QE_12) | 12519.4 A. (QE_34) | 12716.7 A. (QE_12) |  | - | - | - | - | - | - | - | - | - |
| 6 | MBBL2180 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 7 | MBBL2180 | 10600 A. (D1_U) | 12607.4 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 8 | MBAR1323 | 12256.2 A. (D1_U) | 12570 A. (D1_U) | 12660 A. (D1_U) |  | - | - | - | - | - | - | - | - | - |
| 9 | SSS636 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 10 | MBAR2317 | 9362.6 A. (D1-L) | 11760.4 A. (D1-U) | 12212.3 A. (D2-L) | 12623.8 A. (D1-L) | 12835.9 A. (D2-l) | 12853 A. | - | - | - | - | - | - | - |
| 11 | MBBR1348 | 11821.4 A. (D1-U) | 12463.1 A. (D2_L) | 12642.5 A. (D2-L) |  | - | - | - | - | - | - | - | - | - |
| 12 | MBAR2236 | 12810.9 A. (D1_U) | 12710.3 A. (D1-U) |  | - | - | - | - | - | - | - | - | - | - |
| 13 | MBBR1262 | 11728 A. (D1_U) | 12178.3 A. (D1_L) | 39 A. (No Data) | 12333.5 A. (D1_L) |  | - | - | - | - | - | - | - | - |
| 14 | MBBR1344 | 12026.8 A. (D1_U) | 12853 A. | - | - | - | - | - | - | - | - | - | - | - |
| 15 | MBAR1297 | 12149.1 A. (D2-L) | 12615 A. (D2-U) | 12000 A. (No Data) | 12853 A. | - | - | - | - | - | - | - | - | - |

Back to Homepage | For print

No  600 series at all this week;  ^09 done but not yet disconnected & removed; 636 ; leak at connection/pumpdown & never tested

SSS182 only after Thermal cycle;

ROCLA Repair on Thurs 4 May Blocked removal of 1297;  D Cluster NOT SHOWN BECAUSE NOTHING HAS HAPPENED THERE

Statistics For end Apr.06					V.Chohan (06_May’ 06)

Statistics For end Apr. 2006

CRYO TIMES

ICS TIMES

MAGNETIC MEASUREMENT TIMES

Percentage of Magnets Tested And Remaining Till End Apr 06

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_2_M | TBC2 | MB1323 : Quench was not detected in PT5.2. As per instruction from Stephane (TC), test repeated. Same problem was seen. | AB/CO | AQA | TC checked the signals. Quench OK. Bug in AQA? |
| Tuesday_2_N | TBF2 | MB1297: magnet OK for the electrical point of view. Magnet is signed for warm up. Cryo leak suspected on this magnet. Investigation at warm by ACR team | MTM |  | Magnet to be signed for stand by at PT14 with comment "electrical test over but signed STANDBY for cryo leak detection" |
| Tuesday_2_N | TBD2 | SSS609: Magnet trained to ultimate PT7 (3900A). Special Test performed by D.Bocian on T-margin. | MTM |  | TC asked to put on Cryo Stand-By and perform HV after the special test. |
| Tuesday_2_N | TBD2 | SSS609: A test named "PT10 Min NRJ 11850A" is noticed in SMTMS with file name "A0605021737.ka01" and comment "done by Giorgio" (part of the special test). AQA shows a current of 37213A with Alert message “DVT exceeds 150 V” for MQY_I (245 V). | MTM |  | TC checked the quench later. Same features for the other provoked quench. TC discussed with A.Siemko. Feature already observed for double MQY. |
| Wednesday_3_M | TBB2 | SSS232: HV at Cold Final. Test Failed for Quadrupole to Ground. Changed HV rack, but the problem exists. | MTM |  | Guy informed. He confirmed the problem that is coming from the MQ. Special investigation by J.Kojak. |
| Wednesday_3_M |  | Video Display on the Bellegarde side for Test Sequence display is not functional |  |  |  |
| Wednesday_3_N | TBC2 | MB1280: Safety matrix was not enabled when interconnection box is changed to C2 position hence quench heater could not be charged for LF heater discharge. | ICS |  | Guy informed. He found that a Vtap in the clamp if broken. He asked ICS to re-open. |
| Fiday_05_A | TBB1 | SSS333: Powering of CD corrector: repeated failure. | MTM |  | Giorgio found that the inductance value was not correctly selected. |
| Fiday_05_A | TBA2 | MB2238: Paper traveler is not available to sign the resistance measurements step after cold test. | ICS? ACR? |  |  |
| Fiday_05_N | TBF1 | MB2254: Open sector (D1_S_07_N) found in PT1(part1). Reading of some other sector (random) was around 280K. | MTM |  | Maryline consulted. She advised to go on: maybe a problem of calibration of some cards. |
| Friday_05_N Sunday_07_N | TBB2 | SSS504: Plexiglass protection sheet for the Quench heater missing. Same problem in TBB1 | MTM |  | Piece of TBB1 taken. P. Legrand will bring a spare. |
| Friday_05_N | TBB2 | SSS504: Clamp connection (CD,EF) check cannot be launched. | MTM |  | Maryline checked.SSS504 had no sextupole circuit. She had to disable the card. |
| Friday_05_N | TBE1 | MB2221: Magnet could not be scanned as one mobic is not able to establish connection with server, and second mobic shows 'EDMS internal error' after login while trying to access mtf site. | IT |  | Maryline informed. She contacted Oracle support. IAP at warm (LF discharge test) was store in the folder of the magnet 1348. Scanning problem solved by Maryline, She deleted the result file and asked to repeat the LF test. |
| Friday_05_N | TBF1 | MB2254: MM1 could not be launched from TEMA: the current does not ramp up. | MTM |  | OP team tried to quit all application and re-launch the test: no success. Maryline informed. She advised to continue with the test program until she came. She found that there is some interlock in the PC and cleared them. |
| Saturday_06_A | TBF1 | MB2254: MM2, Bucking ratio found to be out of specs in one sector. | MTM |  | Maryline contacted. She asked to go on. Stephane checked the results and validate them. He asked to perform a quench before making MM3 because the magnet did not reach 12000 A yet. |
| Saturday_06_A | TBF2 | MB2264: NC generated in CDW1 for problem in TT821 | MTM |  |  |
| Saturday_06_N | TBF1 | MB2254: Extended test loadline_2 kA. Done correctly, MMP can be closed but TEMA cannot be quit. | AB/CO | TEMA | OP team kill TEMA |
| Sunday_07_A | TBB2 | SSS504: HV BP test failed for correctors circuits. Condensation found on the top of CFB | ACR |  | OP team asked ACR team to reduce the helium level on the leads. Then HV passed. |
| Sunday_07_A | TBB1 | SSS229: During clamp connection check, the LF acquisition is not getting triggered. | MTM |  | Maryline informed. She did the test and switched the electronic on TBB2. A trigger on safety matrix is present because the CD circuit is not connected on SSS504. |