### Week 23 starting 5June 06

### SM18 Operation Meeting (Tuesdays) 6 June 06

Bench status  on Monday  05-06-2006 at:  Midday

Bottom of Form Top of Form Bottom of Form

| TBA1 | MBBR1347 | ICS 4 Final connection | 300 K |
| --- | --- | --- | --- |

| TBA2 | MBAR2314 | PT 14 Magnet MAPS | 300 K |
| --- | --- | --- | --- |

| TBB1 | SSS038 | PT 5.2 1.5 KA Quench HF | 1.9 K |
| --- | --- | --- | --- |

| TBB2 | SSS626 | FINISHED SINCE SAT 3rd June morning | 300 K 2P2M |
| --- | --- | --- | --- |

| TBC1 | MBAL2341 | PT 14 Magnet MAPS | 300 K |
| --- | --- | --- | --- |

| TBC2 | MBAR2330 | Prep 5 Cool Down | Cooling |
| --- | --- | --- | --- |

| TBD1 | SSS633 | S4 Field Strength | 4.4 K |
| --- | --- | --- | --- |

| TBD2 | SSS690 | FINISHED since Friday 2nd June 2300 hrs | 300 K 2P2M |
| --- | --- | --- | --- |

| TBE1 | MBAL2340 | PT 13.2 Resist. Meas. | 300 K |
| --- | --- | --- | --- |

| TBE2 | MBAR2273 | Prep 5 Cool Down | Cooling |
| --- | --- | --- | --- |
| TBE2 | MBAR2273 | Prep 5 Cool Down | Cooling |

| TBF1 | MBBR1351 | Prep 5 Cool Down | Cooling |
| --- | --- | --- | --- |

| TBF2 | MBBR2262 | PT 6.2 Training 2 | 1.9 K |
| --- | --- | --- | --- |

Refreshed on 5 /5/ 2006 at 11h58

Top of Form

Free JavaScripts provided
by The JavaScript Source

Bottom of Form

CONFIG :

D8- AS2 –IRSSS2;

Supervisor: nick/        Eq support : 16-3602 :

18 Magnet Tests completed :  11 Dipoles (1  MM),    2 Arc-SSS  ( no   REPEAT)

(16 Magnets in all)		1  500-series ( 505 (

2  600-series [ 647 (313 hrs),  677 (154 hrs) ]

2284 repeated for cryo reason only; 1353 ROBTC

General Points :

- Last weekend problems  (  Magnets Not removed on D2, B2

          - For whole weekend !!

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- News /Situation from (SSS, SSSS)  :  J Halik

- Situation SSSS  :OP issues :

- A.O.B.

Role of Shift Leaders & Visits

Role of Shift Leaders & Special Experiments

Role of Shift Leaders &  ICS, Cryo in SMCR etc

Bench status  on Monday  05-06-2006 at:  Midday

| Magnets under Test on 05-06 11:52:31 AM |
| --- |
| - | Bench | Magnet | Sh. | Tested | Quench | Start | Connect <24h | Cooling <26h | Cold Test <36h | Warming <12h | Disconnect <12h | Total |  |  |  |
| 1 | TBA1 | MBBR1347 |  | 1 | 0 | 03-06 | 43h | - | - | - | - | 43 h. |  |  |  |
| 2 | TBA2 | MBAR2314 |  | 1 | 1 | 31-05 | 2750h | Check Valueh | 2803h | 17h | 10h | 2846 h. |  |  |  |
| 3 | TBB1 | SSS038 |  | 4 | 0 | 01-06 | 51h | 20h | 13h | - | - | 84 h. |  |  |  |
| 4 | TBB2 | SSS626 |  | 1 | 1 | 29-05 | 31h | 35h | 25h | 7h | 51h | 164 h. |  |  |  |
| 5 | TBC1 | MBAL2341 |  | 1 | 3 | 31-05 | 22h | 29h | 24h | 15h | 11h | 109 h. |  |  |  |
| 6 | TBC2 | MBAR2330 |  | 1 | 0 | 04-06 | 45h | 4h | - | - | - | 49 h. |  |  |  |
| 7 | TBD1 | SSS633 |  | 1 | 2 | 01-06 | 47h | 12h | 32h | - | - | 91 h. |  |  |  |
| 8 | TBD2 | SSS690 |  | 1 | 1 | 29-05 | 27h | 0h | 45h | 14h | 66h | 167 h. |  |  |  |
| 9 | TBE1 | MBAL2340 |  | 1 | 3 | 01-06 | 24h | 23h | 18h | 15h | 0h | 84 h. |  |  |  |
| 10 | TBE2 | MBAR2273 |  | 1 | 0 | 02-06 | 38h | 26h | - | - | - | 64 h. |  |  |  |
| 11 | TBF1 | MBBR1351 |  | 1 | 0 | 03-06 | 23h | 20h | - | - | - | 43 h. |  |  |  |
| 12 | TBF2 | MBBR2262 |  | 1 | 2 | 02-06 | 37h | 28h | 15h | - | - | 80 h. |  |  |  |

.

| Magnets under Test from 28-05-2006 to 03-06-2006 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBBL1236 |  |  | 23-05-2006 | 48 h. | 27 h. | 128 h. | 35 h. | 21 h. | 03-06-2006 | 259 h. | Stripping |  |  |
| 2 | TBA2 | MBAR1354 |  |  | 24-05-2006 | 34 h. | 58 h. | 25 h. | 20 h. | 12 h. | 31-05-2006 | 149 h. | Stripping |  |  |
| 3 | TBB1 | SSS086 |  |  | 29-05-2006 | 31 h. | 11 h. | 19 h. | 9 h. | 9 h. | 01-06-2006 | 79 h. | Stripping |  |  |
| 4 | TBB1 | SSS264 |  |  | 25-05-2006 | 33 h. | 15 h. | 29 h. | 9 h. | 10 h. | 29-05-2006 | 96 h. | Stripping |  |  |
| 5 | TBB2 | SSS505 |  |  | 22-05-2006 | 36 h. | 13 h. | 36 h. | 10 h. | 85 h. | 28-05-2006 | 180 h. | Stand By |  |  |
| 6 | TBC1 | MBAR2310 |  |  | 27-05-2006 | 24 h. | 23 h. | 22 h. | 15 h. | 11 h. | 31-05-2006 | 95 h. | Stripping |  |  |
| 7 | TBC2 | MBAR1318 |  |  | 30-05-2006 | 20 h. | 22 h. | 24 h. | 14 h. | 14 h. | 03-06-2006 | 94 h. | Stripping |  |  |
| 8 | TBC2 | MBAR2272 |  |  | 26-05-2006 | 22 h. | 24 h. | 29 h. | 16 h. | 7 h. | 30-05-2006 | 98 h. | Stripping |  |  |
| 9 | TBD1 | SSS647 |  |  | 19-05-2006 | 103 h. | 15 h. | 171 h. | 15 h. | 9 h. | 01-06-2006 | 313 h. | Stand By |  |  |
| 10 | TBD2 | SSS677 |  |  | 22-05-2006 | 32 h. | 16 h. | 78 h. | 9 h. | 19 h. | 29-05-2006 | 154 h. | Stand By |  |  |
| 11 | TBE1 | MBAR2276 |  |  | 28-05-2006 | 18 h. | 24 h. | 21 h. | 14 h. | 18 h. | 01-06-2006 | 95 h. | Stripping |  |  |
| 12 | TBE1 | MBBR2284 |  |  | 23-05-2006 | 28 h. | 25 h. | 20 h. | 24 h. | 17 h. | 28-05-2006 | 114 h. | Stripping |  |  |
| 13 | TBE2 | MBAR1353 |  |  | 25-05-2006 | 22 h. | 40 h. | 41 h. | - | - | 30-05-2006 | 103 h. | Not Found or Thermal cycle |  |  |
| 14 | TBE2 | MBAR1353 |  |  | 30-05-2006 | 0 h. | 32 h. | 21 h. | 20 h. | 14 h. | 02-06-2006 | 87 h. | Stripping |  |  |
| 15 | TBF1 | MBBR2279 |  |  | 22-05-2006 | 51 h. | 32 h. | 45 h. | 15 h. | 14 h. | 29-05-2006 | 157 h. | Stripping |  |  |
| 16 | TBF1 | MBBR2284 |  |  | 29-05-2006 | 20 h. | 35 h. | 22 h. | 18 h. | 25 h. | 03-06-2006 | 120 h. | Stripping |  |  |
| 17 | TBF2 | MBAR2327 |  |  | 29-05-2006 | 19 h. | 34 h. | 13 h. | 20 h. | 8 h. | 01-06-2006 | 94 h. | Stripping |  |  |
| 18 | TBF2 | MBBR2331 |  |  | 24-05-2006 | 27 h. | 28 h. | 18 h. | 15 h. | 9 h. | 28-05-2006 | 97 h. | Stripping |  |  |

### Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBL1236 | 11821.3 A. (D1-L-S01) | 12360 A. (D1-L-S01) | 12784.8 A. (D2_U-L) |  | - | - | - | - | - | - | - | - | - |
| 2 | MBAR1354 | 11331.9 A. (D2_L) | 12448.9 A. (D1_U) | 12649.8 A. (D2_L) |  | - | - | - | - | - | - | - | - | - |
| 3 | SSS086 | 11987.1 A. (QI_12) | 12265.8 A. (QE_12) | 12677.3 A. (QI_12) |  | - | - | - | - | - | - | - | - | - |
| 4 | SSS264 | 11312 A. (QI_12) | 12475 A. (QE-12) | 12771.2 A. (QE_12) |  | - | - | - | - | - | - | - | - | - |
| 5 | SSS505 | 11527.4 A. (QE_12) | 11596.2 A. (QI-12) | 12360.5 A. (QE_34) | 12853 A. | - | - | - | - | - | - | - | - | - |
| 6 | MBAR2310 | 8814.2 A. (D2_L) | 11286.2 A. (D1_L) | 12409.3 A. (D2_L) | 12660.7 A. (D1_U) |  | - | - | - | - | - | - | - | - |
| 7 | MBAR1318 | 12394.8 A. (D1-U) | 12583.9 A. (D1-U) | 12791.6 A. (D1_U) |  | - | - | - | - | - | - | - | - | - |
| 8 | MBAR2272 | 11009.7 A. (D1_L) | 11683.1 A. (D1_U) | 12197.6 A. (D2_L) | 12475.4 A. (D2_U) | 12789.5 A. (D2_L) | 12851.3 A. (D2_U) | 12851 A. | - | - | - | - | - | - |
| 9 | SSS647 | 3622.7 A. (MQY_I_34) | 3759.6 A. (I_Mqy12) | 2935 A. (MQY-E 12) |  | - | - | - | - | - | - | - | - | - |
| 10 | SSS677 | 5424.5 A. (Int_MQMC) | 5727.9 A. (INT_MQMC) | 5779.3 A. (Ext_MQM) | 5820 A. | - | - | - | - | - | - | - | - | - |
| 11 | MBAR2276 | 10928.7 A. (D2_L) | 12189.2 A. (D1_L) | 12526.6 A. (D2_U) |  | - | - | - | - | - | - | - | - | - |
| 12 | MBBR2284 | 11790 A. (D2_U) | 12264.9 A. (D1_U) | 12614.1 A. (D1_U) |  | - | - | - | - | - | - | - | - | - |
| 13 | MBAR1353 | 11226.4 A. (D2_U) | 11548.9 A. (D2_U) | 11742.7 A. (D1_L) | 11963.9 A. (D2_L) | 12300.4 A. (D2_U) | 12429.9 A. (D1_L) | 12594 A. (D2_U) | 12651.2 A. (D1_L) | 12794.6 A. (D1-U) |  | - | - | - |
| 14 | MBAR1353 | 12134.9 A. (D2_L) | 12337.2 A. (D1_L) | 12484.4 A. (D2_U) |  | - | - | - | - | - | - | - | - | - |
| 15 | MBBR2279 | 9843.8 A. (D2_U) | 12030.3 A. (D1_L) | 12547 A. (D2_U) |  | - | - | - | - | - | - | - | - | - |
| 16 | MBBR2284 | 12646.2 A. (D1_U) |  | - | - | - | - | - | - | - | - | - | - | - |
| 17 | MBAR2327 | 12627.2 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 18 | MBBR2331 | 11458.3 A. (D1_L) | 12644.4 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - | - | - | - |

Effective 16 magnets Tested this week

2284 Repeated for Cryo, 1353 ROBTC, 1236 Long MM (marco, juan), ;  505 could Not be removed by ICS (logistics Org.)

647 usual reasons of connections & powering +SSW,    677 had  special tests at cold by someone for beam loss? .. etc)

| Date | Bench | Problem | Type | App. Soft. | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_30_A | TBB1 | SSS08 : HV test front panel not opening in the control room. | AB/CO | HV | To be investigated by AB/CO expert. |
| Tuesday_30_A | TBE1 | MB2276 : No cryo OK for PT4. | ACR |  | ACR team contacted. They reported problem with helium level. Expected to be ok in 45 minutes. |
| Tuesday_30_A | TBA1 | MB1236: Special test on magnetic axis. The team carrying out magnetic measurements on A1 is placing the Laser in the middle of the way to D1 and doing the tests without a proper barricade. Whoever passes by they shout saying that the laser is there and we are not supposed to cross. Also they were blocking operators from carrying out tests. When they were asked by the shift leader to put barricades, the response was very casual asking us to contact Marco. |  |  | Marco was contacted and requested to instruct them to follow safety rules and know the rules of working in sm18. Problem solved. |
| Tuesday_30_A | TBC2 | MB1318: Problems with the MOBIC. It does not read the bar code. Occasionally when it reads, it does not display the number in the corresponding field and the light in the scanner goes off. Tried with both the units but no success. | MTM |  | Finally the magnet details were entered manually. |
| Tuesday_30_A | TBD1 | SSS647: Cryo interlock appearing on TEMA while launching MM2 - LHC Cycle. Tried to Launch from the Sun station down stairs. MM2 could be launched but the magnet quenched at 808A. | MTM |  | Guy asked us to perform one more cleansing cycle and then try again after consulting TC. TC instructed cryo to increase He flow. He asked us to perform training run after getting CRYO OK and if the magnet does not quench below 3000 asked the op team to try running LHC cycle again. |
| Tuesday_30_A | TBD1 | SSS647: Training launched after obtaining CRYO OK but the magnet quenched at 888A. ACR team insisted to perform one more training after implementing more corrective measures. | MTM |  |  |
| Tuesday_30_N | TBD1 | SSS647: Problem of maintaining cryo stability at 4.5K. | ACR |  |  |
| Tuesday_30_N | TBD1 | SSS647: Problem with magnet quenching at low current (about 800A or less) in V_bus_ext and V_bus_int. | MTM |  | The Cryo informed that the problem is not coming from cryogeny. TC discussed with Stephane and decided to cancel the machine cycle. Magnet warm_up. Reason is a long flexible in the N-line that makes a loop and is a reservoir of gaz bubles. Action will be taken among the MTM mechanics to watch out during the connection. |
| Wednesday_31_N | TBA2 | MB2314 : Problem with HV rack MY92 : PLC-HV fault does not allow the test to be carried out. | MTM |  | MY 93 used instead of MY92. Guy was informed; he will look into the problem tomorrow.. |
| Wednesday_31_N | TBE2 | MB1353: Problem with the charge of one quench heater. | IT? |  | Problem will be reported to equipment support in the morning shift. Problem solved by reset of PLCSE. Quench heater are charging now. Problem with the network? |
| Thursday_01_M | TBC1 | MB 2341 : Prep2.1 could not be launched. | IT? |  | Problem solved by resetting the PLCSE. Problem with the network? |
| Thursday_01_A | TBD1 | SSS633: Only magnet name label present. Anti-cryostat and other labels absent. | CRI |  |  |
| Thursday_01_N | TBE1 | MB2340 : Mobic is not accepting the data during the scan Data for E1 and B1 entered manually in MTF. | MTM |  |  |
| Thursday_01_N | TBB2 | SSS626 : NC generated in CDBP for temp sensor not OK | CRI |  | HV fault with Tsensor. |
| Thursday_01_N | TBB1 | SSS038: ICS has informed that, the initial connection of the magnet before Lyra, could not be done completely due to insufficient lengths of some internal wires. | ICS |  | Done Friday Morning by ICS. |
| Thursday_01_N | TBD2 | SSS690 : Black (-) DC cable lead (2A cable) at top of CFB is broken. | MTM |  |  |
| Thursday_01_N | TBF2 | MB2262 : Mobic not working and anticryostat labels absent. | MTM/ICS |  | Only magnet name label read and entered in MTF manually. |
| Friday_02_M | TBB2 | SSS626 : During PT10.1 warm-up quench analysis, AQA alarm appeared, so the magnet is kept on stand-by temp for consultation with TC in morning. | MTM |  | TC checked the data and asked to warm up the magnet. |
| Friday_2_N | TBC1 TBA2 TBB2 | MB2341 :Interlock (No PC Permit)do not allow to lunch PT4 (slow power abort. Same problem. HV panel not displayed. | AB/CO |  | PLCSE reset but no success. Guy called. Communication problem between PLCSE and PLCCU. After Switch off all PLCSE, Reset of PLCCU & restart. OK for TBA & TBB but still a specific communication problem on TBC. RS communication error with PC 14KA. Guy tried to change all rs convertors anf rs card on PC. Finally, AB/CO said to switch to gateway communication to control the PC. OK then. RS problem need to be investigated on Tuesday by AB-CO and/or AB-PO. |
| Saturday_03_M | TBD1 | SSS633 :Vacuum leak detected by ACR in the anti-cryostat. | MTM |  | TC called and expert ACR? Solved by MTM mechanics. |
| Saturday_03_A | TBD2 | SSS690: One pin of DC cable is found broken inside the socket. | MTM |  | Guy contacted. He made a temporary repair of the 20A DC cable connector ((-) lead). The 20A DC leads must be temporarily connected directly to the power leads. He asked to the next shifts: during disconnection, one should not pull the cable, but the connector. to be very careful to the fact that when disconnecting the lead, current must be off (no protection anymore) |
| Saturday_03_A | TBA2 | MB2314 : Mobic problem : In spite of having full charge and full connectivity, the scanned data could not be validated. | AB/CO? |  | Comment by OP team: Frequent problems with mobic, more of a trouble than a facility, leading to manual updating of magnet and accessories info in to the MTF of late |
| Saturday_03_N | TBA1 | MB1347: SMS from ICS to perform Lyra test. OP team observed that WP04 was not uploaded. | ICS |  | WP=4 (HV test) was not done? HV test failed and performed again by ICS |
| Sunday_04_M | TBD1 | SSS633 : the powering of the K-L circuit failed. | MTM | LF | Guy advised to proceed with MQY training. |
| Sunday_04_M | TBC2 | Sun MTA 58 monitor screen is not working. The sun is hung. | AB/CO |  | Guy called and asked OP team to re-boot the sun. |
| Sunday_04_N | TBB1 | SSS038: PT4 not working. The ramp in the power supply cannot start. PLCSE correct. Problem with FB Loop interlock. Corrector test is not working | AB/PO |  | We continue with correctors and will do PT4 later |