### Week 28 starting 17 July 06

### SM18 Operation Meeting (Tuesdays) 18 July 06

Bench status  on Monday  17-07-2006 at:   7/17/2006 9:27:17 AM

| 9 | A1 | MBAL1396 | PT 13.2 Resist. Meas. | 301.35 |
| --- | --- | --- | --- | --- |
| 12 | A2 | B. Repair |  | 300.28 |
| 8 | B1 | SSS142 | ICS 3 Traveller after warm test | 301.89 |
| 2 | B2 | SSS275 | PT 12 Warm Up | 1.90 |
| 7 | C1 | SSS138 | PT 3 QH Measure | 2.35 |
| 3 | C2 | SSS266 | PT 6.6 Training 6 | 1.91 |
| 10 | D1 | SSS636 | MTM Mech non-avail | 300.01 |
| 1 | D2 | SSS645 | PT 6.6 Training 6 | 4.51 HV |
| 6 | E1 | MBAR1326 | ICS 4 Final connection | 302.6 |
| 5 | E2 | MBAL2353 | PT 13.1 IAP @ Warm | 298.04 |
| 11 | F1 | MBBR2296 | PT 13.2 Resist. Meas. | 299.65 |
| 4 | F2 | MBAL2349 | PT 8 Fast De-excitation | 1.92 |

CONFIG :  D 6 - AS 3 -  IRSSS 3

Supervisor: Gijs (Monday) ; Pierre after;Eq support : 16-3602

15 Magnet Tests completed :

8 Dipoles (2290  with Magnetic Mesurment),

5 Arc-SSS  (271 3 times thermal cycle)

0  500-series [ ]

2  600-series [ 676 (166 hrs), 682 after TH Cy (150 hrs) ]

LOT of StandBy’s & lot of HV problems (Real + humidity etc) this week

682: 6 MQTL’s ; probable open circuit in connection between 4 & 5

General Points :

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- News /Situation from (SSS, SSSS)  :  J Halik

- Situation SSSS  :OP issues :

- A.O.B..

| Magnets under Test from 09-07-2006 to 15-07-2006 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBBR1306 |  |  | 06-07-2006 | 41 h. | 37 h. | 64 h. | 15 h. | 13 h. | 13-07-2006 | 170 h. | Stripping |  |  |
| 2 | TBA2 | MBBR2290 |  |  | 06-07-2006 | 29 h. | 29 h. | 16 h. | 17 h. | 26 h. | 11-07-2006 | 117 h. | Stand By::HV CQBP |  |  |
| 3 | TBB1 | SSS258 |  |  | 07-07-2006 | 30 h. | 14 h. | 16 h. | 8 h. | 26 h. | 11-07-2006 | 94 h. | Stripping |  |  |
| 4 | TBB1 | SSS262 |  |  | 11-07-2006 | 33 h. | 16 h. | 23 h. | 9 h. | 20 h. | 15-07-2006 | 101 h. | Stripping |  |  |
| 5 | TBB2 | SSS676 |  |  | 05-07-2006 | 69 h. | 24 h. | 27 h. | 11 h. | 35 h. | 12-07-2006 | 166 h. | Stand By |  |  |
| 6 | TBC1 | SSS047 |  |  | 04-07-2006 | 22 h. | 14 h. | 72 h. | 6 h. | 8 h. | 09-07-2006 | 122 h. | Stand By: HV CQAP |  |  |
| 7 | TBC1 | SSS274 |  |  | 09-07-2006 | 36 h. | 24 h. | 30 h. | 6 h. | 13 h. | 14-07-2006 | 109 h. | Stand By :TT821 |  |  |
| 8 | TBC2 | SSS141 |  |  | 08-07-2006 | 33 h. | 16 h. | 46 h. | 7 h. | 13 h. | 14-07-2006 | 115 h. | Stand By Corr Pwrng |  |  |
| 9 | TBD2 | SSS682 |  |  | 06-07-2006 | 2 h. | 22 h. | 86 h. | 14 h. | 26 h. | 12-07-2006 | 150 h. | Stand By: MQTL repair |  |  |
| 10 | TBE1 | MBAL2354 |  |  | 10-07-2006 | 45 h. | 24 h. | 32 h. | 15 h. | 12 h. | 15-07-2006 | 128 h. | Stripping |  |  |
| 11 | TBE1 | MBBR1315 |  |  | 06-07-2006 | 27 h. | 22 h. | 15 h. | 15 h. | 11 h. | 10-07-2006 | 90 h. | Stand By :TT821 |  |  |
| 12 | TBE2 | MBBR1349 |  |  | 08-07-2006 | 24 h. | 19 h. | 23 h. | 15 h. | 10 h. | 12-07-2006 | 91 h. | Stripping |  |  |
| 13 | TBF1 | MBAL1397 |  |  | 06-07-2006 | 58 h. | 26 h. | 17 h. | 16 h. | 25 h. | 12-07-2006 | 142 h. | Stripping |  |  |
| 14 | TBF2 | MBAL1398 |  |  | 09-07-2006 | 28 h. | 25 h. | 19 h. | 23 h. | 18 h. | 14-07-2006 | 113 h. | Stripping |  |  |
| 15 | TBF2 | MBBR1356 |  |  | 05-07-2006 | 25 h. | 25 h. | 20 h. | 14 h. | 14 h. | 09-07-2006 | 98 h. | Stripping |  |  |

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBR1306 | 11036.2 A. (D1_U) | 11370.5 A. (D1_L) | 11918.4 A. (D2-L) | 12503 A. (D1-U) | 12779.9 A. (D2_L) | 12850 A. | - | - | - | - | - | - | - |
| 2 | MBBR2290 | 12216.5 A. (D1_U) | 12771.2 A. (D2_U) | 12852 A. | - | - | - | - | - | - | - | - | - | - |
| 3 | SSS258 | 12105.6 A. (QE_12) | 12552 A. (QE_12) | 12728 A. (QI-34) |  | - | - | - | - | - | - | - | - | - |
| 4 | SSS262 | 11188.9 A. (QE_34) | 12032.8 A. (QE_34) | 12000 A. (No Data) | 12370 A. (QE-34) |  | - | - | - | - | - | - | - | - |
| 5 | SSS676 | 5491 A. (Ext-34) | 5820 A. | - | - | - | - | - | - | - | - | - | - | - |
| 6 | SSS047 | 10863 A. (QI_34) | 11131.2 A. (QE_12) | 11509.7 A. (QE_34) | 11893.2 A. (QI_12) | 12501.5 A. (QE_12) | 12850 A. | - | - | - | - | - | - | - |
| 7 | SSS274 | 12110.9 A. (QI_12) | 12421 A. (QE_12) | 12676.9 A. (QI_12) |  | - | - | - | - | - | - | - | - | - |
| 8 | SSS141 | 11579.1 A. (QI_12) | 11914.8 A. (QE_12) | 12163.7 A. (QI_12) | 12337.2 A. (QI-12) | 12756 A. (QI-12) | 12850.4 A. (Q1_12) | 12850 A. (Vclamp E) | 12850 A. (Vbus_Ext) | 12850 A. | - | - | - | - |
| 9 | SSS682 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 10 | MBAL2354 | 10281.6 A. (D1_L) | 11409.3 A. (D2_L) | 11880.7 A. (D1_L) | 12450 A. (D2-U) | 12850 A. | - | - | - | - | - | - | - | - |
| 11 | MBBR1315 | 12613.4 A. (D2_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - | - |
| 12 | MBBR1349 | 11323.2 A. (D1_U) | 12301.5 A. (D1_L) | 12294.1 A. (D1-L) | 0 A. (No Data) |  | - | - | - | - | - | - | - | - |
| 13 | MBAL1397 | 12446.6 A. (D2- L) | 12612.1 A. (D1 - L) | 0 A. (No Data) | 12805.9 A. (D1_L) |  | - | - | - | - | - | - | - | - |
| 14 | MBAL1398 | 12359.2 A. (D2_L) | 12317.5 A. (D2_U) | 12517.1 A. (D2_U) |  | - | - | - | - | - | - | - | - | - |
| 15 | MBBR1356 | 11182 A. (D1_L) | 12212.8 A. (D2-U) | 12146.3 A. (D2-L) |  | - | - | - | - | - | - | - | - | - |

1306 took long because of cohabitation with 2290 : both on A cluster at same time;682 : 6 MQTL’s  2nd run had problem ; 1397 long connect time ?

STATISTICS for endJune 2006

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_11_M Tuesday_11_A | TBD1 | SSS636 : MM1+MM2+MM3 -Problem with_T_coil : no power supply. -Power supply of T_coil need to be configured -T_coil: One temperature out of range. - MM1 is not correct -MM2, MM3 results are displaying sectors in over-ranges. | MTM |  | Power supply taken from another bench. Richard advised to proceed. Check by Stephane and Richard. Not correct config_files. For MQY single, test “family_3” config files have to be selected. |
| Tuesday_11_N | TBA2 | MB2347: file of PREP2.1 (IAP@warm LF heater discharge) saved in 2347_2 folder due to mistaken selection in TEMA | MTM |  | Eq sup contacted Richard copied the file in the correct folder. |
| Tuesday_11_N | TBB1 | SSS262: DAQ (LF) did not trigger in Prep2.2C (clamp connection check). | MTM |  | Richard contacted. He informed POTAIM card was adjusted for previous magnet (without CD and EF circuits) hence LF to be triggered manually. Test conducted with manual trigger |
| Tuesday_11_N | TBD1 | SSS636: Over range errors in sector S0405 of Ap2 in MM2. | MTM |  | Richard in consultation with Stephane suggested to validate this in MM2 and MM3 as the concerned sector is in middle and extrapolation is feasible. |
| Tuesday_11_N | TBF2 | MB1398: PT4 (ramp to 1000A) could not be performed. Power application shows an error '222 at serial port checking.' and interlock 'VS_FAULT' remains after reset. Attempt was made with reset of server PLC as well as Power converter without success. | MTM |  | Richard contacted. VS_FAULT is cleared after switching off excess water connections (limit of 3 power supplies at the same time). |
| Tuesday_11_N | TBC2 | SSS141: Quench occurred two times at ultimate (in the clamp). 2nd quench occurred after holding 12850 for 5min and shown by AQA outside the magnet.. Manual MIITS is 19.4 (seems very low @ 12850A). AQA does not register this value. | ACR |  | Problem temporarily solved by increasing the level of helium in current lead). |
| Tuesday_11_N | TBD1 | SSS636 : Warm up quench could not launched. PLC shows fault on QH even though all individual Quench Heaters are charged. | MTM |  | Richard contacted. The cable connecting the quench heaters and the magnet was not present. It has not been connected after Hv test (PT9). |
| Wednesday_12_M | TBA2 | MB2347: Anti-cryostat interlock in cryo application, but no anti-cryostat present on this dipole. Investigation of cryo is going on. | ACR |  | PLC anti-cryostat is not working properly in this cluster |
| Wednesday_12_M | TBC1 | SSS274: HV insulation (CDBP) Sextupole vs Ground failed. Water condensation has been found on CD EF current leads. | MTM |  | OK after having dried the current leads with a fan (by ACR). |
| Wednesday_12_A | TBC2 | SSS141: Corrector powering GH PT5.5, problem. Resistance of the G current lead too high. | MTM |  | Fixed after intervention of A. Ballarino dismounting and re-mounting the radiator of the current lead. |
| Thursday_13_M | TBE2 | MB2353: HW recognition not OK, because of configuration files missing for prep2.2 (IAP @ warm LF). | MTM | HWREC | OK after having restarted TEMA. |
| Thursday_13_M | TBE2 | MB2353 : HV (Prep 3) could not be continued due to problems in SUN server. TBE1 also is affected, as PT2 is due on E1. | IT |  | Richard is investigating the issue. |
| Thursday_13_M | TBC1 | SSS274 : TEMA application is hanging while trying to launch training for warming up the magnet, | IT |  | SUN expert has been called. |
| Thursday_13_M | TBB1 | SSS262 : Sun machine is not responding in B bench, while trying to launch HV. | IT |  | SUN Expert informed us to not switch off any application or sun because the problem on E can be extended to other sun. AFS server down and IT upgrade. |
| Thursday_13_N | TBB1 | SSS262: During launching PT5.3 (corrector powering CD) VS_Fault present on CD and EF power converter. | AB/PO? |  | Richard called. Emergency button was pressed. |
| Saturday_15_N | TBD2 | SSS645: Current did not ramp in PT4. Error 222 serial port, Test was launched again but file is not getting generated. Triggers on card 5,7,9,13 existing | AB/PO |  | Richard was contacted. Communication failure with RMD power converter. Richard had to reset the DICO. This can be done after exiting power and server applications. |
| Sunday_16_M | TBA2 | MB2347: Paper traveller not available. ICS was informed. E- Traveller signed | ICS? |  | The magnet was removed without paper traveller signed. |
| Sunday_16_M | TBE1 | MB1326: Trigger present on safety matrix. | ICS |  | Richard came: One V_tap soldered in the magnet connection by IC, is open at cold. Richard put a bridge. |
| Sunday_16_M | TBC2 | SSS266: problem with TT821. T sensor shorted to the ground. | MTM |  | Richard informed. He asked to open a NC. |
| Sunday_16_N | TD2 | SSS645: Fourth training quench: current ramped up to 2323 A and dropped down to zero without quench & file generation. | AB/PO? |  | Slow PA abort. Problem with RMD again? |