### vc/Week 31 starting 31 July 06

### SM18 Operation Meeting (Tuesdays) 01 Aug 06

Bench status  on Monday  31-07-2006 at:  7/31/2006 9:41:18 AM

| 7 | A1 | MBBR2362 | Prep 5 Cool Down | 104.81 |
| --- | --- | --- | --- | --- |
| 4 | A2 | MBAR2361 | PT 12 Warm Up | 13.25 |
| 6 | B1 | SSS320 | PT 13.2 Resist. Meas. | 299.83 |
| 11 | B2 |  |  | 300.55 |
| 3 | C1 | SSS312 | PT 6.2 Training 2 | 1.91 |
| 9 | C2 | SSS267 | Prep 5 Cool Down | 156.85 |
| 1 | D1 | SSS650 | waiting for shafts removal | 284.01 |
| 2 | D2 | SSS624 | PT 6.4 Training 4 | 11.53 |
| 10 | E1 | MBAR1358 | PT 14 Magnet MAPS | 302.52 |
| 8 | E2 | MBAL2281 | Prep 5 Cool Down | 92.61 |
| 12 | F1 | MBBR1364 | PT 14 Magnet MAPS | 299.65 |
| 5 | F2 | MBBR1362 | PT1 IAP@1.9K, Part III | 1.90 |

CONFIG :  D 6 - AS 3 -  IRSSS 3 ;

Supervisor: Arjan  ;Eq support : 16-3602

17 Magnet Tests completed

8 Dipoles (2 with MM ,  1 repeat- 2339  was supposed to be for cryo only),

7 Arc-SSS  tests but ONLY 5  magnets( 119 repeated & 257 HV pblms)

0  500-series [ ]

2  600-series [ 623 (148 hrs), 650 1st run - shaft pblms (187 hrs) ]

Major Power Failure at CERN  on Sat morning ; Put us out for 18 + hrs

SSS259 & MB 1364 held up in warming due to this power failure

Big Gate & Cars blocked inside ..resolved only at 17.30 Sat aft

General Points :

C2 Bench was blocked by   sss257 for 2 weeks !!

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- News /Situation from (SSS, SSSS)  :  J Halik

- Situation SSSS  :D1 & D2 issues

- A.O.B..

## Tests

| Magnets under Test from 23-07-2006 to 29-07-2006 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBAL1392 |  |  | 24-07-2006 | 28 h. | 22 h. | 25 h. | 14 h. | 16 h. | 29-07-2006 | 105 h. | Stripping |  |  |
| 2 | TBA1 | MBBR1363 |  |  | 17-07-2006 | 38 h. | 26 h. | 60 h. | 16 h. | 8 h. | 24-07-2006 | 148 h. | Stripping |  |  |
| 3 | TBA2 | MBAL2350 |  |  | 21-07-2006 | 23 h. | 36 h. | 10 h. | 40 h. | 8 h. | 26-07-2006 | 117 h. | Stand By |  |  |
| 4 | TBB1 | SSS119 |  |  | 21-07-2006 | 26 h. | 12 h. | 45 h. | 8 h. | 0 h. | 25-07-2006 | 91 h. | Not Found or Thermal cycle |  |  |
| 5 | TBB1 | SSS119 |  |  | 25-07-2006 | 1 h. | 14 h. | 11 h. | 9 h. | 16 h. | 27-07-2006 | 51 h. | Stripping |  |  |
| 6 | TBB2 | SSS268 |  |  | 21-07-2006 | 34 h. | 26 h. | 35 h. | 8 h. | 11 h. | 26-07-2006 | 114 h. | Stripping |  |  |
| 7 | TBC1 | SSS269 |  |  | 19-07-2006 | 28 h. | 15 h. | 44 h. | 10 h. | 10 h. | 23-07-2006 | 107 h. | Stripping |  |  |
| 8 | TBC1 | SSS321 |  |  | 23-07-2006 | 32 h. | 15 h. | 31 h. | 8 h. | 12 h. | 28-07-2006 | 98 h. | Stripping |  |  |
| 9 | TBC2 | SSS257 |  |  | 18-07-2006 | 23 h. | 21 h. | 122 h. | 27 h. | 0 h. | 27-07-2006 | 193 h. | Not Found or Thermal cycle |  |  |
| 10 | TBC2 | SSS257 |  |  | 27-07-2006 | 0 h. | 12 h. | 6 h. | 23 h. | 33 h. | 29-07-2006 | 74 h. | Stand By |  |  |
| 11 | TBD1 | SSS650 |  |  | 17-07-2006 | 77 h. | 24 h. | 70 h. | 16 h. | 0 h. | 25-07-2006 | 187 h. | Not Found or Thermal cycle |  |  |
| 12 | TBD2 | SSS623 |  |  | 20-07-2006 | 29 h. | 20 h. | 75 h. | 12 h. | 12 h. | 27-07-2006 | 148 h. | Stand By |  |  |
| 13 | TBE1 | MBAL2351 |  |  | 20-07-2006 | 23 h. | 24 h. | 48 h. | 14 h. | 11 h. | 25-07-2006 | 120 h. | Stripping |  |  |
| 14 | TBE2 | MBAL2338 |  |  | 19-07-2006 | 26 h. | 25 h. | 18 h. | 14 h. | 7 h. | 23-07-2006 | 90 h. | Stripping |  |  |
| 15 | TBE2 | MBAR1359 |  |  | 25-07-2006 | 20 h. | 25 h. | 21 h. | 14 h. | 25 h. | 29-07-2006 | 105 h. | Stand By |  |  |
| 16 | TBF1 | MBAL2339 |  |  | 17-07-2006 | 36 h. | 22 h. | 21 h. | 14 h. | 54 h. | 23-07-2006 | 147 h. | Stripping |  |  |
| 17 | TBF2 | MBBL2358 |  |  | 19-07-2006 | 45 h. | 25 h. | 107 h. | 14 h. | 11 h. | 28-07-2006 | 202 h. | Stripping |  |  |

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAL1392 | 11953 A. (D1-L) | 12404 A. (D1-U) |  | - | - | - | - | - | - | - | - | - | - |
| 2 | MBBR1363 | 10124.8 A. (D2_U) | 12622 A. (D2-U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 3 | MBAL2350 |  | - | - | - | - | - | - | - | - | - | - | - | - |
| 4 | SSS119 | 11520.5 A. (QE_34) | 11912.9 A. (QE_12) | 12195 A. (QE-34) | 12280 A. (QE-34) | 12375 A. (QE-34) | 12363.7 A. (QE_12) | 12472.1 A. (QI_34) | 12416 A. (QI_12) | 12523.1 A. (QI_12) | 12559.3 A. (QI_12) |  | - | - |
| 5 | SSS119 | 12344.5 A. (QE_34) | 12540 A. (QE_12) |  | - | - | - | - | - | - | - | - | - | - |
| 6 | SSS268 | 11276.1 A. (QI-12) | 11425.8 A. (QI-12) | 12128.5 A. (QI-12) | 12850 A. | - | - | - | - | - | - | - | - | - |
| 7 | SSS269 | 11395.8 A. (QI_12) | 11811.5 A. (QI_12) | 12018.1 A. (QE_34) | 12491.5 A. (QE_12) | 12451.6 A. (QI_12) | 12700.4 A. (QI_34) | 12851.7 A. (QI_12) | 12850 A. | - | - | - | - | - |
| 8 | SSS321 | 11143.8 A. (QE_34) | 11604.8 A. (QE_34) | 12350.5 A. (QE_12) | 12612.6 A. (QE_34) |  | - | - | - | - | - | - | - | - |
| 9 | SSS257 | 12030.5 A. (QE_34) | 12804.8 A. (QI_34) | 31.1 A. (No Data) | 12852 A. | 12850 A. | - | - | - | - | - | - | - | - | - |
| 10 | SSS257 |  | - | - | - | - | - | - | - | - | - | - | - | - |  |
| 11 | SSS650 | 4318.7 A. (Ex_MQ_34) | 4335 A. (IN_MQ_12) | 4403.6 A. (Ex_MQ34) | 4433.1 A. (I-MQML34) | 4484.3 A. (I-MQML34) | 4500.3 A. (E_MQML12) | 4511.7 A. (Ext-M_12) | 4524.3 A. (Int-MQ12) | 4570.8 A. (Int-MQ34) | 4625.8 A. (I_MQML) | 4650 A. | - | - |  |
| 12 | SSS623 | 3905.4 A. (Ext_MQML) | 4134.3 A. (E_MQML12) | 4211.9 A. (E_MQML34) | 4260 A. (Int_MQML) | 4366.6 A. (E_MQML34) | 4518.6 A. (E_MQML12) | 4579 A. (I_MQML12) | 4629.6 A. (I_MQML34) | 4638.6 A. (I_MQML12) | 4650 A. | - | - | - |  |
| 13 | MBAL2351 | 11149 A. (D2_L) | 11248.9 A. (D1_L) | 12093.2 A. (D2_U) | 12310.6 A. (D2_U) | 12391 A. (D2_U) | 12507.9 A. (D2_U) | 12497.5 A. (D2_U) | 12660.4 A. (D2_U) | 12850 A. | - | - | - | - |  |
| 14 | MBAL2338 | 11336.6 A. (D2-U) | 12340.7 A. (D1-U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |  |
| 15 | MBAR1359 | 12540 A. (D2_U) | 12759.3 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |  |
| 16 | MBAL2339 | 11597 A. (D2_U) | 12302.1 A. (D2_U) | 12416.7 A. (D1_L) | 12511 A. (D1_U) |  | - | - | - | - | - | - | - | - |  |
| 17 | MBBL2358 | 10352.5 A. (D2_L) | 12008 A. (D1_U) | 12601 A. (D1-U) | 12850 A. | - | - | - | - | - | - | - | - | - |  |

257 stayed on bench previous w’end without tests unnecessarily & we could have offloaded it earlier : overall ~half month of C2 Time

2358 & special hall probe m’ments:  one whole night magnet was kept cold without any tests

1363 had MM ; 623 had long training; 2351 had long training

650 .. long training to ultimate; shafts could not be put at cold in 1st run & so magnet had to be warmed up, readjustments made ..& then start 2nd run.. magnet now finished only on Monday 31st July…disconnection start

25 AUG 06 (Morning) to 31 AUG 06 (Night)

| 25A(TUE) |
| --- |
| TBF2 |  | The power converter tripped as the power supply for the controller module overheated. Guy was called. | He attended to it & by this time the temperature reduced and the converter was functional again. |

| 26M(WED) |
| --- |
| TBA2 | ES | TBA2: MBAL2350 – E-traveler for PT13.1 and PT13.2 is signed but actual tests are not conducted. | This is as per request from Richard (Eq. Support). HV test PT2 has failed and Richard is carrying further investigation |

| 25N(TUE) |
| --- |
| TBC1 | ES | TBC1: During IAP at Cold (AC Compensation), all potaim cards are getting triggered immediately after applying the ac voltage to AC patch panel. Guy was. | 2 potaim card (V-clamp and V supra) were triggering when applied AC voltage is above 0.5Vp-p. Guy reported that voltage taps were inversed inside CFB. A special connector is fitted to care of this. Tests are being resumed |
| TBD1 | CR | TBD1: Temp too high AC Bench D1: Informed Equipment Support in consultation with Cryo. | CRYO said this is faulty, because of disconnection of anticryostats for repair, and OP team gets an SMS every hour. This cannot be cancelled |

| 26N(FRI) |
| --- |
| TBA1 | ES | TBA1:PT4 could not be launched due to Hardware Recognition Error. Guy came and checked that TT821 of MB2361 is missing in mtf, so HR create an error while lunching the on TBA1.. | Guy switched off HWR to go on with test on A1. Tests on A2 will wait until the data is updated in MTF on Thursday |
| TBC1 | ES | TBC1: During PT6.4 (warm up quench), no data was available on LF. So there was no QH information available. But from all other AQA information, it was concluded that the magnet can be warmed up | Eq sup may please check to restore LF data if possible |

| 27M(THU) |
| --- |
| TBC2 | ES | TBC2:- PT1-Part could not be launched. Permanent triggers. Eq. Sup confirmed about two missing voltage taps. Problem being investigated by them | After investigation equipment support instructed CRYO team to warm up the magnet naturally. Magnet to be signed for stand by as the voltage taps problem still persists. |

| 28N(FRI) |
| --- |
| TBE1 | ES | TBE1 : Problem in CDAP: Passive bus bar internal vs. external failed. Run recovery attempted three times but the step was failing. Guy contacted. | ES solved the problem. |
| TBD1 | ES | TBD1 : Magnetic measurement program giving alert (MOTOR PROBLEM!) while performing Full load line at position (1340,1338). Guy contacted. | It was found that there was over range error both in aperture 1 and 2. also, the standard deviation corr. to shaft rotation was very high |

| 29M(SAT) |
| --- |
|  |  | Major Power Supply Interruption. |  |
| TBD1 |  | TBD1: MM2 (Previous shift informed us to repeat MM2). During second run there was a problem with data acquisition system and measurement was aborted after 1Hour. There was a power interruption immediately after aborting the test. | PS related problem |

| 29A(SUN) |
| --- |
|  |  | As per instructions from Nick, contacted Jerome for problems related to servers and displays. SMTMS and sm18op site is operational but MTF is still not accessible. Contacted IT helpdesk and got the information that database servers will be operational only after one hour. | PS related |
| TBA1 |  | TBA1:TBE2: New magnets not scanned as mobic is not operational as well as MTF application is not accessible. |  |

| 29N(SUN) |
| --- |
| D1&B1 |  | TBD1 & TBB1: No water supply for at least 2 hours so no powering tests can be launched. | PS related problems |
| TBB1 |  | TBB1: Water interlock on PLCSE vanished. However there is still a water interlock in Cryo process |  |
| TBC |  | TBC: The display of sunmta58 (Power Application M/C's) is not functional. Thus unable to launch Prep2.2C from tema. Equipment Support to be contacted in Morning. |  |

| 30M(SUN) |
| --- |
| TBD1 |  | Nick instructed us to warm up D1 - the shafts will be removed at warm. |  |

| 31M(MON) |
| --- |
| TBF | ES | TBF: no current reading on F bench, after troubleshooting Equ. support called the AB/PO piquet. | They found a problem inside the PC |

| 31A(MON) |
| --- |
| TBC1 | ES | TBC1: IJ current leads seems to have too much contact resistance that is why the current cannot reach 60A because as it reaches 30A the voltage level for the detection is reached also (4V). | Equ.support called TC to send a technician from AT/MEL on the ground and check it. |

| 31N(MON) |
| --- |
| TBD2 |  | TBD2: As per last shift, HV test problems probably due to MY92 HV rack battery low. Being tried with new rack. Informed to Eq.Sup as there is shortage of racks. | ES to take action |
| TBE2 |  | TBE2: Lot of condensation at the current leads on CFB, so HV could not be carried out. Heater installed and being moved on all leads time to time. | Condensation problems |

- Frequent Water condensation at CFB.

- Trigger is generated as soon as PT1(AC comp) cable is connected.

- One HV rack was repaired & put back in service.  Still problems are reported.

- Power supply interruption related problems