### VC-EV/Week 40 starting 02 October 06

### SM18 Operation Meeting (Tuesdays) 03 Oct 06

Bench status  on Sunday-01-2006 at 10/1/2006 11:43:22 AM

| 10 | A1 | MBBL1381 | ICS 4 Final connection | 297.86 |
| --- | --- | --- | --- | --- |
| 3 | A2 | MBBL2337 | PTE Min. NRJ 3KA Diode Test | 1.91 |
| 4 | B1 | SSS351 | ICS 6 Disconnect Magnet WP07 | 298.58 |
| 7 | B2 | SSS242 | Prep 5 Cool Down | 4.13 HV |
| 9 | C1 | SSS355 | Prep 5 Cool Down | 93.82 |
| 2 | C2 | SSS341 | ICS 2 Connect Magnet (ICS) | 297.41 |
| 11 | D1 | SSS317 | Prep 3 CDW1 HV | 296.97 |
| 6 | D2 | SSS666 | PT 13.2 Resist. Meas. | 297.41 |
| 5 | E1 | MBAL2344 | PT 2 HV Cold Test | 1.90 |
| 8 | E2 | MBBL1378 | Prep 5 Cool Down | 133.48 |
| 1 | F1 | MBBR2383 | Prep 3 CDW1 HV | 296.43 |
| 12 | F2 | MBBL2178 | ICS 4 Final connection | 297.24 |

CONFIG :  D 6 - AS 5 -  IRSSS 1 ;

Supervisor:  Marta  ;Eq support : 16-3602

19 Magnet Tests completed ( 16 Magnet Stripped )

11 Dipoles tests

6 Arc-SSS  tests (S139 120 h. (with MM ), S242 ROBTC)

0  500-series

2  600-series [661  (225 hrs), 673  (183 hrs)]]

General Points:

- Situation from Cryogenics:  J-P Lamboy or ACR Rep.

- Situation SSSS  :D1 & D2 issues Halik Jerzy;

- SSS351: Investigation HV

- SSS242 cryo difficulties to cooling down…

- A.O.B..

# From 9/24/2006 to 9/30/2006

Tests Information

| Magnets under Test from 24-09-2006 to 30-09-2006 |
| --- |
| - | Bench | Magnet | Shafts | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |
| 1 | TBA1 | MBAL1384 |  | 25-09-2006 | 23 h. | 21 h. | 18 h. | 15 h. | 21 h. | 29-09-2006 | 98 h. | Stripping |
| 2 | TBA1 | MBBR2380 |  | 20-09-2006 | 26 h. | 20 h. | 22 h. | 22 h. | 26 h. | 25-09-2006 | 116 h. | Stripping |
| 3 | TBA2 | MBBR1370 |  | 22-09-2006 | 40 h. | 23 h. | 42 h. | 23 h. | 19 h. | 28-09-2006 | 147 h. | Stripping |
| 4 | TBB1 | SSS359 |  | 23-09-2006 | 20 h. | 15 h. | 47 h. | 9 h. | 8 h. | 27-09-2006 | 99 h. | Stripping |
| 5 | TBB2 | SSS139 |  | 16-09-2006 | 27 h. | 12 h. | 120 h. | 18 h. | 41 h. | 25-09-2006 | 218 h. | Stripping |
| 6 | TBB2 | SSS242(2) |  | 25-09-2006 | 38 h. | 13 h. | 55 h. | 8 h. | 0 h. | 30-09-2006 | 114 h. | Thermal cycle |
| 7 | TBC1 | SSS144 |  | 24-09-2006 | 25 h. | 17 h. | 32 h. | 7 h. | 12 h. | 28-09-2006 | 93 h. | Stripping |
| 8 | TBC1 | SSS244 |  | 20-09-2006 | 25 h. | 25 h. | 14 h. | 7 h. | 16 h. | 24-09-2006 | 87 h. | Stripping |
| 9 | TBC2 | SSS236(2) |  | 23-09-2006 | 29 h. | 24 h. | 10 h. | 11 h. | 12 h. | 27-09-2006 | 86 h. | Stripping |
| 10 | TBC2 | SSS256 |  | 27-09-2006 | 25 h. | 14 h. | 19 h. | 6 h. | 16 h. | 30-09-2006 | 80 h. | Stripping |
| 11 | TBD1 | SSS661 |  | 20-09-2006 | 59 h. | 15 h. | 111 h. | 17 h. | 23 h. | 29-09-2006 | 225 h. | Stand By |
| 12 | TBD2 | SSS673 |  | 18-09-2006 | 74 h. | 22 h. | 34 h. | 15 h. | 38 h. | 26-09-2006 | 183 h. | Stand By |
| 13 | TBE1 | MBAR1376 |  | 23-09-2006 | 29 h. | 30 h. | 17 h. | 14 h. | 19 h. | 28-09-2006 | 109 h. | Stripping |
| 14 | TBE2 | MBBL1377 |  | 19-09-2006 | 25 h. | 20 h. | 30 h. | 16 h. | 29 h. | 24-09-2006 | 120 h. | Stripping |
| 15 | TBE2 | MBBR2375 |  | 24-09-2006 | 37 h. | 22 h. | 17 h. | 13 h. | 10 h. | 28-09-2006 | 99 h. | Stripping |
| 16 | TBF1 | MBBL1379 |  | 22-09-2006 | 28 h. | 20 h. | 16 h. | 17 h. | 8 h. | 26-09-2006 | 89 h. | Stripping |
| 17 | TBF1 | MBBR2386 |  | 26-09-2006 | 24 h. | 34 h. | 11 h. | 17 h. | 21 h. | 30-09-2006 | 107 h. | Stripping |
| 18 | TBF2 | MBBR1342 |  | 15-09-2006 | 40 h. | 21 h. | 123 h. | 25 h. | 2 h. | 24-09-2006 | 211 h. | Stripping |
| 19 | TBF2 | MBBR2385 |  | 24-09-2006 | 49 h. | 34 h. | 12 h. | 17 h. | 14 h. | 29-09-2006 | 126 h. | Stripping |

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAL1384 | 10245.4 A. (D1_L) | 12120.7 A. (D1_L) | 12471.8 A. (D1_L) |  | - | - | - | - | - | - |
| 2 | MBBR2380 | 11320 A. (D2 - L) | 12117 A. (D2 - U) | 12850 A. | - | - | - | - | - | - | - |
| 3 | MBBR1370 | 10828.1 A. (D1_L) | 12170.1 A. (D1_U) | 12113 A. (D2-U) |  | - | - | - | - | - | - |
| 4 | SSS359 | 10622 A. (QI_12) | 12387.1 A. (QI_34) | 11967 A. (No Data) | 12531 A. (QI_12) |  | - | - | - | - | - |
| 5 | SSS139 | 11060 A. (qi_12) | 11754.3 A. (QE - 12) | 11942.2 A. (QI -12) | 12087.5 A. (QI_34) | 12202 A. (QI_12) | 12382 A. (QI_12) | 12574 A. (QI_12) | 12692 A. (QI_12) | 12768.4 A. (QE_12) | 12838.4 A. (QI_12) | 12850 A. |
| 6 | SSS242 | 12281.8 A. (QI-12) | 12850 A. | - | - | - | - | - | - | - | - |  |
| 7 | SSS144 | 10890.8 A. (QE_34) | 11205.6 A. (QE_34) | 12213.6 A. (QI-34) | 12623.2 A. (QI_34) | 12797 A. (QE-34) | 12735.2 A. (QI_34) | 12745.9 A. (QI_34) | 12850 A. | - | - |  |
| 8 | SSS244 | 12426.5 A. (QE_34) | 12850 A. | - | - | - | - | - | - | - | - |  |
| 9 | SSS236 | 12312.8 A. (QI_34) | 12850 A. | - | - | - | - | - | - | - | - |  |
| 10 | SSS256 | 11934.2 A. (QE_12) | 12136.1 A. (QE_12) | 12578.9 A. (QE_12) |  | - | - | - | - | - | - |  |
| 11 | SSS661 | 3559.1 A. (E-MQY_12) | 3562 A. (I-MQY-34) | 3684.1 A. (I-MQY-12) | 3750.5 A. (No Data) | 3792.2 A. (Ex_MQ_34) | 3817.2 A. (Ex_MQY12) | 3838.8 A. (In_MQY12) | 3862.1 A. (Ex_MQY12) | 3795.9 A. (E_MQY_34) | 3891 A. (In_MQY12) | 3900 A. |
| 12 | SSS673 | 3850 A. (E.MQM.2) | 4130 A. (E.MQM.2) | 4232 A. (E.MQM.1) | 4287.8 A. (Ex-MQM_1) | 4429.8 A. (Ex-MQM-1) | 4537.8 A. (I_MQM_2) | 4590.5 A. (E_MQM_2) | 4642.2 A. (I_MQM_2) | 4650 A. | - |  |
| 13 | MBAR1376 | 11754.6 A. (D1_L) | 12614.6 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - |  |
| 14 | MBBL1377 | 12538.6 A. (D1_L) | 12821.2 A. (D2_L) | 12859 A. | - | - | - | - | - | - | - |  |
| 15 | MBBR2375 | 11516 A. (D2-U) | 12816.5 A. (D1_U) | 12850 A. | - | - | - | - | - | - | - |  |
| 16 | MBBL1379 | 11705.5 A. (D2_U) | 12198 A. (D1_U) | 12831.1 A. (D1_L) |  | - | - | - | - | - | - |  |
| 17 | MBBR2386 | 12121.5 A. (D1_L) | 12850 A. | - | - | - | - | - | - | - | - |  |
| 18 | MBBR1342 | 12301 A. (D2_U) | 12850 A. | - | - | - | - | - | - | - | - |  |
| 19 | MBBR2385 | 11415.8 A. (D1_U) | 12203.9 A. (D2_U) | 12741.4 A. (D1_L) |  | - | - | - | - | - | - |  |

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_26_A | TBD1 | S4 661 : while doing Extended MM 6 during ramp down the application reported motor/encoder error. On advice of Stephan we are proceeding with the further tests. Further it was observed that on the AP1 motor/encoder toruqe loss LED was glowing. To check if it will cause problem during the next MM.. | MTM |  | Contacted Fedrico. He suggested to do Hardware test. On performing this test the error vanished |
| Wenesday_27_A | TBD2 | Scanning the new magnet on D2 perturbed hardware recognition. | MTM |  | Called Federico. HWREC cannot find TT821 data in MTF. Federico skipped the errors inside config files. Test will proceed. Eq. Support will inform reponsible for MTF data. |
| Wenesday_27_N | TBB1 | During Prep 2.2C (CD clamp connection test) problem in LF Data acquisition, LF is not getting armed. ICS contacted and they cheched and informed op team that connections are OK.Equipment support contacted. | MTM |  | Federico came, software pbm maybe, now it works |
| Wenesday_27_N Thursday_28_M | TBB2 | PT 5.4 (EF powering) failed -- power supply problem (Invertor failure Indicator was glowing) | MTM |  | Richard was informed. He solved the problem.The test was carried out |
| Wenesday_27_N | TBA1 | Cryo not OK hence PT9 i.e. CDAP can not be launched | ACR |  | Waiting 2 h cryo Ok, after PT8 without quench. |
| Thursday_28_M | TBB2 | Problem in carrying out PT-9. The test is failing at first step i.e. Quadrapole Vs Ground. Richard is investigating. The first step was s kipped and it was restarted from the second step.It is ok. | MTM |  | Richard will investigated the problem.Keep the rack connected. |
| Thursday_28_A Friday_29_M | TBB2 | S242 HV test CQAP failed at Quadrupole Vs. ground step. Richard asked OP team to maintain the magnet on 1.9K until 14:30 for investigation by A.Siemko. Richard called again at 17:00Hrs and asked OP team to keep the magnet on standby till tommorrow morning for investigation by A. Siemko. | MTM |  | The short between magnet and ground detected yesterday has mysteriously disappeared. Please repeat a PT5.2 quench and then the CQAP HV sequence to check if the problem reappears. Richard |
| Friday_29_M | TBC1 | (SSS355) There was problem of connection on V_Supra_lead and V_clamp_lead during clamp connection test CD (Prep2.2C). ICS contacted. After the connection was made OK by ICS, there was another problem of V_clamp_F trigger. | MTM/ICS |  | The lead wire was soldered properly but the wire was damaged internally. Euipment support told ICs to connect lead wire and it was connected. |
| Friday_29_N | TBA2 | While doing Prep-5, DC cable at CFB could not be plugged, as it's +ve pin is broken and stucked in female counterpart. | MTM |  | So AC cable used. |
| Saturday_30_M | TBD2 | QH voltage OK was not coming after all the heaters are charged. It was found that back panel connection of Quench Heater 3 was broken. Informed Frederico. He will come and attend to the problem. | MTM |  | Federico repaired the broken connector. Problem solved. |
| Saturday_30_M | TBE2 | Alarm anticryostat Federico informed. | AB/CO | ART | Federico checked ART system, it looks good then he checked CRYO PLC (error on communication module was found) so he restarted CRYO PLC and the problem was solved. |
| Saturday_30_M | TBD1 | Leak on the X line. ICS found that the line is not straight. | ICS |  | The afternoon shift will have to disconnect the magnet and re-align it.connection |
| Saturday_30_A | TBB2 | SSS242: Magnet is warm, However due to problem with the cryo the magnet is in cryo standby and has not started cooling down. Waiting for Cryo team to resolve the problem | CRYO | IAP | SSS242 Cryo problem resolved. Magnet cooling down. |
| Sunday_31_M | TBD2 | SSS 666: During the CQW2 test There is bad remark in template at row ,Do we test Diplole corrector, even all ok are present. Quench Heaters resistance test (CQW2) at YT311 and YT321 ( has been retested and varified with MTR) has marginal deviation from the values of CQW1 | MTM |  | ??? |
| Sunday_31_M | TBA2 | Cluster A. Checking quench analysis of the current magnet (2337) and previous magnets, we find systematically warnings of "HF/LF current discrepancy", and the main voltage tap signals do not correspond to the usual 200mV/10ms trigger delay. (NB. The quenches are detected correctly by the hardware) TBA2 Training. Much noise on voltage taps, and DirVt voltage jump warning. Federico suggested continuing and to call him if the next quench is worse. | MTM |  | Called Federico about quench analysis on cluster A. He suspects an acquisition timing problem and will look at it tomorrow. |