### EV- /Week 44 starting 30 October 06

### SM18 Operation Meeting (Tuesdays) 31 Oct 06

Bench status  on Monday 30-oct-2006 at:

10/30/2006 7:02:06PM

| 7 | A1 | MBBR2398 | Prep 5 Cool Down | 126.43 |
| --- | --- | --- | --- | --- |
| 3 | A2 | MBAR1413 | PT 6.3 Training 3 | 2.35 |
| 1 | B1 | SSS 514 | PT 3 QH Measure | 2.35 |
| 10 | B2 | SSS515 | PT 13.2 Resist. Meas. | 297.41 |
| 12 | C1 | SSS308 | PT 14 Magnet MAPS | 295.54 |
| 6 | C2 | SSS237 | PT 12 Warm Up | 210.2 |
| 2 | D1 | SSS521 | Faulty Quench | 1.90 |
| 5 | D2 | SSS679 | PT5.3: Powering Corrector CD | 2.35 |
| 8 | E1 | MBAR1411 | ICS 3 Traveller after warm test | 296.79 |
| 9 | E2 | MBAL2350 | PT 13.2 Resist. Meas. | 298.04 |
| 11 | F1 | MBAR2372 | Finished Ready to go Final scan | 295.36 |
| 4 | F2 | MBAR2270 | PT 1 IAP @ cold | 2.03 HV |

CONFIG :  D 6 - AS 2 -  IRSSS 4 ;

Supervisor:  Vittorio (Begin of week), Nick (End of week) ;

Eq support 16-3602 : Salem till Friday, Giogio Saterday/Sunday

11 Magnet Tests completed ( 9 Magnet Stripped )

7 Dipoles tests	(1410 with MM)

1 Arc-SSS (352)

2  500-series (520 Trained the training ramp 10 + SSW + MM)

1  600-series []

General Points:

- Situation from Cryogenics:  J-P Lamboy or ACR Rep :

  - Saterday/Monday Cryo Problem  (Only 1 or 2 Magnets @cold at the same time)?

  - Several faulty quench on D1, during week end.

- SSW on D1 not possible to launch test Sunday afternoon.

- MM on 1413 shaft not rotating properly.

- Situation SSSS  :D1 & D2 issues J Halik

- A.O.B..

| Magnets under Test from 22-10-2006 to 28-10-2006 |
| --- |
| - | Bench | Magnet | Shafts | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |
| 1 | TBA1 | MBAL2342 |  | 23-10-2006 | 22 h. | 23 h. | 32 h. | 17 h. | 39 h. | 28-10-2006 | 133 h. | Stripping |
| 2 | TBA1 | MBBR2382 |  | 14-10-2006 | 35 h. | 30 h. | 100 h. | 16 h. | 14 h. | 23-10-2006 | 195 h. | Stripping |
| 3 | TBA2 | MBAR1410 |  | 19-10-2006 | 37 h. | 31 h. | 71 h. | 25 h. | 7 h. | 26-10-2006 | 171 h. | Stripping |
| 4 | TBB1 | SSS512 |  | 20-10-2006 | 27 h. | 15 h. | 20 h. | 9 h. | 29 h. | 25-10-2006 | 100 h. | Stripping |
| 5 | TBB2 | SSS662(2) |  | 18-10-2006 | 56 h. | 29 h. | 14 h. | 10 h. | 19 h. | 24-10-2006 | 128 h. | Stand By |
| 6 | TBC2 | SSS352 |  | 16-10-2006 | 50 h. | 51 h. | 22 h. | 8 h. | 9 h. | 24-10-2006 | 140 h. | Stripping |
| 7 | TBD1 | SSS520 |  | 13-10-2006 | 30 h. | 13 h. | 140 h. | 12 h. | 45 h. | 23-10-2006 | 240 h. | Stand By |
| 8 | TBE1 | MBBR2392 |  | 16-10-2006 | 28 h. | 81 h. | 17 h. | 17 h. | 8 h. | 22-10-2006 | 151 h. | Stripping |
| 9 | TBE2 | MBAL1382 |  | 16-10-2006 | 22 h. | 95 h. | 30 h. | 14 h. | 28 h. | 24-10-2006 | 189 h. | Stripping |
| 10 | TBF1 | MBBR2396 |  | 18-10-2006 | 35 h. | 47 h. | 25 h. | 15 h. | 25 h. | 24-10-2006 | 147 h. | Stripping |
| 11 | TBF2 | MBBR2397 |  | 22-10-2006 | 36 h. | 26 h. | 35 h. | 16 h. | 16 h. | 27-10-2006 | 129 h. | Stripping |

Quench Information

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBAL2342 | 11986.6 A. (D1-L) | 12377.9 A. (D1-U) | 8604.6 A. (No Data) | 12635.7 A. (D1_U) |  | - | - | - | - | - |
| 2 | MBBR2382 | 11203.3 A. (D2_U) | 11919.7 A. (D1-L) | 12814.2 A. (D2_U) | 12850 A. | - | - | - | - | - | - |
| 3 | MBAR1410 | 11378 A. (D2-L) | 12075 A. (D2-L) | 12643.3 A. (D1_U) | 12850 A. | - | - | - | - | - | - |
| 4 | SSS512 | 12562.9 A. (QI_12) | 12852 A. | - | - | - | - | - | - | - | - |
| 5 | SSS662 | 5629.3 A. (INT_12) | 5820 A. | - | - | - | - | - | - | - | - |
| 6 | SSS352 | 11290.7 A. (QE_12) | 12361.1 A. (QE_12) | 12845.4 A. (QI_12) |  | - | - | - | - | - | - |
| 7 | SSS520 | 10849.8 A. (QI-34) | 11781.9 A. (QI-34) | 12104.4 A. (QE_12) | 12383.9 A. (QI_34) | 12549.1 A. (QE-12) | 12668.1 A. (QI_12) | 12662 A. (QI_34) | 12796.4 A. (28.7) | 12847 A. (QE_12) | 12850 A. |
| 8 | MBBR2392 | 12164.3 A. (D2_U) | 12339 A. (D1-L) | 12850 A. | - | - | - | - | - | - | - |
| 9 | MBAL1382 | 10679.6 A. (D1_L) | 12067.2 A. (D1_L) | 12621.2 A. (D1_U) |  | - | - | - | - | - | - |
| 10 | MBBR2396 | 11723 A. (D2_U) | 12702.8 A. (D2_U) | 12850 A. | - | - | - | - | - | - | - |
| 11 | MBBR2397 | 10950 A. (D2-U) | 11670 A. (D2_L) | 12173 A. (D2_L) | 11948.6 A. (D2_L) | 12226.4 A. (D2_L) | 12342.8 A. (D2_L) | 12732.7 A. (D2_U) | 12833.9 A. (D1_U) | 12850 A. | - |

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Tuesday_24_N | TBB2 | SSS515: trigger in EF clamp. | ICS |  | OP team asked ICS to redo the connection. |
| Wednesday_25_M | TBA1 | MB2342: Quench validation delays are wrong. Problem not solved. MXI/VME error loading hf DAQ Richard replaced Potaim card n2 and HF DAQ card n4 and now the validation time is OK | MTM |  | Richard changed Potaim card No.1 to see if that solves the problem Richard changed card 4 in the HF crate. |
| Wednesday_25_N | TBD1 | SSS521: Alarm anticryostat. ACR team informed OP team | ICS |  | Blue cable not connected. Connected by OP team. |
| Wednesday_25_N | TBE2 | MB2350 : Impossible to launch Prep2.1 IAP LF, because cryo is not in the good phase (HV 4.5 K phase) | ACR |  |  |
| Wednesday_25_N | TBF2 | MB2397: PT10.1 Provoked quench @ Nominal HF and LF are in 2 different files. | AB/CO | Hf,lf | Magnet warmed up because analysis could be done and everything is OK. Richard is however informed. |
| Wednesday_25_N | TBE1 | MB2325: No trigger detected by AQA during PTE. | AB/CO |  |  |
| Thursday_26_M | TBC2 | SSS237: The scanning of the magnet name is not getting validated due to database problem. | IT |  | Label was wrong. NC opened. Tests prepared manually. |
| Thursday_26_A | TBB2 | SSS515: Temperature anti-cryostat too low in aperture 2. | MTM |  | It looks like the gamelle is missing in aperture 2 or the connector inside cryostat is not connected by mistake. Richard asked to go on with the cooldwon and the tests. |
| Thursday_26_N | TBC2 | SSS237: Not possible to do Lyre test, because ICS unreachable. | ICS |  | OP team waited next team around 5:00 No manpower available for Lyra test. Done at 10h00. |
| Saturday_28_M | TBB1 | SSS514: Lyra test was done without clamp connection check. OP team asked ICS but they say no information from previous shift.. | ICS |  |  |
| Saturday_28_M | TBB1 | SSS514: OP team executed the clamp connection check for EF but for CD circuit check not possible as special Adopter is required for V-taps of 2 MQTL | MTM |  | Richard informed. Richard put this adaptor in place at the back of the second potaim crate. It will be removed at the end of the test campaign. Clamp test on CD can be done. |
| Saturday_28_A | TBB1 | SSS514: During IAP at warm prep2.2, TT821 Temperature sensor Error Tmax=267.6 Tmin=267. | MTM |  | Richard contacted. He suggested checking if the error appeared at in IAP cold also then only generating a NC. |
| Saturday_28_A | TBD1 | SSS521: Problem on Corrector GH. Quench one (gf01)57A then detraining (gf02)55A and last (gf08)50A. | MTM |  | Investigation and TC asked to continue the training. |
| Saturday_28_A | All | Too many magnets at 1.9 K with only one pumping unit (4 magnets). Leak in the second one, worl at 18g/s. The liquefactor cannot purify the return gaz at full power. | ACR |  | ACR team asked OP team to put some magnets in stand_by. |
| Saturday_28_N | TBD1 | SSS521: SSW installation. The small white wire from the SSW return wire that goes into the plastic tube was broken. | MTM |  | Repaired with a “bricolage”. |
| Sunday_29_M | TBB2 | SSS515: Ultimate reached on 11th ramp. However, current ramped down automatically to zero without any file generation after about 4 minutes. | MTM |  | OP team executed another ramp to ultimate. |
| Sunday_29_M | TBD1 | SSS521: SSW measurements cannot be done because the motor controller could not read the wire tension correctly. | MTM |  | Richard came and changed the motor controller in Aperture 2. |
| Sunday_29_A | TBD1 | SSS521: SSW measurements failed. Integrator was not picking up the signal at-all. | MTM |  | Richard informed. He suggested checking the Input cable to integrator to see any damage. Control by Guy Monday. Reset everything. OK. |
| Sunday_29_N | TBA2 | MB1413: Bad rotation in Aperture 2 during MM2. Kevin tried to improve it by rectifying the position of the end-cover. Still RMS at 3% for the rotation. | MTM |  | Stéphane asked to Mechanics to check the TRU/Shaft connection. |
| Sunday_29_N | TBD1 | SSS521: .Training of the MQ, two faulty quenches at 11000 A. Quench in the bus-bars. | ACR |  | Stéphane asked Lionel to tune the cooling parameters. Then OK. |
| Sunday_29_N | TBB2 | SSS515: Corrector Magnet (E-F) Voltage measured at MQTL EE861 & EE961 is non-conform. (Reverse sign wrt the template). | MTM |  | Richard or Salem will confirm the measurements.. |