### VC- /Week 50 starting   18Dec06

### SM18 Operation Meeting (Tuesdays) 19 dec  06

Bench status  on monday 18-Dec-2006 at   12/18/2006 2:56:26 PM

| 2 | A1 | MBAR2411 | PT 6.1 Training 1 | 1.90 |
| --- | --- | --- | --- | --- |
| 7 | A2 | MBAR2416 | PT 14 Magnet MAPS | 295.18 |
| 5 | B1 | SSS519 | ICS 5 Traveller Last | 296.88 |
| 1 | B2 | SSS516 | PT1 IAP@1.9K, Part III | 1.90 |
| 8 | C1 |  |  | 295.09 |
| 3 | C2 | SSS370 | Thermal Cycle | 2.88 HV |
| 4 | D1 | SSS669 | PT 6.9 Training 9 | 14.47 |
| 6 | D2 | SSS695 | PT 13.2 Resist. Meas. | 295.09 |
| 9 | E1 | MBAR2402 | Prep 5 Cool Down | 239.86 |
| 10 | E2 | MBBR2394 | Prep 3 CDW1 HV | 294.73 |
| 11 | F1 | MBBR1345 | PT 14 Magnet MAPS | 294.55 |
| 12 | F2 | MBAL1406 | PT 14 Magnet MAPS | 295.89 |

CONFIG :  D 6 - AS 2 -  IRSSS 4 ;

Supervisor:  Vittorio ;        Eq support 16-3602 :

Last Quench on Thursday 21st Dec at 1700 hrs ..  Shutdown Friday evening 22nd ( no night shift)

10 Magnet Tests completed ( 7 Magnet Stripped ) ;

5 Dipoles tests

2 Arc-SSS (  365, 368 )

1 500-series (  525 (245 hrs) )

2  600-series [ 674( 187 hrs)   694 (311 hrs)]

Situation from Cryogenics:  J-P Lamboy .

General Points:

- Situation SSSS  :D1 & D2 , B1 B2 issues J Halik : start situation  8 Jan ?

  - PUT 509 on B ..   any other 5-series ?

  - Emptying D cluster completely before Friday 22nd ?

- Eq Support Issues last Sat night & HWREC ..  (reboot like W-XP ? , no solution in 3 yrs?)

- A.O.B

TERMINATION TICK OFF COUNTER

( SPECIAL & NORMAL SSS )

| MAGNET | ARRIVAL | DEPARTURE | Tick if Done |
| --- | --- | --- | --- |
| 525 | 05/12/2006 | 15 dec 06 |  |
| 509 |  |  |  |
| 523 |  |  |  |
| 516 | 15 dec 06 |  |  |
| 519 | 08/12/2006 | Waiting to go |  |
| 529 |  |  |  |
| 530 |  |  |  |
| 531 |  |  |  |
| 669 | 14 dec 06 | Waiting to go |  |
| 695 | 11/12/2006 | Waiting to go |  |
| 370 | 12 dec 06 17Dec :ROBTC |  |  |
| 361 |  |  |  |
| 099 | 18/12/2006 |  |  |
| 374 |  |  |  |
| 356 |  |  |  |

Updated :  December 20, 2006

TERMINATION TICK OFF COUNTER

( DIPOLES )

| MAGNET | ARRIVAL | DEPARTURE | TICK IF DONE |
| --- | --- | --- | --- |
| 1345 | 10-12-2006 | Waiting to Depart |  |
| 2263 |  |  |  |
| 2267 |  |  |  |
| 2268 |  |  |  |
| 2326 |  |  |  |
| 2328 |  |  |  |
| 2417 | 09-12-2006 | 16-12-2006 |  |
| 2418 |  |  |  |
| 2419 |  |  |  |
| 2420 |  |  |  |
| 1406 | 12-12-2006 | Waiting to Depart |  |
| 2363 | 30-11-2006 | 13-12-2006 |  |
| 2394 | 16-12-2006 |  |  |
| 2399 |  |  |  |
| 2402 | 16-12-2006 |  |  |
| 2407 |  |  |  |
| 2408 |  |  |  |
| 2411 | 13-12-2006 |  |  |
| 2414 |  |  |  |
| 2415 | 07-12-2006 | 13-12-2006 |  |
| 2416 | 05-12-2006 | Waiting to Depart |  |
| 2524 |  |  |  |
| 2384 |  |  |  |

1334

## Tests

| Magnets under Test from 10-12-2006 to 16-12-2006 |
| --- |
| - | Bench | Magnet | Shafts | + | Start Date | Connect Time | Cooling Time | Cold Test Time | Warming Time | Disconnect Time | End Date | Total | Signed for |  |  |
| 1 | TBA1 | MBBR2363 |  |  | 30-11-2006 | 58 h. | 81 h. | 66 h. | 16 h. | 56 h. | 13-12-2006 | 277 h. | Stripping |  |  |
| 2 | TBB2 | SSS525 |  |  | 05-12-2006 | 97 h. | -19 h. | 124 h. | 13 h. | 30 h. | 15-12-2006 | 245 h. | Stripping |  |  |
| 3 | TBC1 | SSS368 |  |  | 05-12-2006 | 29 h. | 17 h. | 40 h. | 11 h. | 76 h. | 12-12-2006 | 173 h. | Stripping |  |  |
| 4 | TBC2 | SSS365 |  |  | 29-11-2006 | 79 h. | 78 h. | 14 h. | 6 h. | 115 h. | 11-12-2006 | 292 h. | Stripping |  |  |
| 5 | TBD1 | SSS674 |  |  | 05-12-2006 | 86 h. | 14 h. | 63 h. | 11 h. | 13 h. | 13-12-2006 | 187 h. | Stand By |  |  |
| 6 | TBD2 | SSS694 |  |  | 28-11-2006 | 60 h. | 107 h. | 50 h. | 20 h. | 74 h. | 11-12-2006 | 311 h. | Stand By |  |  |
| 7 | TBE1 | MBAR2415 |  |  | 07-12-2006 | 39 h. | 19 h. | 15 h. | 14 h. | 20 h. | 13-12-2006 | 107 h. | Stripping |  |  |
| 8 | TBE2 | MBBR2417 |  |  | 09-12-2006 | 42 h. | 23 h. | 17 h. | 16 h. | 52 h. | 16-12-2006 | 150 h. | Stripping |  |  |
| 9 | TBF1 | MBAR1414 |  |  | 30-11-2006 | 40 h. | 101 h. | 21 h. | 16 h. | 17 h. | 10-12-2006 | 195 h. | Stripping |  |  |
| 10 | TBF2 | MBAR2413 |  |  | 01-12-2006 | 74 h. | 70 h. | 21 h. | 14 h. | 39 h. | 11-12-2006 | 218 h. | Stand By |  |  |

| - | Magnet | Q1 | Q2 | Q3 | Q4 | Q5 | Q6 | Q7 | Q8 | Q9 | Q10 | Q11 | Q12 | Q13 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | MBBR2363 | 12031.3 A. (D1_U) | 12849.4 A. (No Data) | 12850 A. | - | - | - | - | - | - | - | - | - | - |
| 2 | SSS525 | 11050.4 A. (QI_34) | 12226.8 A. (QE_34) | 12477.7 A. (QE-34) | 12623.5 A. (QE_34) | 12750 A. (QE-34) | 12832.5 A. (QI-34) | 12850 A. | - | - | - | - | - | - |
| 3 | SSS368 | 10989 A. (QI-12) | 11126.6 A. (QE-34) | 12223.2 A. (QE_12) | 12406 A. (QE-34) | 12525 A. (QE_34) | 12684.6 A. (QE_34) | 12850 A. | - | - | - | - | - | - |
| 4 | SSS365 | 12541.8 A. (QE_34) | 12585.3 A. (QI_34) | 12750.7 A. (QI-34) |  | - | - | - | - | - | - | - | - | - |
| 5 | SSS674 | 3660.3 A. (Int_MQM) | 3838.8 A. (Ext_MQM) | 3927.2 A. (Int_MQM) | 4128.2 A. (Int_MQM) | 4216 A. (Int_MQM) | 4311.2 A. (Ext_MQM) | 4374.9 A. (Ext_MQML) | 4406.1 A. (Int_MQM) | 4434.4 A. (EXT_MQM) | 4453 A. (Ext-MQM) | 4471 A. (I-MQM) | 4485 A. (E-MQML) | 4504.6 A. (I_MQML) | 4551.5 A. (INT-MQM) | 4572.7 A. (Ext_MQM) | 4604.1 A. (Int-MQM) | 4626 A. (E-MQTL) |  |
| 6 | SSS694 |  | - | - | - | - | - | - | - | - | - | - | - | - |  |  |  |  |  |
| 7 | MBAR2415 | 11336.5 A. (D1_L) | 12169.7 A. (D2_U) | 12765.8 A. (D2_L) |  | - | - | - | - | - | - | - | - | - |  |  |  |  |  |
| 8 | MBBR2417 | 11422 A. (D1-U) | 11996 A. (D2-U) | 12584 A. (D1-U) | 12850 A. | - | - | - | - | - | - | - | - | - |  |  |  |  |  |
| 9 | MBAR1414 | 11066.9 A. (D1_L) | 11077.6 A. (D2-U) | 12420.1 A. (D2_L) | 12690.3 A. (D1_U) |  | - | - | - | - | - | - | - | - |  |  |  |  |  |
| 10 | MBAR2413 | 11944.6 A. (D2_L) | 12840.8 A. (D2_U) | 12850 A. | - | - | - | - | - | - | - | - | - | - |  |  |  |  |  |

| Date | Bench | Problem | Type | Application/ soft | Action/advise |
| --- | --- | --- | --- | --- | --- |
| Wednesday_13_A | TBD1 | SSS674 : Resistance values were slightly out of range in PT13.1. | MTM |  | Guy was informed. He did the test again and informed Op team not to open NC and upload the file as it is. |
| Wednesday_13_N | TBB1 | SSS519: This MQ is an MQ without a heat exchanger and has to be tested at 4K. No training information is available for this case. | MTM |  | Guy and TC contacted. TC will get information from experts to start the training. TC asked meanwhile to perform the other tests. (HV, IAP, hf..) Pierre gave the information Thursday morning. |
| Wednesday_13_N | TBD2 | SSS695: Leak in the connection (MRB side, beam flange) found during the pumping. | MTM |  | MTM mechanic team changed the O-R joint |
| Wednesday_13_N | TBF1 | MB1345: While carrying out PT1-Part-III,Ac voltage of 8 volts is being feed. This voltage is not increasing beyond 1 Volt P-P. OP tried to change the pulse generator, change cable, verify of all the connections, but not success. | MTM |  | Guy contacted. |
| Thursday_14_N | TBA1 | MB2411: HV test failed at warm. Water in the CFB? | MTM |  |  |
| Friday_15_N Sunday_17_A | TBB1 | SSS519: Long training of MQ at 4.5 K (no exchanger). Ultimate not reached after 9 quenches. OP waiting fir TC instruction (ROBTC?) Magnet reached 11000 A after 20 Quenches. | MTM |  | TC and Andzrej advised to continue the training up to I=11000 A. |
| Friday_15_N | TBB1 | SSS519: PT8 not defined for this MQ at 4.5 K PT8 failed 2 times: quench in the clamp after several minutes. Same problem Sunday at the end of the training. MQ did not sustain 11000 A for 10 min, quench n the clamp | MTM ACR |  | TC advised to do PT8 at 10500 A. Problem of cool down? TC asked to warm_up the magnet. |
| Saturday_16_M | All | Cryo OK lost. | ACR |  | Flow meter down in one pumping unit. Cryo team worked with the second group. |
| Saturday_16_M | TBB2 | SSS516: Persistent trigger due to V_clamp_E this is continuing since 5AM morning. Hence EF clamp connection could not be done. | ICS |  | Guy came and found that - cable connection for the test was wrong +EF-V clamp Vtap was inverted. |
| Saturday_16_N | TBE1 | MB2402: HWREC error message: error in config files. Problem with IAP : lf application generating 2 sets of files. | MTM |  | OP shift leader contacted Guy. Guy said that he did not understand the origin of problem. He advised to wait 1 hour and re-start the test. OP shift leader reboot the SUN station and the application was unblocked. |
| Sunday_17_M | TBA1 | MB2411: Shaft in Ap1 cannot be introduced. | ICS |  | TC contacted. He advised to perform the other tests and wait for MM Monday. MTM mechanics found that the TRU was not aligned correctly by ICS. |
| Sunday_17_M | TBD | SSS669 and SSS695: Inversion the bench name in the display for the on-going tests. |  |  | OP team will ask Guy to correct the sequence. |
| Sunday_17_M | TBA1 | MB2411: still no cryo OK with this magnet | ACR |  |  |
| Sunday_17_M | TBB1 | SSS519: Ultimate current reached for MQ at 4.5 K but quench in bus-bar | ACR |  | TC asked to repeat the quench (cryo difficulty to cool down the clamp region) Quench repeated: same results. TC asked to warm-up the magnet. |
| Sunday_17_A | All | Problem with e-log book. The earlier entries sometimes get deleted. |  |  | ? |
| Sunday_17_N | TBE2 | MB2394: Capacitance Error during HV at warm Dipole-QH vs Gnd. Other steps are ok | MTM | HV | Guy called. QH investigation sequence all OK. Guy checked the IFS box and all connections. The Dipole-QH vs Ground still with capacitance problem. Another rack has been tried without success. Guy waited for Richard consultation Monday. |
| Sunday_17_N | All | Cryo team informed that due to pumping problem 1.9K temperature is not possible (second unit of pumping was stopped because of suspect noise). So no training possible on 1.9K Magnets. |  |  | Monday morning: intervention to repair the flow meter of the pumping unit 1. |
| Monday_18_A | TBE2 | MB2394: During HV test MB vs ground displays a high leak current. |  |  | Richard asked the cryo-team to purge the CFB circuits to remove the humidity. Still the problem. Richard asked to cool down the magnet. |